## Overview
Objective:
Validate phantom_curves.py automatic extraction of PDDs and profiles from RayStation phantom.
Method:
PatientID: Watertank cube [COMMISSIONING]
| Case: | LSt Watertank Phantom
| Plan: | Validation
Plan contains two beams.
15x10
MLC Striped pattern
Profiles and PDDs automatically extracted using script.
PDD and profiles at 10cm depth for 15x10 manually extracted using line dose tool.
Plotted on tabs overleaf for comparison.
Process repeated for INPLANE profile only for MLC Striped pattern at 10cm depth.
Conclusion:
Some residual error between RayStation line dose tool and script.
Cannot not be attributed to limits of fine motor skills.
Profiles are extracted at the right depth, with the right orientation.
Data can be opened in PTW Scanned Data Analysis.

## PDD
SCRIPT | [5 empty] | RayStation
BEGIN_SCAN_DATA | [5 empty] | #######################################################
| FORMAT=CC-Export V1.9 |  |  |  |  | #RayStationVersion: 13.1.0.144
| FILE_CREATION_DATE= |  |  |  |  | #PatientName: RayTraining 40x40x40 Phantom
| LAST_MODIFIED=11-Jan-2023 10:50:14 |  |  |  |  | #PatientId: Watertank cube
| BEGIN_SCAN 1 |  |  |  |  | #CoordinateSystem: IEC patient coordinate system
|  | TASK_NAME=RayStation Phantom Profiles & PDDs |  |  |  | #######################################################
|  | PROGRAM=RayStation |  |  |  | #LineName: 1
|  | COMMENT=RayStation Script Export |  |  |  | #DoseName: Beam dose: 15x10 (Validation, CT 1)
|  | MEAS_DATE=11-Jan-2023 10:50:14 |  |  |  | #DoseEngine: Clinical: Collapsed Cone v5.7
|  | LINAC=RayStation |  |  |  | #TreatmentMachine: RSL_TrueBeam Commission time: 11/20/2018 14:29:57
|  | MODALITY=X |  |  |  | #X [cm] | Y [cm] | Z [cm] | Dose [Gy] | COORD | cGy
|  | ISOCENTER=1000.00 |  |  |  | -0.027 | 25 | 30.146 | 0.005 | -30.146 | 0.5
|  | INPLANE_AXIS=Y |  |  |  | -0.027 | 25 | 30.136 | 0.029 | -30.136 | 2.9
|  | CROSSPLANE_AXIS=X |  |  |  | -0.027 | 25 | 30.126 | 0.052 | -30.126 | 5.2
|  | DEPTH_AXIS=Z |  |  |  | -0.027 | 25 | 30.116 | 0.075 | -30.116 | 7.5
|  | INPLANE_AXIS_DIR=GUN_TARGET |  |  |  | -0.027 | 25 | 30.106 | 0.099 | -30.106 | 9.9
|  | CROSSPLANE_AXIS_DIR=LEFT_RIGHT |  |  |  | -0.027 | 25 | 30.097 | 0.122 | -30.097 | 12.2
|  | DEPTH_AXIS_DIR=UP_DOWN |  |  |  | -0.027 | 25 | 30.087 | 0.145 | -30.087 | 14.5
|  | ENERGY=6.3 |  |  |  | -0.027 | 25 | 30.077 | 0.168 | -30.077 | 16.8
|  | NOMINAL_DMAX=20.00 |  |  |  | -0.027 | 25 | 30.067 | 0.192 | -30.067 | 19.2
|  | SSD=1000.0 |  |  |  | -0.027 | 25 | 30.057 | 0.215 | -30.057 | 21.5
|  | SCD=1000.00 |  |  |  | -0.027 | 25 | 30.047 | 0.238 | -30.047 | 23.8
|  | WEDGE_ANGLE=0.00 |  |  |  | -0.027 | 25 | 30.037 | 0.262 | -30.037 | 26.2
|  | FIELD_INPLANE=150.0 |  |  |  | -0.028 | 25 | 30.027 | 0.285 | -30.027 | 28.5
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -0.028 | 25 | 30.017 | 0.308 | -30.017 | 30.8
|  | FIELD_TYPE=IRREGULAR |  |  |  | -0.028 | 25 | 30.007 | 0.331 | -30.007 | 33.1
|  | GANTRY=0.00 |  |  |  | -0.028 | 25 | 29.997 | 0.355 | -29.997 | 35.5
|  | GANTRY_UPRIGHT_POSITION=0 |  |  |  | -0.028 | 25 | 29.987 | 0.378 | -29.987 | 37.8
|  | GANTRY_ROTATION=CW |  |  |  | -0.028 | 25 | 29.977 | 0.401 | -29.977 | 40.1
|  | COLL_ANGLE=0.00 |  |  |  | -0.028 | 25 | 29.967 | 0.425 | -29.967 | 42.5
|  | REF_FIELD_DEPTH=50.00 |  |  |  | -0.028 | 25 | 29.957 | 0.448 | -29.957 | 44.8
|  | REF_FIELD_DEFINED=ISOCENTER |  |  |  | -0.028 | 25 | 29.947 | 0.468 | -29.947 | 46.8
|  | FIELD_INPLANE=150.0 |  |  |  | -0.028 | 25 | 29.937 | 0.476 | -29.937 | 47.6
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -0.028 | 25 | 29.927 | 0.484 | -29.927 | 48.4
|  | SCAN_CURVETYPE=PDD |  |  |  | -0.028 | 25 | 29.917 | 0.492 | -29.917 | 49.2
|  | SCAN_OFFAXIS_INPLANE=0.00 |  |  |  | -0.028 | 25 | 29.907 | 0.5 | -29.907 | 50
|  | SCAN_OFFAXIS_CROSSPLANE=0.00 |  |  |  | -0.028 | 25 | 29.897 | 0.509 | -29.897 | 50.9
|  | SCAN_ANGLE=0.00 |  |  |  | -0.028 | 25 | 29.887 | 0.517 | -29.887 | 51.7
|  | SCAN_DIAGONAL=NOT_DIAGONAL |  |  |  | -0.028 | 25 | 29.877 | 0.525 | -29.877 | 52.5
|  | SCAN_DIRECTION=NEGATIVE |  |  |  | -0.028 | 25 | 29.867 | 0.533 | -29.867 | 53.3
|  | MEAS_MEDIUM=WATER |  |  |  | -0.028 | 25 | 29.857 | 0.542 | -29.857 | 54.2
|  | MEAS_TIME=1.000 |  |  |  | -0.028 | 25 | 29.847 | 0.55 | -29.847 | 55
|  | MEAS_UNIT=cGy |  |  |  | -0.028 | 25 | 29.837 | 0.558 | -29.837 | 55.8
|  | INCLINATION_ANGLE=0.00 |  |  |  | -0.028 | 25 | 29.827 | 0.566 | -29.827 | 56.6
|  | PRESSURE=1013.20 |  |  |  | -0.028 | 25 | 29.817 | 0.574 | -29.817 | 57.4
|  | TEMPERATURE=20.00 |  |  |  | -0.028 | 25 | 29.807 | 0.583 | -29.807 | 58.3
|  | NORM_TEMPERATURE=20.00 |  |  |  | -0.028 | 25 | 29.797 | 0.591 | -29.797 | 59.1
|  | CORRECTION_FACTOR=1.0000 |  |  |  | -0.028 | 25 | 29.787 | 0.599 | -29.787 | 59.9
|  | BEGIN_DATA |  |  |  | -0.028 | 25 | 29.777 | 0.607 | -29.777 | 60.7
|  |  | -30.05 | 23.289 | 1 | -0.028 | 25 | 29.767 | 0.615 | -29.767 | 61.5
|  |  | -30 | 34.934 | 1 | -0.028 | 25 | 29.757 | 0.624 | -29.757 | 62.4
|  |  | -29.95 | 46.577 | 1 | -0.028 | 25 | 29.747 | 0.631 | -29.747 | 63.1
|  |  | -29.9 | 50.683 | 1 | -0.028 | 25 | 29.737 | 0.634 | -29.737 | 63.4
|  |  | -29.85 | 54.789 | 1 | -0.028 | 25 | 29.727 | 0.638 | -29.727 | 63.8
|  |  | -29.8 | 58.894 | 1 | -0.028 | 25 | 29.717 | 0.642 | -29.717 | 64.2
|  |  | -29.75 | 63 | 1 | -0.028 | 25 | 29.707 | 0.645 | -29.707 | 64.5
|  |  | -29.7 | 64.803 | 1 | -0.028 | 25 | 29.697 | 0.649 | -29.697 | 64.9
|  |  | -29.65 | 66.606 | 1 | -0.028 | 25 | 29.687 | 0.652 | -29.687 | 65.2
|  |  | -29.6 | 68.409 | 1 | -0.028 | 25 | 29.677 | 0.656 | -29.677 | 65.6
|  |  | -29.55 | 70.212 | 1 | -0.028 | 25 | 29.667 | 0.66 | -29.667 | 66
|  |  | -29.5 | 71.172 | 1 | -0.028 | 25 | 29.657 | 0.663 | -29.657 | 66.3
|  |  | -29.45 | 72.132 | 1 | -0.028 | 25 | 29.647 | 0.667 | -29.647 | 66.7
|  |  | -29.4 | 73.092 | 1 | -0.028 | 25 | 29.637 | 0.67 | -29.637 | 67
|  |  | -29.35 | 74.052 | 1 | -0.028 | 25 | 29.627 | 0.674 | -29.627 | 67.4
|  |  | -29.3 | 74.541 | 1 | -0.028 | 25 | 29.617 | 0.678 | -29.617 | 67.8
|  |  | -29.25 | 75.03 | 1 | -0.028 | 25 | 29.607 | 0.681 | -29.607 | 68.1
|  |  | -29.2 | 75.519 | 1 | -0.028 | 25 | 29.597 | 0.685 | -29.597 | 68.5
|  |  | -29.15 | 76.008 | 1 | -0.028 | 25 | 29.587 | 0.688 | -29.587 | 68.8
|  |  | -29.1 | 76.294 | 1 | -0.028 | 25 | 29.577 | 0.692 | -29.577 | 69.2
|  |  | -29.05 | 76.579 | 1 | -0.028 | 25 | 29.567 | 0.696 | -29.567 | 69.6
|  |  | -29 | 76.864 | 1 | -0.028 | 25 | 29.557 | 0.699 | -29.557 | 69.9
|  |  | -28.95 | 77.149 | 1 | -0.028 | 25 | 29.547 | 0.703 | -29.547 | 70.3
|  |  | -28.9 | 77.268 | 1 | -0.028 | 25 | 29.537 | 0.704 | -29.537 | 70.4
|  |  | -28.85 | 77.387 | 1 | -0.028 | 25 | 29.527 | 0.706 | -29.527 | 70.6
|  |  | -28.8 | 77.506 | 1 | -0.028 | 25 | 29.517 | 0.708 | -29.517 | 70.8
|  |  | -28.75 | 77.625 | 1 | -0.028 | 25 | 29.507 | 0.71 | -29.507 | 71
|  |  | -28.7 | 77.609 | 1 | -0.028 | 25 | 29.497 | 0.712 | -29.497 | 71.2
|  |  | -28.65 | 77.592 | 1 | -0.028 | 25 | 29.487 | 0.714 | -29.487 | 71.4
|  |  | -28.6 | 77.576 | 1 | -0.028 | 25 | 29.477 | 0.716 | -29.477 | 71.6
|  |  | -28.55 | 77.56 | 1 | -0.028 | 25 | 29.467 | 0.718 | -29.467 | 71.8
|  |  | -28.5 | 77.518 | 1 | -0.028 | 25 | 29.457 | 0.72 | -29.457 | 72
|  |  | -28.45 | 77.476 | 1 | -0.028 | 25 | 29.447 | 0.722 | -29.447 | 72.2
|  |  | -28.4 | 77.434 | 1 | -0.028 | 25 | 29.437 | 0.724 | -29.437 | 72.4
|  |  | -28.35 | 77.392 | 1 | -0.028 | 25 | 29.427 | 0.726 | -29.427 | 72.6
|  |  | -28.3 | 77.279 | 1 | -0.028 | 25 | 29.417 | 0.728 | -29.417 | 72.8
|  |  | -28.25 | 77.167 | 1 | -0.028 | 25 | 29.407 | 0.729 | -29.407 | 72.9
|  |  | -28.2 | 77.055 | 1 | -0.028 | 25 | 29.397 | 0.731 | -29.397 | 73.1
|  |  | -28.15 | 76.943 | 1 | -0.028 | 25 | 29.387 | 0.733 | -29.387 | 73.3
|  |  | -28.1 | 76.837 | 1 | -0.028 | 25 | 29.377 | 0.735 | -29.377 | 73.5
|  |  | -28.05 | 76.73 | 1 | -0.028 | 25 | 29.367 | 0.737 | -29.367 | 73.7
|  |  | -28 | 76.624 | 1 | -0.028 | 25 | 29.357 | 0.739 | -29.357 | 73.9
|  |  | -27.95 | 76.517 | 1 | -0.028 | 25 | 29.347 | 0.741 | -29.347 | 74.1
|  |  | -27.9 | 76.365 | 1 | -0.028 | 25 | 29.337 | 0.742 | -29.337 | 74.2
|  |  | -27.85 | 76.213 | 1 | -0.029 | 25 | 29.327 | 0.743 | -29.327 | 74.3
|  |  | -27.8 | 76.061 | 1 | -0.029 | 25 | 29.317 | 0.744 | -29.317 | 74.4
|  |  | -27.75 | 75.909 | 1 | -0.029 | 25 | 29.307 | 0.745 | -29.307 | 74.5
|  |  | -27.7 | 75.77 | 1 | -0.029 | 25 | 29.297 | 0.746 | -29.297 | 74.6
|  |  | -27.65 | 75.631 | 1 | -0.029 | 25 | 29.287 | 0.747 | -29.287 | 74.7
|  |  | -27.6 | 75.493 | 1 | -0.029 | 25 | 29.277 | 0.748 | -29.277 | 74.8
|  |  | -27.55 | 75.354 | 1 | -0.029 | 25 | 29.267 | 0.749 | -29.267 | 74.9
|  |  | -27.5 | 75.213 | 1 | -0.029 | 25 | 29.257 | 0.75 | -29.257 | 75
|  |  | -27.45 | 75.072 | 1 | -0.029 | 25 | 29.247 | 0.751 | -29.247 | 75.1
|  |  | -27.4 | 74.93 | 1 | -0.029 | 25 | 29.237 | 0.751 | -29.237 | 75.1
|  |  | -27.35 | 74.789 | 1 | -0.029 | 25 | 29.227 | 0.752 | -29.227 | 75.2
|  |  | -27.3 | 74.614 | 1 | -0.029 | 25 | 29.217 | 0.753 | -29.217 | 75.3
|  |  | -27.25 | 74.439 | 1 | -0.029 | 25 | 29.207 | 0.754 | -29.207 | 75.4
|  |  | -27.2 | 74.265 | 1 | -0.029 | 25 | 29.197 | 0.755 | -29.197 | 75.5
|  |  | -27.15 | 74.09 | 1 | -0.029 | 25 | 29.187 | 0.756 | -29.187 | 75.6
|  |  | -27.1 | 73.942 | 1 | -0.029 | 25 | 29.177 | 0.757 | -29.177 | 75.7
|  |  | -27.05 | 73.794 | 1 | -0.029 | 25 | 29.167 | 0.758 | -29.167 | 75.8
|  |  | -27 | 73.646 | 1 | -0.029 | 25 | 29.157 | 0.759 | -29.157 | 75.9
|  |  | -26.95 | 73.498 | 1 | -0.029 | 25 | 29.147 | 0.76 | -29.147 | 76
|  |  | -26.9 | 73.328 | 1 | -0.029 | 25 | 29.137 | 0.761 | -29.137 | 76.1
|  |  | -26.85 | 73.158 | 1 | -0.029 | 25 | 29.127 | 0.761 | -29.127 | 76.1
|  |  | -26.8 | 72.987 | 1 | -0.029 | 25 | 29.117 | 0.762 | -29.117 | 76.2
|  |  | -26.75 | 72.817 | 1 | -0.029 | 25 | 29.107 | 0.762 | -29.107 | 76.2
|  |  | -26.7 | 72.672 | 1 | -0.029 | 25 | 29.097 | 0.763 | -29.097 | 76.3
|  |  | -26.65 | 72.527 | 1 | -0.029 | 25 | 29.087 | 0.764 | -29.087 | 76.4
|  |  | -26.6 | 72.382 | 1 | -0.029 | 25 | 29.077 | 0.764 | -29.077 | 76.4
|  |  | -26.55 | 72.237 | 1 | -0.029 | 25 | 29.067 | 0.765 | -29.067 | 76.5
|  |  | -26.5 | 72.092 | 1 | -0.029 | 25 | 29.057 | 0.765 | -29.057 | 76.5
|  |  | -26.45 | 71.946 | 1 | -0.029 | 25 | 29.047 | 0.766 | -29.047 | 76.6
|  |  | -26.4 | 71.801 | 1 | -0.029 | 25 | 29.037 | 0.766 | -29.037 | 76.6
|  |  | -26.35 | 71.655 | 1 | -0.029 | 25 | 29.027 | 0.767 | -29.027 | 76.7
|  |  | -26.3 | 71.503 | 1 | -0.029 | 25 | 29.017 | 0.768 | -29.017 | 76.8
|  |  | -26.25 | 71.351 | 1 | -0.029 | 25 | 29.007 | 0.768 | -29.007 | 76.8
|  |  | -26.2 | 71.199 | 1 | -0.029 | 25 | 28.997 | 0.769 | -28.997 | 76.9
|  |  | -26.15 | 71.047 | 1 | -0.029 | 25 | 28.987 | 0.769 | -28.987 | 76.9
|  |  | -26.1 | 70.886 | 1 | -0.029 | 25 | 28.977 | 0.77 | -28.977 | 77
|  |  | -26.05 | 70.726 | 1 | -0.029 | 25 | 28.967 | 0.77 | -28.967 | 77
|  |  | -26 | 70.566 | 1 | -0.029 | 25 | 28.957 | 0.771 | -28.957 | 77.1
|  |  | -25.95 | 70.406 | 1 | -0.029 | 25 | 28.947 | 0.772 | -28.947 | 77.2
|  |  | -25.9 | 70.246 | 1 | -0.029 | 25 | 28.937 | 0.772 | -28.937 | 77.2
|  |  | -25.85 | 70.087 | 1 | -0.029 | 25 | 28.927 | 0.772 | -28.927 | 77.2
|  |  | -25.8 | 69.928 | 1 | -0.029 | 25 | 28.917 | 0.772 | -28.917 | 77.2
|  |  | -25.75 | 69.769 | 1 | -0.029 | 25 | 28.907 | 0.772 | -28.907 | 77.2
|  |  | -25.7 | 69.611 | 1 | -0.029 | 25 | 28.897 | 0.773 | -28.897 | 77.3
|  |  | -25.65 | 69.453 | 1 | -0.029 | 25 | 28.887 | 0.773 | -28.887 | 77.3
|  |  | -25.6 | 69.295 | 1 | -0.029 | 25 | 28.877 | 0.773 | -28.877 | 77.3
|  |  | -25.55 | 69.137 | 1 | -0.029 | 25 | 28.867 | 0.773 | -28.867 | 77.3
|  |  | -25.5 | 68.98 | 1 | -0.029 | 25 | 28.857 | 0.774 | -28.857 | 77.4
|  |  | -25.45 | 68.823 | 1 | -0.029 | 25 | 28.847 | 0.774 | -28.847 | 77.4
|  |  | -25.4 | 68.666 | 1 | -0.029 | 25 | 28.837 | 0.774 | -28.837 | 77.4
|  |  | -25.35 | 68.509 | 1 | -0.029 | 25 | 28.827 | 0.774 | -28.827 | 77.4
|  |  | -25.3 | 68.345 | 1 | -0.029 | 25 | 28.817 | 0.775 | -28.817 | 77.5
|  |  | -25.25 | 68.182 | 1 | -0.029 | 25 | 28.807 | 0.775 | -28.807 | 77.5
|  |  | -25.2 | 68.019 | 1 | -0.029 | 25 | 28.797 | 0.775 | -28.797 | 77.5
|  |  | -25.15 | 67.856 | 1 | -0.029 | 25 | 28.787 | 0.775 | -28.787 | 77.5
|  |  | -25.1 | 67.694 | 1 | -0.029 | 25 | 28.777 | 0.776 | -28.777 | 77.6
|  |  | -25.05 | 67.532 | 1 | -0.029 | 25 | 28.767 | 0.776 | -28.767 | 77.6
|  |  | -25 | 67.37 | 1 | -0.029 | 25 | 28.757 | 0.776 | -28.757 | 77.6
|  |  | -24.95 | 67.208 | 1 | -0.029 | 25 | 28.747 | 0.776 | -28.747 | 77.6
|  |  | -24.9 | 67.048 | 1 | -0.029 | 25 | 28.737 | 0.776 | -28.737 | 77.6
|  |  | -24.85 | 66.887 | 1 | -0.029 | 25 | 28.727 | 0.776 | -28.727 | 77.6
|  |  | -24.8 | 66.726 | 1 | -0.029 | 25 | 28.717 | 0.776 | -28.717 | 77.6
|  |  | -24.75 | 66.565 | 1 | -0.029 | 25 | 28.707 | 0.776 | -28.707 | 77.6
|  |  | -24.7 | 66.405 | 1 | -0.029 | 25 | 28.697 | 0.776 | -28.697 | 77.6
|  |  | -24.65 | 66.245 | 1 | -0.029 | 25 | 28.687 | 0.776 | -28.687 | 77.6
|  |  | -24.6 | 66.086 | 1 | -0.029 | 25 | 28.677 | 0.776 | -28.677 | 77.6
|  |  | -24.55 | 65.926 | 1 | -0.029 | 25 | 28.667 | 0.776 | -28.667 | 77.6
|  |  | -24.5 | 65.772 | 1 | -0.029 | 25 | 28.657 | 0.776 | -28.657 | 77.6
|  |  | -24.45 | 65.617 | 1 | -0.029 | 25 | 28.647 | 0.776 | -28.647 | 77.6
|  |  | -24.4 | 65.463 | 1 | -0.029 | 25 | 28.637 | 0.776 | -28.637 | 77.6
|  |  | -24.35 | 65.308 | 1 | -0.03 | 25 | 28.627 | 0.776 | -28.627 | 77.6
|  |  | -24.3 | 65.155 | 1 | -0.03 | 25 | 28.617 | 0.776 | -28.617 | 77.6
|  |  | -24.25 | 65.001 | 1 | -0.03 | 25 | 28.607 | 0.776 | -28.607 | 77.6
|  |  | -24.2 | 64.848 | 1 | -0.03 | 25 | 28.597 | 0.776 | -28.597 | 77.6
|  |  | -24.15 | 64.694 | 1 | -0.03 | 25 | 28.587 | 0.776 | -28.587 | 77.6
|  |  | -24.1 | 64.542 | 1 | -0.03 | 25 | 28.577 | 0.776 | -28.577 | 77.6
|  |  | -24.05 | 64.39 | 1 | -0.03 | 25 | 28.567 | 0.776 | -28.567 | 77.6
|  |  | -24 | 64.237 | 1 | -0.03 | 25 | 28.557 | 0.776 | -28.557 | 77.6
|  |  | -23.95 | 64.085 | 1 | -0.03 | 25 | 28.547 | 0.776 | -28.547 | 77.6
|  |  | -23.9 | 63.933 | 1 | -0.03 | 25 | 28.537 | 0.775 | -28.537 | 77.5
|  |  | -23.85 | 63.782 | 1 | -0.03 | 25 | 28.527 | 0.775 | -28.527 | 77.5
|  |  | -23.8 | 63.631 | 1 | -0.03 | 25 | 28.517 | 0.775 | -28.517 | 77.5
|  |  | -23.75 | 63.479 | 1 | -0.03 | 25 | 28.507 | 0.775 | -28.507 | 77.5
|  |  | -23.7 | 63.325 | 1 | -0.03 | 25 | 28.497 | 0.775 | -28.497 | 77.5
|  |  | -23.65 | 63.17 | 1 | -0.03 | 25 | 28.487 | 0.775 | -28.487 | 77.5
|  |  | -23.6 | 63.015 | 1 | -0.03 | 25 | 28.477 | 0.775 | -28.477 | 77.5
|  |  | -23.55 | 62.861 | 1 | -0.03 | 25 | 28.467 | 0.775 | -28.467 | 77.5
|  |  | -23.5 | 62.707 | 1 | -0.03 | 25 | 28.457 | 0.775 | -28.457 | 77.5
|  |  | -23.45 | 62.554 | 1 | -0.03 | 25 | 28.447 | 0.775 | -28.447 | 77.5
|  |  | -23.4 | 62.4 | 1 | -0.03 | 25 | 28.437 | 0.775 | -28.437 | 77.5
|  |  | -23.35 | 62.246 | 1 | -0.03 | 25 | 28.427 | 0.775 | -28.427 | 77.5
|  |  | -23.3 | 62.094 | 1 | -0.03 | 25 | 28.417 | 0.774 | -28.417 | 77.4
|  |  | -23.25 | 61.941 | 1 | -0.03 | 25 | 28.407 | 0.774 | -28.407 | 77.4
|  |  | -23.2 | 61.789 | 1 | -0.03 | 25 | 28.397 | 0.774 | -28.397 | 77.4
|  |  | -23.15 | 61.636 | 1 | -0.03 | 25 | 28.387 | 0.774 | -28.387 | 77.4
|  |  | -23.1 | 61.485 | 1 | -0.03 | 25 | 28.377 | 0.774 | -28.377 | 77.4
|  |  | -23.05 | 61.333 | 1 | -0.03 | 25 | 28.367 | 0.774 | -28.367 | 77.4
|  |  | -23 | 61.181 | 1 | -0.03 | 25 | 28.357 | 0.774 | -28.357 | 77.4
|  |  | -22.95 | 61.03 | 1 | -0.03 | 25 | 28.347 | 0.774 | -28.347 | 77.4
|  |  | -22.9 | 60.879 | 1 | -0.03 | 25 | 28.337 | 0.774 | -28.337 | 77.4
|  |  | -22.85 | 60.728 | 1 | -0.03 | 25 | 28.327 | 0.773 | -28.327 | 77.3
|  |  | -22.8 | 60.577 | 1 | -0.03 | 25 | 28.317 | 0.773 | -28.317 | 77.3
|  |  | -22.75 | 60.426 | 1 | -0.03 | 25 | 28.308 | 0.773 | -28.308 | 77.3
|  |  | -22.7 | 60.276 | 1 | -0.03 | 25 | 28.298 | 0.773 | -28.298 | 77.3
|  |  | -22.65 | 60.126 | 1 | -0.03 | 25 | 28.288 | 0.773 | -28.288 | 77.3
|  |  | -22.6 | 59.976 | 1 | -0.03 | 25 | 28.278 | 0.772 | -28.278 | 77.2
|  |  | -22.55 | 59.826 | 1 | -0.03 | 25 | 28.268 | 0.772 | -28.268 | 77.2
|  |  | -22.5 | 59.677 | 1 | -0.03 | 25 | 28.258 | 0.772 | -28.258 | 77.2
|  |  | -22.45 | 59.528 | 1 | -0.03 | 25 | 28.248 | 0.772 | -28.248 | 77.2
|  |  | -22.4 | 59.379 | 1 | -0.03 | 25 | 28.238 | 0.771 | -28.238 | 77.1
|  |  | -22.35 | 59.23 | 1 | -0.03 | 25 | 28.228 | 0.771 | -28.228 | 77.1
|  |  | -22.3 | 59.082 | 1 | -0.03 | 25 | 28.218 | 0.771 | -28.218 | 77.1
|  |  | -22.25 | 58.934 | 1 | -0.03 | 25 | 28.208 | 0.771 | -28.208 | 77.1
|  |  | -22.2 | 58.786 | 1 | -0.03 | 25 | 28.198 | 0.771 | -28.198 | 77.1
|  |  | -22.15 | 58.638 | 1 | -0.03 | 25 | 28.188 | 0.77 | -28.188 | 77
|  |  | -22.1 | 58.49 | 1 | -0.03 | 25 | 28.178 | 0.77 | -28.178 | 77
|  |  | -22.05 | 58.341 | 1 | -0.03 | 25 | 28.168 | 0.77 | -28.168 | 77
|  |  | -22 | 58.193 | 1 | -0.03 | 25 | 28.158 | 0.77 | -28.158 | 77
|  |  | -21.95 | 58.044 | 1 | -0.03 | 25 | 28.148 | 0.769 | -28.148 | 76.9
|  |  | -21.9 | 57.897 | 1 | -0.03 | 25 | 28.138 | 0.769 | -28.138 | 76.9
|  |  | -21.85 | 57.75 | 1 | -0.03 | 25 | 28.128 | 0.769 | -28.128 | 76.9
|  |  | -21.8 | 57.602 | 1 | -0.03 | 25 | 28.118 | 0.769 | -28.118 | 76.9
|  |  | -21.75 | 57.455 | 1 | -0.03 | 25 | 28.108 | 0.769 | -28.108 | 76.9
|  |  | -21.7 | 57.308 | 1 | -0.03 | 25 | 28.098 | 0.768 | -28.098 | 76.8
|  |  | -21.65 | 57.162 | 1 | -0.03 | 25 | 28.088 | 0.768 | -28.088 | 76.8
|  |  | -21.6 | 57.015 | 1 | -0.03 | 25 | 28.078 | 0.768 | -28.078 | 76.8
|  |  | -21.55 | 56.869 | 1 | -0.03 | 25 | 28.068 | 0.768 | -28.068 | 76.8
|  |  | -21.5 | 56.723 | 1 | -0.03 | 25 | 28.058 | 0.767 | -28.058 | 76.7
|  |  | -21.45 | 56.578 | 1 | -0.03 | 25 | 28.048 | 0.767 | -28.048 | 76.7
|  |  | -21.4 | 56.432 | 1 | -0.03 | 25 | 28.038 | 0.767 | -28.038 | 76.7
|  |  | -21.35 | 56.287 | 1 | -0.03 | 25 | 28.028 | 0.767 | -28.028 | 76.7
|  |  | -21.3 | 56.139 | 1 | -0.03 | 25 | 28.018 | 0.767 | -28.018 | 76.7
|  |  | -21.25 | 55.991 | 1 | -0.03 | 25 | 28.008 | 0.766 | -28.008 | 76.6
|  |  | -21.2 | 55.844 | 1 | -0.03 | 25 | 27.998 | 0.766 | -27.998 | 76.6
|  |  | -21.15 | 55.696 | 1 | -0.03 | 25 | 27.988 | 0.766 | -27.988 | 76.6
|  |  | -21.1 | 55.549 | 1 | -0.03 | 25 | 27.978 | 0.766 | -27.978 | 76.6
|  |  | -21.05 | 55.403 | 1 | -0.03 | 25 | 27.968 | 0.766 | -27.968 | 76.6
|  |  | -21 | 55.256 | 1 | -0.03 | 25 | 27.958 | 0.765 | -27.958 | 76.5
|  |  | -20.95 | 55.11 | 1 | -0.03 | 25 | 27.948 | 0.765 | -27.948 | 76.5
|  |  | -20.9 | 54.964 | 1 | -0.03 | 25 | 27.938 | 0.765 | -27.938 | 76.5
|  |  | -20.85 | 54.818 | 1 | -0.031 | 25 | 27.928 | 0.765 | -27.928 | 76.5
|  |  | -20.8 | 54.672 | 1 | -0.031 | 25 | 27.918 | 0.764 | -27.918 | 76.4
|  |  | -20.75 | 54.527 | 1 | -0.031 | 25 | 27.908 | 0.764 | -27.908 | 76.4
|  |  | -20.7 | 54.382 | 1 | -0.031 | 25 | 27.898 | 0.764 | -27.898 | 76.4
|  |  | -20.65 | 54.237 | 1 | -0.031 | 25 | 27.888 | 0.763 | -27.888 | 76.3
|  |  | -20.6 | 54.093 | 1 | -0.031 | 25 | 27.878 | 0.763 | -27.878 | 76.3
|  |  | -20.55 | 53.948 | 1 | -0.031 | 25 | 27.868 | 0.763 | -27.868 | 76.3
|  |  | -20.5 | 53.81 | 1 | -0.031 | 25 | 27.858 | 0.762 | -27.858 | 76.2
|  |  | -20.45 | 53.673 | 1 | -0.031 | 25 | 27.848 | 0.762 | -27.848 | 76.2
|  |  | -20.4 | 53.535 | 1 | -0.031 | 25 | 27.838 | 0.762 | -27.838 | 76.2
|  |  | -20.35 | 53.398 | 1 | -0.031 | 25 | 27.828 | 0.761 | -27.828 | 76.1
|  |  | -20.3 | 53.261 | 1 | -0.031 | 25 | 27.818 | 0.761 | -27.818 | 76.1
|  |  | -20.25 | 53.124 | 1 | -0.031 | 25 | 27.808 | 0.761 | -27.808 | 76.1
|  |  | -20.2 | 52.988 | 1 | -0.031 | 25 | 27.798 | 0.761 | -27.798 | 76.1
|  |  | -20.15 | 52.851 | 1 | -0.031 | 25 | 27.788 | 0.76 | -27.788 | 76
|  |  | -20.1 | 52.715 | 1 | -0.031 | 25 | 27.778 | 0.76 | -27.778 | 76
|  |  | -20.05 | 52.579 | 1 | -0.031 | 25 | 27.768 | 0.76 | -27.768 | 76
|  |  | -20 | 52.444 | 1 | -0.031 | 25 | 27.758 | 0.759 | -27.758 | 75.9
|  |  | -19.95 | 52.308 | 1 | -0.031 | 25 | 27.748 | 0.759 | -27.748 | 75.9
|  |  | -19.9 | 52.173 | 1 | -0.031 | 25 | 27.738 | 0.759 | -27.738 | 75.9
|  |  | -19.85 | 52.038 | 1 | -0.031 | 25 | 27.728 | 0.758 | -27.728 | 75.8
|  |  | -19.8 | 51.903 | 1 | -0.031 | 25 | 27.718 | 0.758 | -27.718 | 75.8
|  |  | -19.75 | 51.768 | 1 | -0.031 | 25 | 27.708 | 0.758 | -27.708 | 75.8
|  |  | -19.7 | 51.633 | 1 | -0.031 | 25 | 27.698 | 0.758 | -27.698 | 75.8
|  |  | -19.65 | 51.497 | 1 | -0.031 | 25 | 27.688 | 0.757 | -27.688 | 75.7
|  |  | -19.6 | 51.361 | 1 | -0.031 | 25 | 27.678 | 0.757 | -27.678 | 75.7
|  |  | -19.55 | 51.226 | 1 | -0.031 | 25 | 27.668 | 0.757 | -27.668 | 75.7
|  |  | -19.5 | 51.091 | 1 | -0.031 | 25 | 27.658 | 0.757 | -27.658 | 75.7
|  |  | -19.45 | 50.956 | 1 | -0.031 | 25 | 27.648 | 0.756 | -27.648 | 75.6
|  |  | -19.4 | 50.822 | 1 | -0.031 | 25 | 27.638 | 0.756 | -27.638 | 75.6
|  |  | -19.35 | 50.687 | 1 | -0.031 | 25 | 27.628 | 0.756 | -27.628 | 75.6
|  |  | -19.3 | 50.553 | 1 | -0.031 | 25 | 27.618 | 0.755 | -27.618 | 75.5
|  |  | -19.25 | 50.419 | 1 | -0.031 | 25 | 27.608 | 0.755 | -27.608 | 75.5
|  |  | -19.2 | 50.285 | 1 | -0.031 | 25 | 27.598 | 0.755 | -27.598 | 75.5
|  |  | -19.15 | 50.152 | 1 | -0.031 | 25 | 27.588 | 0.755 | -27.588 | 75.5
|  |  | -19.1 | 50.019 | 1 | -0.031 | 25 | 27.578 | 0.754 | -27.578 | 75.4
|  |  | -19.05 | 49.886 | 1 | -0.031 | 25 | 27.568 | 0.754 | -27.568 | 75.4
|  |  | -19 | 49.753 | 1 | -0.031 | 25 | 27.558 | 0.754 | -27.558 | 75.4
|  |  | -18.95 | 49.62 | 1 | -0.031 | 25 | 27.548 | 0.753 | -27.548 | 75.3
|  |  | -18.9 | 49.49 | 1 | -0.031 | 25 | 27.538 | 0.753 | -27.538 | 75.3
|  |  | -18.85 | 49.361 | 1 | -0.031 | 25 | 27.528 | 0.753 | -27.528 | 75.3
|  |  | -18.8 | 49.232 | 1 | -0.031 | 25 | 27.518 | 0.753 | -27.518 | 75.3
|  |  | -18.75 | 49.103 | 1 | -0.031 | 25 | 27.508 | 0.752 | -27.508 | 75.2
|  |  | -18.7 | 48.974 | 1 | -0.031 | 25 | 27.498 | 0.752 | -27.498 | 75.2
|  |  | -18.65 | 48.846 | 1 | -0.031 | 25 | 27.488 | 0.752 | -27.488 | 75.2
|  |  | -18.6 | 48.718 | 1 | -0.031 | 25 | 27.478 | 0.752 | -27.478 | 75.2
|  |  | -18.55 | 48.589 | 1 | -0.031 | 25 | 27.468 | 0.751 | -27.468 | 75.1
|  |  | -18.5 | 48.462 | 1 | -0.031 | 25 | 27.458 | 0.751 | -27.458 | 75.1
|  |  | -18.45 | 48.334 | 1 | -0.031 | 25 | 27.448 | 0.751 | -27.448 | 75.1
|  |  | -18.4 | 48.207 | 1 | -0.031 | 25 | 27.438 | 0.75 | -27.438 | 75
|  |  | -18.35 | 48.079 | 1 | -0.031 | 25 | 27.428 | 0.75 | -27.428 | 75
|  |  | -18.3 | 47.952 | 1 | -0.031 | 25 | 27.418 | 0.75 | -27.418 | 75
|  |  | -18.25 | 47.826 | 1 | -0.031 | 25 | 27.408 | 0.75 | -27.408 | 75
|  |  | -18.2 | 47.699 | 1 | -0.031 | 25 | 27.398 | 0.749 | -27.398 | 74.9
|  |  | -18.15 | 47.572 | 1 | -0.031 | 25 | 27.388 | 0.749 | -27.388 | 74.9
|  |  | -18.1 | 47.448 | 1 | -0.031 | 25 | 27.378 | 0.749 | -27.378 | 74.9
|  |  | -18.05 | 47.323 | 1 | -0.031 | 25 | 27.368 | 0.748 | -27.368 | 74.8
|  |  | -18 | 47.199 | 1 | -0.031 | 25 | 27.358 | 0.748 | -27.358 | 74.8
|  |  | -17.95 | 47.074 | 1 | -0.031 | 25 | 27.348 | 0.748 | -27.348 | 74.8
|  |  | -17.9 | 46.95 | 1 | -0.031 | 25 | 27.338 | 0.747 | -27.338 | 74.7
|  |  | -17.85 | 46.827 | 1 | -0.031 | 25 | 27.328 | 0.747 | -27.328 | 74.7
|  |  | -17.8 | 46.703 | 1 | -0.031 | 25 | 27.318 | 0.747 | -27.318 | 74.7
|  |  | -17.75 | 46.579 | 1 | -0.031 | 25 | 27.308 | 0.746 | -27.308 | 74.6
|  |  | -17.7 | 46.457 | 1 | -0.031 | 25 | 27.298 | 0.746 | -27.298 | 74.6
|  |  | -17.65 | 46.334 | 1 | -0.031 | 25 | 27.288 | 0.746 | -27.288 | 74.6
|  |  | -17.6 | 46.211 | 1 | -0.031 | 25 | 27.278 | 0.745 | -27.278 | 74.5
|  |  | -17.55 | 46.088 | 1 | -0.031 | 25 | 27.268 | 0.745 | -27.268 | 74.5
|  |  | -17.5 | 45.966 | 1 | -0.031 | 25 | 27.258 | 0.745 | -27.258 | 74.5
|  |  | -17.45 | 45.844 | 1 | -0.031 | 25 | 27.248 | 0.744 | -27.248 | 74.4
|  |  | -17.4 | 45.722 | 1 | -0.031 | 25 | 27.238 | 0.744 | -27.238 | 74.4
|  |  | -17.35 | 45.6 | 1 | -0.032 | 25 | 27.228 | 0.744 | -27.228 | 74.4
|  |  | -17.3 | 45.475 | 1 | -0.032 | 25 | 27.218 | 0.743 | -27.218 | 74.3
|  |  | -17.25 | 45.351 | 1 | -0.032 | 25 | 27.208 | 0.743 | -27.208 | 74.3
|  |  | -17.2 | 45.227 | 1 | -0.032 | 25 | 27.198 | 0.743 | -27.198 | 74.3
|  |  | -17.15 | 45.103 | 1 | -0.032 | 25 | 27.188 | 0.742 | -27.188 | 74.2
|  |  | -17.1 | 44.98 | 1 | -0.032 | 25 | 27.178 | 0.742 | -27.178 | 74.2
|  |  | -17.05 | 44.857 | 1 | -0.032 | 25 | 27.168 | 0.742 | -27.168 | 74.2
|  |  | -17 | 44.733 | 1 | -0.032 | 25 | 27.158 | 0.741 | -27.158 | 74.1
|  |  | -16.95 | 44.61 | 1 | -0.032 | 25 | 27.148 | 0.741 | -27.148 | 74.1
|  |  | -16.9 | 44.488 | 1 | -0.032 | 25 | 27.138 | 0.741 | -27.138 | 74.1
|  |  | -16.85 | 44.365 | 1 | -0.032 | 25 | 27.128 | 0.74 | -27.128 | 74
|  |  | -16.8 | 44.243 | 1 | -0.032 | 25 | 27.118 | 0.74 | -27.118 | 74
|  |  | -16.75 | 44.12 | 1 | -0.032 | 25 | 27.108 | 0.74 | -27.108 | 74
|  |  | -16.7 | 43.999 | 1 | -0.032 | 25 | 27.098 | 0.739 | -27.098 | 73.9
|  |  | -16.65 | 43.877 | 1 | -0.032 | 25 | 27.088 | 0.739 | -27.088 | 73.9
|  |  | -16.6 | 43.755 | 1 | -0.032 | 25 | 27.078 | 0.739 | -27.078 | 73.9
|  |  | -16.55 | 43.634 | 1 | -0.032 | 25 | 27.068 | 0.738 | -27.068 | 73.8
|  |  | -16.5 | 43.517 | 1 | -0.032 | 25 | 27.058 | 0.738 | -27.058 | 73.8
|  |  | -16.45 | 43.4 | 1 | -0.032 | 25 | 27.048 | 0.738 | -27.048 | 73.8
|  |  | -16.4 | 43.282 | 1 | -0.032 | 25 | 27.038 | 0.738 | -27.038 | 73.8
|  |  | -16.35 | 43.165 | 1 | -0.032 | 25 | 27.028 | 0.737 | -27.028 | 73.7
|  |  | -16.3 | 43.049 | 1 | -0.032 | 25 | 27.018 | 0.737 | -27.018 | 73.7
|  |  | -16.25 | 42.933 | 1 | -0.032 | 25 | 27.008 | 0.737 | -27.008 | 73.7
|  |  | -16.2 | 42.816 | 1 | -0.032 | 25 | 26.998 | 0.736 | -26.998 | 73.6
|  |  | -16.15 | 42.7 | 1 | -0.032 | 25 | 26.988 | 0.736 | -26.988 | 73.6
|  |  | -16.1 | 42.585 | 1 | -0.032 | 25 | 26.978 | 0.736 | -26.978 | 73.6
|  |  | -16.05 | 42.469 | 1 | -0.032 | 25 | 26.968 | 0.736 | -26.968 | 73.6
|  |  | -16 | 42.353 | 1 | -0.032 | 25 | 26.958 | 0.735 | -26.958 | 73.5
|  |  | -15.95 | 42.238 | 1 | -0.032 | 25 | 26.948 | 0.735 | -26.948 | 73.5
|  |  | -15.9 | 42.123 | 1 | -0.032 | 25 | 26.938 | 0.735 | -26.938 | 73.5
|  |  | -15.85 | 42.008 | 1 | -0.032 | 25 | 26.928 | 0.734 | -26.928 | 73.4
|  |  | -15.8 | 41.893 | 1 | -0.032 | 25 | 26.918 | 0.734 | -26.918 | 73.4
|  |  | -15.75 | 41.779 | 1 | -0.032 | 25 | 26.908 | 0.734 | -26.908 | 73.4
|  |  | -15.7 | 41.664 | 1 | -0.032 | 25 | 26.898 | 0.733 | -26.898 | 73.3
|  |  | -15.65 | 41.549 | 1 | -0.032 | 25 | 26.888 | 0.733 | -26.888 | 73.3
|  |  | -15.6 | 41.434 | 1 | -0.032 | 25 | 26.878 | 0.733 | -26.878 | 73.3
|  |  | -15.55 | 41.32 | 1 | -0.032 | 25 | 26.868 | 0.732 | -26.868 | 73.2
|  |  | -15.5 | 41.206 | 1 | -0.032 | 25 | 26.858 | 0.732 | -26.858 | 73.2
|  |  | -15.45 | 41.092 | 1 | -0.032 | 25 | 26.848 | 0.732 | -26.848 | 73.2
|  |  | -15.4 | 40.978 | 1 | -0.032 | 25 | 26.838 | 0.731 | -26.838 | 73.1
|  |  | -15.35 | 40.864 | 1 | -0.032 | 25 | 26.828 | 0.731 | -26.828 | 73.1
|  |  | -15.3 | 40.75 | 1 | -0.032 | 25 | 26.818 | 0.731 | -26.818 | 73.1
|  |  | -15.25 | 40.637 | 1 | -0.032 | 25 | 26.808 | 0.73 | -26.808 | 73
|  |  | -15.2 | 40.524 | 1 | -0.032 | 25 | 26.798 | 0.73 | -26.798 | 73
|  |  | -15.15 | 40.411 | 1 | -0.032 | 25 | 26.788 | 0.729 | -26.788 | 72.9
|  |  | -15.1 | 40.298 | 1 | -0.032 | 25 | 26.778 | 0.729 | -26.778 | 72.9
|  |  | -15.05 | 40.186 | 1 | -0.032 | 25 | 26.768 | 0.729 | -26.768 | 72.9
|  |  | -15 | 40.073 | 1 | -0.032 | 25 | 26.758 | 0.728 | -26.758 | 72.8
|  |  | -14.95 | 39.961 | 1 | -0.032 | 25 | 26.748 | 0.728 | -26.748 | 72.8
|  |  | -14.9 | 39.851 | 1 | -0.032 | 25 | 26.738 | 0.728 | -26.738 | 72.8
|  |  | -14.85 | 39.741 | 1 | -0.032 | 25 | 26.728 | 0.728 | -26.728 | 72.8
|  |  | -14.8 | 39.631 | 1 | -0.032 | 25 | 26.718 | 0.727 | -26.718 | 72.7
|  |  | -14.75 | 39.521 | 1 | -0.032 | 25 | 26.708 | 0.727 | -26.708 | 72.7
|  |  | -14.7 | 39.412 | 1 | -0.032 | 25 | 26.698 | 0.727 | -26.698 | 72.7
|  |  | -14.65 | 39.303 | 1 | -0.032 | 25 | 26.688 | 0.726 | -26.688 | 72.6
|  |  | -14.6 | 39.193 | 1 | -0.032 | 25 | 26.678 | 0.726 | -26.678 | 72.6
|  |  | -14.55 | 39.084 | 1 | -0.032 | 25 | 26.668 | 0.726 | -26.668 | 72.6
|  |  | -14.5 | 38.976 | 1 | -0.032 | 25 | 26.658 | 0.726 | -26.658 | 72.6
|  |  | -14.45 | 38.867 | 1 | -0.032 | 25 | 26.648 | 0.725 | -26.648 | 72.5
|  |  | -14.4 | 38.759 | 1 | -0.032 | 25 | 26.638 | 0.725 | -26.638 | 72.5
|  |  | -14.35 | 38.65 | 1 | -0.032 | 25 | 26.628 | 0.725 | -26.628 | 72.5
|  |  | -14.3 | 38.542 | 1 | -0.032 | 25 | 26.618 | 0.724 | -26.618 | 72.4
|  |  | -14.25 | 38.435 | 1 | -0.032 | 25 | 26.608 | 0.724 | -26.608 | 72.4
|  |  | -14.2 | 38.327 | 1 | -0.032 | 25 | 26.598 | 0.724 | -26.598 | 72.4
|  |  | -14.15 | 38.219 | 1 | -0.032 | 25 | 26.588 | 0.723 | -26.588 | 72.3
|  |  | -14.1 | 38.115 | 1 | -0.032 | 25 | 26.578 | 0.723 | -26.578 | 72.3
|  |  | -14.05 | 38.01 | 1 | -0.032 | 25 | 26.568 | 0.723 | -26.568 | 72.3
|  |  | -14 | 37.906 | 1 | -0.032 | 25 | 26.558 | 0.723 | -26.558 | 72.3
|  |  | -13.95 | 37.802 | 1 | -0.032 | 25 | 26.548 | 0.722 | -26.548 | 72.2
|  |  | -13.9 | 37.698 | 1 | -0.032 | 25 | 26.538 | 0.722 | -26.538 | 72.2
|  |  | -13.85 | 37.594 | 1 | -0.033 | 25 | 26.528 | 0.722 | -26.528 | 72.2
|  |  | -13.8 | 37.491 | 1 | -0.033 | 25 | 26.519 | 0.721 | -26.519 | 72.1
|  |  | -13.75 | 37.387 | 1 | -0.033 | 25 | 26.509 | 0.721 | -26.509 | 72.1
|  |  | -13.7 | 37.284 | 1 | -0.033 | 25 | 26.499 | 0.721 | -26.499 | 72.1
|  |  | -13.65 | 37.181 | 1 | -0.033 | 25 | 26.489 | 0.721 | -26.489 | 72.1
|  |  | -13.6 | 37.078 | 1 | -0.033 | 25 | 26.479 | 0.72 | -26.479 | 72
|  |  | -13.55 | 36.975 | 1 | -0.033 | 25 | 26.469 | 0.72 | -26.469 | 72
|  |  | -13.5 | 36.873 | 1 | -0.033 | 25 | 26.459 | 0.72 | -26.459 | 72
|  |  | -13.45 | 36.77 | 1 | -0.033 | 25 | 26.449 | 0.719 | -26.449 | 71.9
|  |  | -13.4 | 36.668 | 1 | -0.033 | 25 | 26.439 | 0.719 | -26.439 | 71.9
|  |  | -13.35 | 36.565 | 1 | -0.033 | 25 | 26.429 | 0.719 | -26.429 | 71.9
|  |  | -13.3 | 36.465 | 1 | -0.033 | 25 | 26.419 | 0.719 | -26.419 | 71.9
|  |  | -13.25 | 36.364 | 1 | -0.033 | 25 | 26.409 | 0.718 | -26.409 | 71.8
|  |  | -13.2 | 36.263 | 1 | -0.033 | 25 | 26.399 | 0.718 | -26.399 | 71.8
|  |  | -13.15 | 36.163 | 1 | -0.033 | 25 | 26.389 | 0.718 | -26.389 | 71.8
|  |  | -13.1 | 36.063 | 1 | -0.033 | 25 | 26.379 | 0.717 | -26.379 | 71.7
|  |  | -13.05 | 35.963 | 1 | -0.033 | 25 | 26.369 | 0.717 | -26.369 | 71.7
|  |  | -13 | 35.863 | 1 | -0.033 | 25 | 26.359 | 0.717 | -26.359 | 71.7
|  |  | -12.95 | 35.763 | 1 | -0.033 | 25 | 26.349 | 0.717 | -26.349 | 71.7
|  |  | -12.9 | 35.663 | 1 | -0.033 | 25 | 26.339 | 0.716 | -26.339 | 71.6
|  |  | -12.85 | 35.564 | 1 | -0.033 | 25 | 26.329 | 0.716 | -26.329 | 71.6
|  |  | -12.8 | 35.464 | 1 | -0.033 | 25 | 26.319 | 0.716 | -26.319 | 71.6
|  |  | -12.75 | 35.365 | 1 | -0.033 | 25 | 26.309 | 0.715 | -26.309 | 71.5
|  |  | -12.7 | 35.266 | 1 | -0.033 | 25 | 26.299 | 0.715 | -26.299 | 71.5
|  |  | -12.65 | 35.167 | 1 | -0.033 | 25 | 26.289 | 0.715 | -26.289 | 71.5
|  |  | -12.6 | 35.068 | 1 | -0.033 | 25 | 26.279 | 0.714 | -26.279 | 71.4
|  |  | -12.55 | 34.97 | 1 | -0.033 | 25 | 26.269 | 0.714 | -26.269 | 71.4
|  |  | -12.5 | 34.876 | 1 | -0.033 | 25 | 26.259 | 0.714 | -26.259 | 71.4
|  |  | -12.45 | 34.782 | 1 | -0.033 | 25 | 26.249 | 0.713 | -26.249 | 71.3
|  |  | -12.4 | 34.688 | 1 | -0.033 | 25 | 26.239 | 0.713 | -26.239 | 71.3
|  |  | -12.35 | 34.594 | 1 | -0.033 | 25 | 26.229 | 0.713 | -26.229 | 71.3
|  |  | -12.3 | 34.5 | 1 | -0.033 | 25 | 26.219 | 0.713 | -26.219 | 71.3
|  |  | -12.25 | 34.407 | 1 | -0.033 | 25 | 26.209 | 0.712 | -26.209 | 71.2
|  |  | -12.2 | 34.313 | 1 | -0.033 | 25 | 26.199 | 0.712 | -26.199 | 71.2
|  |  | -12.15 | 34.22 | 1 | -0.033 | 25 | 26.189 | 0.712 | -26.189 | 71.2
|  |  | -12.1 | 34.127 | 1 | -0.033 | 25 | 26.179 | 0.711 | -26.179 | 71.1
|  |  | -12.05 | 34.034 | 1 | -0.033 | 25 | 26.169 | 0.711 | -26.169 | 71.1
|  |  | -12 | 33.941 | 1 | -0.033 | 25 | 26.159 | 0.711 | -26.159 | 71.1
|  |  | -11.95 | 33.848 | 1 | -0.033 | 25 | 26.149 | 0.71 | -26.149 | 71
|  |  | -11.9 | 33.756 | 1 | -0.033 | 25 | 26.139 | 0.71 | -26.139 | 71
|  |  | -11.85 | 33.664 | 1 | -0.033 | 25 | 26.129 | 0.71 | -26.129 | 71
|  |  | -11.8 | 33.572 | 1 | -0.033 | 25 | 26.119 | 0.709 | -26.119 | 70.9
|  |  | -11.75 | 33.479 | 1 | -0.033 | 25 | 26.109 | 0.709 | -26.109 | 70.9
|  |  | -11.7 | 33.387 | 1 | -0.033 | 25 | 26.099 | 0.709 | -26.099 | 70.9
|  |  | -11.65 | 33.295 | 1 | -0.033 | 25 | 26.089 | 0.708 | -26.089 | 70.8
|  |  | -11.6 | 33.203 | 1 | -0.033 | 25 | 26.079 | 0.708 | -26.079 | 70.8
|  |  | -11.55 | 33.111 | 1 | -0.033 | 25 | 26.069 | 0.708 | -26.069 | 70.8
|  |  | -11.5 | 33.02 | 1 | -0.033 | 25 | 26.059 | 0.708 | -26.059 | 70.8
|  |  | -11.45 | 32.929 | 1 | -0.033 | 25 | 26.049 | 0.707 | -26.049 | 70.7
|  |  | -11.4 | 32.837 | 1 | -0.033 | 25 | 26.039 | 0.707 | -26.039 | 70.7
|  |  | -11.35 | 32.746 | 1 | -0.033 | 25 | 26.029 | 0.707 | -26.029 | 70.7
|  |  | -11.3 | 32.655 | 1 | -0.033 | 25 | 26.019 | 0.706 | -26.019 | 70.6
|  |  | -11.25 | 32.564 | 1 | -0.033 | 25 | 26.009 | 0.706 | -26.009 | 70.6
|  |  | -11.2 | 32.473 | 1 | -0.033 | 25 | 25.999 | 0.706 | -25.999 | 70.6
|  |  | -11.15 | 32.383 | 1 | -0.033 | 25 | 25.989 | 0.705 | -25.989 | 70.5
|  |  | -11.1 | 32.292 | 1 | -0.033 | 25 | 25.979 | 0.705 | -25.979 | 70.5
|  |  | -11.05 | 32.202 | 1 | -0.033 | 25 | 25.969 | 0.705 | -25.969 | 70.5
|  |  | -11 | 32.112 | 1 | -0.033 | 25 | 25.959 | 0.704 | -25.959 | 70.4
|  |  | -10.95 | 32.021 | 1 | -0.033 | 25 | 25.949 | 0.704 | -25.949 | 70.4
|  |  | -10.9 | 31.934 | 1 | -0.033 | 25 | 25.939 | 0.704 | -25.939 | 70.4
|  |  | -10.85 | 31.847 | 1 | -0.033 | 25 | 25.929 | 0.703 | -25.929 | 70.3
|  |  | -10.8 | 31.76 | 1 | -0.033 | 25 | 25.919 | 0.703 | -25.919 | 70.3
|  |  | -10.75 | 31.673 | 1 | -0.033 | 25 | 25.909 | 0.703 | -25.909 | 70.3
|  |  | -10.7 | 31.586 | 1 | -0.033 | 25 | 25.899 | 0.702 | -25.899 | 70.2
|  |  | -10.65 | 31.499 | 1 | -0.033 | 25 | 25.889 | 0.702 | -25.889 | 70.2
|  |  | -10.6 | 31.413 | 1 | -0.033 | 25 | 25.879 | 0.702 | -25.879 | 70.2
|  |  | -10.55 | 31.326 | 1 | -0.033 | 25 | 25.869 | 0.701 | -25.869 | 70.1
|  |  | -10.5 | 31.24 | 1 | -0.033 | 25 | 25.859 | 0.701 | -25.859 | 70.1
|  |  | -10.45 | 31.154 | 1 | -0.033 | 25 | 25.849 | 0.701 | -25.849 | 70.1
|  |  | -10.4 | 31.068 | 1 | -0.033 | 25 | 25.839 | 0.7 | -25.839 | 70
|  |  | -10.35 | 30.982 | 1 | -0.034 | 25 | 25.829 | 0.7 | -25.829 | 70
|  |  | -10.3 | 30.896 | 1 | -0.034 | 25 | 25.819 | 0.7 | -25.819 | 70
|  |  | -10.25 | 30.811 | 1 | -0.034 | 25 | 25.809 | 0.7 | -25.809 | 70
|  |  | -10.2 | 30.725 | 1 | -0.034 | 25 | 25.799 | 0.699 | -25.799 | 69.9
|  |  | -10.15 | 30.64 | 1 | -0.034 | 25 | 25.789 | 0.699 | -25.789 | 69.9
|  |  | -10.1 | 30.556 | 1 | -0.034 | 25 | 25.779 | 0.699 | -25.779 | 69.9
|  |  | -10.05 | 30.473 | 1 | -0.034 | 25 | 25.769 | 0.698 | -25.769 | 69.8
|  |  | -10 | 30.39 | 1 | -0.034 | 25 | 25.759 | 0.698 | -25.759 | 69.8
|  |  | -9.95 | 30.306 | 1 | -0.034 | 25 | 25.749 | 0.698 | -25.749 | 69.8
|  |  | -9.9 | 30.224 | 1 | -0.034 | 25 | 25.739 | 0.697 | -25.739 | 69.7
|  |  | -9.85 | 30.141 | 1 | -0.034 | 25 | 25.729 | 0.697 | -25.729 | 69.7
|  |  | -9.8 | 30.058 | 1 | -0.034 | 25 | 25.719 | 0.697 | -25.719 | 69.7
|  |  | -9.75 | 29.976 | 1 | -0.034 | 25 | 25.709 | 0.696 | -25.709 | 69.6
|  |  | -9.7 | 29.893 | 1 | -0.034 | 25 | 25.699 | 0.696 | -25.699 | 69.6
|  |  | -9.65 | 29.811 | 1 | -0.034 | 25 | 25.689 | 0.696 | -25.689 | 69.6
|  |  | -9.6 | 29.729 | 1 | -0.034 | 25 | 25.679 | 0.695 | -25.679 | 69.5
|  |  | -9.55 | 29.647 | 1 | -0.034 | 25 | 25.669 | 0.695 | -25.669 | 69.5
|  |  | -9.5 | 29.565 | 1 | -0.034 | 25 | 25.659 | 0.695 | -25.659 | 69.5
|  |  | -9.45 | 29.483 | 1 | -0.034 | 25 | 25.649 | 0.694 | -25.649 | 69.4
|  |  | -9.4 | 29.401 | 1 | -0.034 | 25 | 25.639 | 0.694 | -25.639 | 69.4
|  |  | -9.35 | 29.32 | 1 | -0.034 | 25 | 25.629 | 0.694 | -25.629 | 69.4
|  |  | -9.3 | 29.239 | 1 | -0.034 | 25 | 25.619 | 0.694 | -25.619 | 69.4
|  |  | -9.25 | 29.159 | 1 | -0.034 | 25 | 25.609 | 0.693 | -25.609 | 69.3
|  |  | -9.2 | 29.078 | 1 | -0.034 | 25 | 25.599 | 0.693 | -25.599 | 69.3
|  |  | -9.15 | 28.998 | 1 | -0.034 | 25 | 25.589 | 0.693 | -25.589 | 69.3
|  |  | -9.1 | 28.918 | 1 | -0.034 | 25 | 25.579 | 0.692 | -25.579 | 69.2
|  |  | -9.05 | 28.838 | 1 | -0.034 | 25 | 25.569 | 0.692 | -25.569 | 69.2
|  |  | -9 | 28.758 | 1 | -0.034 | 25 | 25.559 | 0.692 | -25.559 | 69.2
|  |  | -8.95 | 28.678 | 1 | -0.034 | 25 | 25.549 | 0.691 | -25.549 | 69.1
|  |  | -8.9 | 28.599 | 1 | -0.034 | 25 | 25.539 | 0.691 | -25.539 | 69.1
|  |  | -8.85 | 28.52 | 1 | -0.034 | 25 | 25.529 | 0.691 | -25.529 | 69.1
|  |  | -8.8 | 28.44 | 1 | -0.034 | 25 | 25.519 | 0.69 | -25.519 | 69
|  |  | -8.75 | 28.361 | 1 | -0.034 | 25 | 25.509 | 0.69 | -25.509 | 69
|  |  | -8.7 | 28.282 | 1 | -0.034 | 25 | 25.499 | 0.69 | -25.499 | 69
|  |  | -8.65 | 28.203 | 1 | -0.034 | 25 | 25.489 | 0.689 | -25.489 | 68.9
|  |  | -8.6 | 28.124 | 1 | -0.034 | 25 | 25.479 | 0.689 | -25.479 | 68.9
|  |  | -8.55 | 28.045 | 1 | -0.034 | 25 | 25.469 | 0.689 | -25.469 | 68.9
|  |  | -8.5 | 27.969 | 1 | -0.034 | 25 | 25.459 | 0.688 | -25.459 | 68.8
|  |  | -8.45 | 27.893 | 1 | -0.034 | 25 | 25.449 | 0.688 | -25.449 | 68.8
|  |  | -8.4 | 27.817 | 1 | -0.034 | 25 | 25.439 | 0.688 | -25.439 | 68.8
|  |  | -8.35 | 27.741 | 1 | -0.034 | 25 | 25.429 | 0.688 | -25.429 | 68.8
|  |  | -8.3 | 27.666 | 1 | -0.034 | 25 | 25.419 | 0.687 | -25.419 | 68.7
|  |  | -8.25 | 27.59 | 1 | -0.034 | 25 | 25.409 | 0.687 | -25.409 | 68.7
|  |  | -8.2 | 27.515 | 1 | -0.034 | 25 | 25.399 | 0.687 | -25.399 | 68.7
|  |  | -8.15 | 27.439 | 1 | -0.034 | 25 | 25.389 | 0.686 | -25.389 | 68.6
|  |  | -8.1 | 27.364 | 1 | -0.034 | 25 | 25.379 | 0.686 | -25.379 | 68.6
|  |  | -8.05 | 27.289 | 1 | -0.034 | 25 | 25.369 | 0.686 | -25.369 | 68.6
|  |  | -8 | 27.214 | 1 | -0.034 | 25 | 25.359 | 0.685 | -25.359 | 68.5
|  |  | -7.95 | 27.139 | 1 | -0.034 | 25 | 25.349 | 0.685 | -25.349 | 68.5
|  |  | -7.9 | 27.065 | 1 | -0.034 | 25 | 25.339 | 0.685 | -25.339 | 68.5
|  |  | -7.85 | 26.99 | 1 | -0.034 | 25 | 25.329 | 0.684 | -25.329 | 68.4
|  |  | -7.8 | 26.916 | 1 | -0.034 | 25 | 25.319 | 0.684 | -25.319 | 68.4
|  |  | -7.75 | 26.841 | 1 | -0.034 | 25 | 25.309 | 0.684 | -25.309 | 68.4
|  |  | -7.7 | 26.767 | 1 | -0.034 | 25 | 25.299 | 0.683 | -25.299 | 68.3
|  |  | -7.65 | 26.692 | 1 | -0.034 | 25 | 25.289 | 0.683 | -25.289 | 68.3
|  |  | -7.6 | 26.618 | 1 | -0.034 | 25 | 25.279 | 0.683 | -25.279 | 68.3
|  |  | -7.55 | 26.543 | 1 | -0.034 | 25 | 25.269 | 0.682 | -25.269 | 68.2
|  |  | -7.5 | 26.469 | 1 | -0.034 | 25 | 25.259 | 0.682 | -25.259 | 68.2
|  |  | -7.45 | 26.395 | 1 | -0.034 | 25 | 25.249 | 0.682 | -25.249 | 68.2
|  |  | -7.4 | 26.321 | 1 | -0.034 | 25 | 25.239 | 0.681 | -25.239 | 68.1
|  |  | -7.35 | 26.247 | 1 | -0.034 | 25 | 25.229 | 0.681 | -25.229 | 68.1
|  |  | -7.3 | 26.174 | 1 | -0.034 | 25 | 25.219 | 0.681 | -25.219 | 68.1
|  |  | -7.25 | 26.1 | 1 | -0.034 | 25 | 25.209 | 0.68 | -25.209 | 68
|  |  | -7.2 | 26.027 | 1 | -0.034 | 25 | 25.199 | 0.68 | -25.199 | 68
|  |  | -7.15 | 25.953 | 1 | -0.034 | 25 | 25.189 | 0.68 | -25.189 | 68
|  |  | -7.1 | 25.88 | 1 | -0.034 | 25 | 25.179 | 0.679 | -25.179 | 67.9
|  |  | -7.05 | 25.807 | 1 | -0.034 | 25 | 25.169 | 0.679 | -25.169 | 67.9
|  |  | -7 | 25.734 | 1 | -0.034 | 25 | 25.159 | 0.679 | -25.159 | 67.9
|  |  | -6.95 | 25.661 | 1 | -0.034 | 25 | 25.149 | 0.679 | -25.149 | 67.9
|  |  | -6.9 | 25.591 | 1 | -0.034 | 25 | 25.139 | 0.678 | -25.139 | 67.8
|  |  | -6.85 | 25.52 | 1 | -0.035 | 25 | 25.129 | 0.678 | -25.129 | 67.8
|  |  | -6.8 | 25.45 | 1 | -0.035 | 25 | 25.119 | 0.678 | -25.119 | 67.8
|  |  | -6.75 | 25.379 | 1 | -0.035 | 25 | 25.109 | 0.677 | -25.109 | 67.7
|  |  | -6.7 | 25.309 | 1 | -0.035 | 25 | 25.099 | 0.677 | -25.099 | 67.7
|  |  | -6.65 | 25.239 | 1 | -0.035 | 25 | 25.089 | 0.677 | -25.089 | 67.7
|  |  | -6.6 | 25.169 | 1 | -0.035 | 25 | 25.079 | 0.676 | -25.079 | 67.6
|  |  | -6.55 | 25.099 | 1 | -0.035 | 25 | 25.069 | 0.676 | -25.069 | 67.6
|  |  | -6.5 | 25.029 | 1 | -0.035 | 25 | 25.059 | 0.676 | -25.059 | 67.6
|  |  | -6.45 | 24.96 | 1 | -0.035 | 25 | 25.049 | 0.675 | -25.049 | 67.5
|  |  | -6.4 | 24.89 | 1 | -0.035 | 25 | 25.039 | 0.675 | -25.039 | 67.5
|  |  | -6.35 | 24.82 | 1 | -0.035 | 25 | 25.029 | 0.675 | -25.029 | 67.5
|  |  | -6.3 | 24.751 | 1 | -0.035 | 25 | 25.019 | 0.674 | -25.019 | 67.4
|  |  | -6.25 | 24.682 | 1 | -0.035 | 25 | 25.009 | 0.674 | -25.009 | 67.4
|  |  | -6.2 | 24.613 | 1 | -0.035 | 25 | 24.999 | 0.674 | -24.999 | 67.4
|  |  | -6.15 | 24.544 | 1 | -0.035 | 25 | 24.989 | 0.673 | -24.989 | 67.3
|  |  | -6.1 | 24.477 | 1 | -0.035 | 25 | 24.979 | 0.673 | -24.979 | 67.3
|  |  | -6.05 | 24.41 | 1 | -0.035 | 25 | 24.969 | 0.673 | -24.969 | 67.3
|  |  | -6 | 24.343 | 1 | -0.035 | 25 | 24.959 | 0.672 | -24.959 | 67.2
|  |  | -5.95 | 24.276 | 1 | -0.035 | 25 | 24.949 | 0.672 | -24.949 | 67.2
|  |  | -5.9 | 24.209 | 1 | -0.035 | 25 | 24.939 | 0.672 | -24.939 | 67.2
|  |  | -5.85 | 24.142 | 1 | -0.035 | 25 | 24.929 | 0.671 | -24.929 | 67.1
|  |  | -5.8 | 24.076 | 1 | -0.035 | 25 | 24.919 | 0.671 | -24.919 | 67.1
|  |  | -5.75 | 24.009 | 1 | -0.035 | 25 | 24.909 | 0.671 | -24.909 | 67.1
|  |  | -5.7 | 23.943 | 1 | -0.035 | 25 | 24.899 | 0.67 | -24.899 | 67
|  |  | -5.65 | 23.877 | 1 | -0.035 | 25 | 24.889 | 0.67 | -24.889 | 67
|  |  | -5.6 | 23.811 | 1 | -0.035 | 25 | 24.879 | 0.67 | -24.879 | 67
|  |  | -5.55 | 23.744 | 1 | -0.035 | 25 | 24.869 | 0.669 | -24.869 | 66.9
|  |  | -5.5 | 23.679 | 1 | -0.035 | 25 | 24.859 | 0.669 | -24.859 | 66.9
|  |  | -5.45 | 23.613 | 1 | -0.035 | 25 | 24.849 | 0.669 | -24.849 | 66.9
|  |  | -5.4 | 23.547 | 1 | -0.035 | 25 | 24.839 | 0.669 | -24.839 | 66.9
|  |  | -5.35 | 23.481 | 1 | -0.035 | 25 | 24.829 | 0.668 | -24.829 | 66.8
|  |  | -5.3 | 23.417 | 1 | -0.035 | 25 | 24.819 | 0.668 | -24.819 | 66.8
|  |  | -5.25 | 23.352 | 1 | -0.035 | 25 | 24.809 | 0.668 | -24.809 | 66.8
|  |  | -5.2 | 23.288 | 1 | -0.035 | 25 | 24.799 | 0.667 | -24.799 | 66.7
|  |  | -5.15 | 23.224 | 1 | -0.035 | 25 | 24.789 | 0.667 | -24.789 | 66.7
|  |  | -5.1 | 23.16 | 1 | -0.035 | 25 | 24.779 | 0.667 | -24.779 | 66.7
|  |  | -5.05 | 23.096 | 1 | -0.035 | 25 | 24.769 | 0.666 | -24.769 | 66.6
|  |  | -5 | 23.032 | 1 | -0.035 | 25 | 24.759 | 0.666 | -24.759 | 66.6
|  |  | -4.95 | 22.968 | 1 | -0.035 | 25 | 24.749 | 0.666 | -24.749 | 66.6
|  |  | -4.9 | 22.904 | 1 | -0.035 | 25 | 24.739 | 0.665 | -24.739 | 66.5
|  |  | -4.85 | 22.841 | 1 | -0.035 | 25 | 24.73 | 0.665 | -24.73 | 66.5
|  |  | -4.8 | 22.777 | 1 | -0.035 | 25 | 24.72 | 0.665 | -24.72 | 66.5
|  |  | -4.75 | 22.713 | 1 | -0.035 | 25 | 24.71 | 0.664 | -24.71 | 66.4
|  |  | -4.7 | 22.65 | 1 | -0.035 | 25 | 24.7 | 0.664 | -24.7 | 66.4
|  |  | -4.65 | 22.587 | 1 | -0.035 | 25 | 24.69 | 0.664 | -24.69 | 66.4
|  |  | -4.6 | 22.524 | 1 | -0.035 | 25 | 24.68 | 0.663 | -24.68 | 66.3
|  |  | -4.55 | 22.46 | 1 | -0.035 | 25 | 24.67 | 0.663 | -24.67 | 66.3
|  |  | -4.5 | 22.4 | 1 | -0.035 | 25 | 24.66 | 0.663 | -24.66 | 66.3
|  |  | -4.45 | 22.34 | 1 | -0.035 | 25 | 24.65 | 0.662 | -24.65 | 66.2
|  |  | -4.4 | 22.28 | 1 | -0.035 | 25 | 24.64 | 0.662 | -24.64 | 66.2
|  |  | -4.35 | 22.219 | 1 | -0.035 | 25 | 24.63 | 0.662 | -24.63 | 66.2
|  |  | -4.3 | 22.159 | 1 | -0.035 | 25 | 24.62 | 0.661 | -24.62 | 66.1
|  |  | -4.25 | 22.1 | 1 | -0.035 | 25 | 24.61 | 0.661 | -24.61 | 66.1
|  |  | -4.2 | 22.04 | 1 | -0.035 | 25 | 24.6 | 0.661 | -24.6 | 66.1
|  |  | -4.15 | 21.98 | 1 | -0.035 | 25 | 24.59 | 0.66 | -24.59 | 66
|  |  | -4.1 | 21.92 | 1 | -0.035 | 25 | 24.58 | 0.66 | -24.58 | 66
|  |  | -4.05 | 21.861 | 1 | -0.035 | 25 | 24.57 | 0.66 | -24.57 | 66
|  |  | -4 | 21.801 | 1 | -0.035 | 25 | 24.56 | 0.66 | -24.56 | 66
|  |  | -3.95 | 21.742 | 1 | -0.035 | 25 | 24.55 | 0.659 | -24.55 | 65.9
|  |  | -3.9 | 21.682 | 1 | -0.035 | 25 | 24.54 | 0.659 | -24.54 | 65.9
|  |  | -3.85 | 21.623 | 1 | -0.035 | 25 | 24.53 | 0.659 | -24.53 | 65.9
|  |  | -3.8 | 21.564 | 1 | -0.035 | 25 | 24.52 | 0.658 | -24.52 | 65.8
|  |  | -3.75 | 21.505 | 1 | -0.035 | 25 | 24.51 | 0.658 | -24.51 | 65.8
|  |  | -3.7 | 21.446 | 1 | -0.035 | 25 | 24.5 | 0.658 | -24.5 | 65.8
|  |  | -3.65 | 21.388 | 1 | -0.035 | 25 | 24.49 | 0.657 | -24.49 | 65.7
|  |  | -3.6 | 21.329 | 1 | -0.035 | 25 | 24.48 | 0.657 | -24.48 | 65.7
|  |  | -3.55 | 21.271 | 1 | -0.035 | 25 | 24.47 | 0.657 | -24.47 | 65.7
|  |  | -3.5 | 21.213 | 1 | -0.035 | 25 | 24.46 | 0.656 | -24.46 | 65.6
|  |  | -3.45 | 21.154 | 1 | -0.035 | 25 | 24.45 | 0.656 | -24.45 | 65.6
|  |  | -3.4 | 21.096 | 1 | -0.035 | 25 | 24.44 | 0.656 | -24.44 | 65.6
|  |  | -3.35 | 21.038 | 1 | -0.036 | 25 | 24.43 | 0.656 | -24.43 | 65.6
|  |  | -3.3 | 20.98 | 1 | -0.036 | 25 | 24.42 | 0.655 | -24.42 | 65.5
|  |  | -3.25 | 20.923 | 1 | -0.036 | 25 | 24.41 | 0.655 | -24.41 | 65.5
|  |  | -3.2 | 20.865 | 1 | -0.036 | 25 | 24.4 | 0.655 | -24.4 | 65.5
|  |  | -3.15 | 20.807 | 1 | -0.036 | 25 | 24.39 | 0.654 | -24.39 | 65.4
|  |  | -3.1 | 20.749 | 1 | -0.036 | 25 | 24.38 | 0.654 | -24.38 | 65.4
|  |  | -3.05 | 20.692 | 1 | -0.036 | 25 | 24.37 | 0.654 | -24.37 | 65.4
|  |  | -3 | 20.634 | 1 | -0.036 | 25 | 24.36 | 0.653 | -24.36 | 65.3
|  |  | -2.95 | 20.577 | 1 | -0.036 | 25 | 24.35 | 0.653 | -24.35 | 65.3
|  |  | -2.9 | 20.522 | 1 | -0.036 | 25 | 24.34 | 0.653 | -24.34 | 65.3
|  |  | -2.85 | 20.466 | 1 | -0.036 | 25 | 24.33 | 0.652 | -24.33 | 65.2
|  |  | -2.8 | 20.411 | 1 | -0.036 | 25 | 24.32 | 0.652 | -24.32 | 65.2
|  |  | -2.75 | 20.356 | 1 | -0.036 | 25 | 24.31 | 0.652 | -24.31 | 65.2
|  |  | -2.7 | 20.301 | 1 | -0.036 | 25 | 24.3 | 0.652 | -24.3 | 65.2
|  |  | -2.65 | 20.246 | 1 | -0.036 | 25 | 24.29 | 0.651 | -24.29 | 65.1
|  |  | -2.6 | 20.191 | 1 | -0.036 | 25 | 24.28 | 0.651 | -24.28 | 65.1
|  |  | -2.55 | 20.136 | 1 | -0.036 | 25 | 24.27 | 0.651 | -24.27 | 65.1
|  |  | -2.5 | 20.081 | 1 | -0.036 | 25 | 24.26 | 0.65 | -24.26 | 65
|  |  | -2.45 | 20.026 | 1 | -0.036 | 25 | 24.25 | 0.65 | -24.25 | 65
|  |  | -2.4 | 19.972 | 1 | -0.036 | 25 | 24.24 | 0.65 | -24.24 | 65
|  |  | -2.35 | 19.917 | 1 | -0.036 | 25 | 24.23 | 0.649 | -24.23 | 64.9
|  |  | -2.3 | 19.863 | 1 | -0.036 | 25 | 24.22 | 0.649 | -24.22 | 64.9
|  |  | -2.25 | 19.808 | 1 | -0.036 | 25 | 24.21 | 0.649 | -24.21 | 64.9
|  |  | -2.2 | 19.754 | 1 | -0.036 | 25 | 24.2 | 0.648 | -24.2 | 64.8
|  |  | -2.15 | 19.7 | 1 | -0.036 | 25 | 24.19 | 0.648 | -24.19 | 64.8
|  |  | -2.1 | 19.647 | 1 | -0.036 | 25 | 24.18 | 0.648 | -24.18 | 64.8
|  |  | -2.05 | 19.594 | 1 | -0.036 | 25 | 24.17 | 0.648 | -24.17 | 64.8
|  |  | -2 | 19.541 | 1 | -0.036 | 25 | 24.16 | 0.647 | -24.16 | 64.7
|  |  | -1.95 | 19.488 | 1 | -0.036 | 25 | 24.15 | 0.647 | -24.15 | 64.7
|  |  | -1.9 | 19.436 | 1 | -0.036 | 25 | 24.14 | 0.647 | -24.14 | 64.7
|  |  | -1.85 | 19.383 | 1 | -0.036 | 25 | 24.13 | 0.646 | -24.13 | 64.6
|  |  | -1.8 | 19.331 | 1 | -0.036 | 25 | 24.12 | 0.646 | -24.12 | 64.6
|  |  | -1.75 | 19.278 | 1 | -0.036 | 25 | 24.11 | 0.646 | -24.11 | 64.6
|  |  | -1.7 | 19.226 | 1 | -0.036 | 25 | 24.1 | 0.645 | -24.1 | 64.5
|  |  | -1.65 | 19.174 | 1 | -0.036 | 25 | 24.09 | 0.645 | -24.09 | 64.5
|  |  | -1.6 | 19.122 | 1 | -0.036 | 25 | 24.08 | 0.645 | -24.08 | 64.5
|  |  | -1.55 | 19.069 | 1 | -0.036 | 25 | 24.07 | 0.644 | -24.07 | 64.4
|  |  | -1.5 | 19.017 | 1 | -0.036 | 25 | 24.06 | 0.644 | -24.06 | 64.4
|  |  | -1.45 | 18.966 | 1 | -0.036 | 25 | 24.05 | 0.644 | -24.05 | 64.4
|  |  | -1.4 | 18.914 | 1 | -0.036 | 25 | 24.04 | 0.644 | -24.04 | 64.4
|  |  | -1.35 | 18.862 | 1 | -0.036 | 25 | 24.03 | 0.643 | -24.03 | 64.3
|  |  | -1.3 | 18.811 | 1 | -0.036 | 25 | 24.02 | 0.643 | -24.02 | 64.3
|  |  | -1.25 | 18.76 | 1 | -0.036 | 25 | 24.01 | 0.643 | -24.01 | 64.3
|  |  | -1.2 | 18.709 | 1 | -0.036 | 25 | 24 | 0.642 | -24 | 64.2
|  |  | -1.15 | 18.658 | 1 | -0.036 | 25 | 23.99 | 0.642 | -23.99 | 64.2
|  |  | -1.1 | 18.607 | 1 | -0.036 | 25 | 23.98 | 0.642 | -23.98 | 64.2
|  |  | -1.05 | 18.556 | 1 | -0.036 | 25 | 23.97 | 0.641 | -23.97 | 64.1
|  |  | -1 | 18.505 | 1 | -0.036 | 25 | 23.96 | 0.641 | -23.96 | 64.1
|  |  | -0.95 | 18.454 | 1 | -0.036 | 25 | 23.95 | 0.641 | -23.95 | 64.1
|  |  | -0.9 | 18.404 | 1 | -0.036 | 25 | 23.94 | 0.641 | -23.94 | 64.1
|  |  | -0.85 | 18.354 | 1 | -0.036 | 25 | 23.93 | 0.64 | -23.93 | 64
|  |  | -0.8 | 18.303 | 1 | -0.036 | 25 | 23.92 | 0.64 | -23.92 | 64
|  |  | -0.75 | 18.253 | 1 | -0.036 | 25 | 23.91 | 0.64 | -23.91 | 64
|  |  | -0.7 | 18.202 | 1 | -0.036 | 25 | 23.9 | 0.639 | -23.9 | 63.9
|  |  | -0.65 | 18.152 | 1 | -0.036 | 25 | 23.89 | 0.639 | -23.89 | 63.9
|  |  | -0.6 | 18.102 | 1 | -0.036 | 25 | 23.88 | 0.639 | -23.88 | 63.9
|  |  | -0.55 | 18.052 | 1 | -0.036 | 25 | 23.87 | 0.638 | -23.87 | 63.8
|  |  | -0.5 | 18.004 | 1 | -0.036 | 25 | 23.86 | 0.638 | -23.86 | 63.8
|  |  | -0.45 | 17.955 | 1 | -0.036 | 25 | 23.85 | 0.638 | -23.85 | 63.8
|  |  | -0.4 | 17.907 | 1 | -0.036 | 25 | 23.84 | 0.637 | -23.84 | 63.7
|  |  | -0.35 | 17.859 | 1 | -0.036 | 25 | 23.83 | 0.637 | -23.83 | 63.7
|  |  | -0.3 | 17.811 | 1 | -0.036 | 25 | 23.82 | 0.637 | -23.82 | 63.7
|  |  | -0.25 | 17.763 | 1 | -0.036 | 25 | 23.81 | 0.637 | -23.81 | 63.7
|  |  | -0.2 | 17.715 | 1 | -0.036 | 25 | 23.8 | 0.636 | -23.8 | 63.6
|  |  | -0.15 | 17.667 | 1 | -0.036 | 25 | 23.79 | 0.636 | -23.79 | 63.6
|  |  | -0.1 | 17.619 | 1 | -0.036 | 25 | 23.78 | 0.636 | -23.78 | 63.6
|  |  | -0.05 | 17.571 | 1 | -0.036 | 25 | 23.77 | 0.635 | -23.77 | 63.5
|  |  | 0 | 17.523 | 1 | -0.036 | 25 | 23.76 | 0.635 | -23.76 | 63.5
|  |  | 0.05 | 17.476 | 1 | -0.036 | 25 | 23.75 | 0.635 | -23.75 | 63.5
|  |  | 0.1 | 17.428 | 1 | -0.036 | 25 | 23.74 | 0.634 | -23.74 | 63.4
|  |  | 0.15 | 17.381 | 1 | -0.037 | 25 | 23.73 | 0.634 | -23.73 | 63.4
|  |  | 0.2 | 17.333 | 1 | -0.037 | 25 | 23.72 | 0.634 | -23.72 | 63.4
|  |  | 0.25 | 17.286 | 1 | -0.037 | 25 | 23.71 | 0.634 | -23.71 | 63.4
|  |  | 0.3 | 17.239 | 1 | -0.037 | 25 | 23.7 | 0.633 | -23.7 | 63.3
|  |  | 0.35 | 17.192 | 1 | -0.037 | 25 | 23.69 | 0.633 | -23.69 | 63.3
|  |  | 0.4 | 17.146 | 1 | -0.037 | 25 | 23.68 | 0.633 | -23.68 | 63.3
|  |  | 0.45 | 17.099 | 1 | -0.037 | 25 | 23.67 | 0.632 | -23.67 | 63.2
|  |  | 0.5 | 17.052 | 1 | -0.037 | 25 | 23.66 | 0.632 | -23.66 | 63.2
|  |  | 0.55 | 17.006 | 1 | -0.037 | 25 | 23.65 | 0.632 | -23.65 | 63.2
|  |  | 0.6 | 16.959 | 1 | -0.037 | 25 | 23.64 | 0.631 | -23.64 | 63.1
|  |  | 0.65 | 16.913 | 1 | -0.037 | 25 | 23.63 | 0.631 | -23.63 | 63.1
|  |  | 0.7 | 16.866 | 1 | -0.037 | 25 | 23.62 | 0.631 | -23.62 | 63.1
|  |  | 0.75 | 16.82 | 1 | -0.037 | 25 | 23.61 | 0.63 | -23.61 | 63
|  |  | 0.8 | 16.774 | 1 | -0.037 | 25 | 23.6 | 0.63 | -23.6 | 63
|  |  | 0.85 | 16.728 | 1 | -0.037 | 25 | 23.59 | 0.63 | -23.59 | 63
|  |  | 0.9 | 16.682 | 1 | -0.037 | 25 | 23.58 | 0.63 | -23.58 | 63
|  |  | 0.95 | 16.636 | 1 | -0.037 | 25 | 23.57 | 0.629 | -23.57 | 62.9
|  |  | 1 | 16.59 | 1 | -0.037 | 25 | 23.56 | 0.629 | -23.56 | 62.9
|  |  | 1.05 | 16.544 | 1 | -0.037 | 25 | 23.55 | 0.629 | -23.55 | 62.9
|  |  | 1.1 | 16.499 | 1 | -0.037 | 25 | 23.54 | 0.628 | -23.54 | 62.8
|  |  | 1.15 | 16.455 | 1 | -0.037 | 25 | 23.53 | 0.628 | -23.53 | 62.8
|  |  | 1.2 | 16.41 | 1 | -0.037 | 25 | 23.52 | 0.628 | -23.52 | 62.8
|  |  | 1.25 | 16.366 | 1 | -0.037 | 25 | 23.51 | 0.627 | -23.51 | 62.7
|  |  | 1.3 | 16.322 | 1 | -0.037 | 25 | 23.5 | 0.627 | -23.5 | 62.7
|  |  | 1.35 | 16.278 | 1 | -0.037 | 25 | 23.49 | 0.627 | -23.49 | 62.7
|  |  | 1.4 | 16.233 | 1 | -0.037 | 25 | 23.48 | 0.626 | -23.48 | 62.6
|  |  | 1.45 | 16.189 | 1 | -0.037 | 25 | 23.47 | 0.626 | -23.47 | 62.6
|  |  | 1.5 | 16.145 | 1 | -0.037 | 25 | 23.46 | 0.626 | -23.46 | 62.6
|  |  | 1.55 | 16.101 | 1 | -0.037 | 25 | 23.45 | 0.626 | -23.45 | 62.6
|  |  | 1.6 | 16.058 | 1 | -0.037 | 25 | 23.44 | 0.625 | -23.44 | 62.5
|  |  | 1.65 | 16.014 | 1 | -0.037 | 25 | 23.43 | 0.625 | -23.43 | 62.5
|  |  | 1.7 | 15.97 | 1 | -0.037 | 25 | 23.42 | 0.625 | -23.42 | 62.5
|  |  | 1.75 | 15.926 | 1 | -0.037 | 25 | 23.41 | 0.624 | -23.41 | 62.4
|  |  | 1.8 | 15.883 | 1 | -0.037 | 25 | 23.4 | 0.624 | -23.4 | 62.4
|  |  | 1.85 | 15.839 | 1 | -0.037 | 25 | 23.39 | 0.624 | -23.39 | 62.4
|  |  | 1.9 | 15.797 | 1 | -0.037 | 25 | 23.38 | 0.623 | -23.38 | 62.3
|  |  | 1.95 | 15.755 | 1 | -0.037 | 25 | 23.37 | 0.623 | -23.37 | 62.3
|  |  | 2 | 15.712 | 1 | -0.037 | 25 | 23.36 | 0.623 | -23.36 | 62.3
|  |  | 2.05 | 15.67 | 1 | -0.037 | 25 | 23.35 | 0.622 | -23.35 | 62.2
|  |  | 2.1 | 15.628 | 1 | -0.037 | 25 | 23.34 | 0.622 | -23.34 | 62.2
|  |  | 2.15 | 15.586 | 1 | -0.037 | 25 | 23.33 | 0.622 | -23.33 | 62.2
|  |  | 2.2 | 15.544 | 1 | -0.037 | 25 | 23.32 | 0.622 | -23.32 | 62.2
|  |  | 2.25 | 15.502 | 1 | -0.037 | 25 | 23.31 | 0.621 | -23.31 | 62.1
|  |  | 2.3 | 15.46 | 1 | -0.037 | 25 | 23.3 | 0.621 | -23.3 | 62.1
|  |  | 2.35 | 15.419 | 1 | -0.037 | 25 | 23.29 | 0.621 | -23.29 | 62.1
|  |  | 2.4 | 15.377 | 1 | -0.037 | 25 | 23.28 | 0.62 | -23.28 | 62
|  |  | 2.45 | 15.335 | 1 | -0.037 | 25 | 23.27 | 0.62 | -23.27 | 62
|  |  | 2.5 | 15.294 | 1 | -0.037 | 25 | 23.26 | 0.62 | -23.26 | 62
|  |  | 2.55 | 15.252 | 1 | -0.037 | 25 | 23.25 | 0.619 | -23.25 | 61.9
|  |  | 2.6 | 15.21 | 1 | -0.037 | 25 | 23.24 | 0.619 | -23.24 | 61.9
|  |  | 2.65 | 15.169 | 1 | -0.037 | 25 | 23.23 | 0.619 | -23.23 | 61.9
|  |  | 2.7 | 15.128 | 1 | -0.037 | 25 | 23.22 | 0.618 | -23.22 | 61.8
|  |  | 2.75 | 15.088 | 1 | -0.037 | 25 | 23.21 | 0.618 | -23.21 | 61.8
|  |  | 2.8 | 15.047 | 1 | -0.037 | 25 | 23.2 | 0.618 | -23.2 | 61.8
|  |  | 2.85 | 15.007 | 1 | -0.037 | 25 | 23.19 | 0.618 | -23.19 | 61.8
|  |  | 2.9 | 14.967 | 1 | -0.037 | 25 | 23.18 | 0.617 | -23.18 | 61.7
|  |  | 2.95 | 14.926 | 1 | -0.037 | 25 | 23.17 | 0.617 | -23.17 | 61.7
|  |  | 3 | 14.886 | 1 | -0.037 | 25 | 23.16 | 0.617 | -23.16 | 61.7
|  |  | 3.05 | 14.846 | 1 | -0.037 | 25 | 23.15 | 0.616 | -23.15 | 61.6
|  |  | 3.1 | 14.806 | 1 | -0.037 | 25 | 23.14 | 0.616 | -23.14 | 61.6
|  |  | 3.15 | 14.766 | 1 | -0.037 | 25 | 23.13 | 0.616 | -23.13 | 61.6
|  |  | 3.2 | 14.726 | 1 | -0.037 | 25 | 23.12 | 0.615 | -23.12 | 61.5
|  |  | 3.25 | 14.686 | 1 | -0.037 | 25 | 23.11 | 0.615 | -23.11 | 61.5
|  |  | 3.3 | 14.646 | 1 | -0.037 | 25 | 23.1 | 0.615 | -23.1 | 61.5
|  |  | 3.35 | 14.606 | 1 | -0.037 | 25 | 23.09 | 0.615 | -23.09 | 61.5
|  |  | 3.4 | 14.566 | 1 | -0.037 | 25 | 23.08 | 0.614 | -23.08 | 61.4
|  |  | 3.45 | 14.526 | 1 | -0.037 | 25 | 23.07 | 0.614 | -23.07 | 61.4
|  |  | 3.5 | 14.488 | 1 | -0.037 | 25 | 23.06 | 0.614 | -23.06 | 61.4
|  |  | 3.55 | 14.449 | 1 | -0.037 | 25 | 23.05 | 0.613 | -23.05 | 61.3
|  |  | 3.6 | 14.41 | 1 | -0.037 | 25 | 23.04 | 0.613 | -23.04 | 61.3
|  |  | 3.65 | 14.372 | 1 | -0.038 | 25 | 23.03 | 0.613 | -23.03 | 61.3
|  |  | 3.7 | 14.333 | 1 | -0.038 | 25 | 23.02 | 0.612 | -23.02 | 61.2
|  |  | 3.75 | 14.295 | 1 | -0.038 | 25 | 23.01 | 0.612 | -23.01 | 61.2
|  |  | 3.8 | 14.257 | 1 | -0.038 | 25 | 23 | 0.612 | -23 | 61.2
|  |  | 3.85 | 14.218 | 1 | -0.038 | 25 | 22.99 | 0.611 | -22.99 | 61.1
|  |  | 3.9 | 14.18 | 1 | -0.038 | 25 | 22.98 | 0.611 | -22.98 | 61.1
|  |  | 3.95 | 14.142 | 1 | -0.038 | 25 | 22.97 | 0.611 | -22.97 | 61.1
|  |  | 4 | 14.104 | 1 | -0.038 | 25 | 22.96 | 0.611 | -22.96 | 61.1
|  |  | 4.05 | 14.065 | 1 | -0.038 | 25 | 22.95 | 0.61 | -22.95 | 61
|  |  | 4.1 | 14.028 | 1 | -0.038 | 25 | 22.941 | 0.61 | -22.941 | 61
|  |  | 4.15 | 13.99 | 1 | -0.038 | 25 | 22.931 | 0.61 | -22.931 | 61
|  |  | 4.2 | 13.952 | 1 | -0.038 | 25 | 22.921 | 0.609 | -22.921 | 60.9
|  |  | 4.25 | 13.914 | 1 | -0.038 | 25 | 22.911 | 0.609 | -22.911 | 60.9
|  |  | 4.3 | 13.876 | 1 | -0.038 | 25 | 22.901 | 0.609 | -22.901 | 60.9
|  |  | 4.35 | 13.839 | 1 | -0.038 | 25 | 22.891 | 0.608 | -22.891 | 60.8
|  |  | 4.4 | 13.802 | 1 | -0.038 | 25 | 22.881 | 0.608 | -22.881 | 60.8
|  |  | 4.45 | 13.764 | 1 | -0.038 | 25 | 22.871 | 0.608 | -22.871 | 60.8
|  |  | 4.5 | 13.727 | 1 | -0.038 | 25 | 22.861 | 0.608 | -22.861 | 60.8
|  |  | 4.55 | 13.69 | 1 | -0.038 | 25 | 22.851 | 0.607 | -22.851 | 60.7
|  |  | 4.6 | 13.653 | 1 | -0.038 | 25 | 22.841 | 0.607 | -22.841 | 60.7
|  |  | 4.65 | 13.616 | 1 | -0.038 | 25 | 22.831 | 0.607 | -22.831 | 60.7
|  |  | 4.7 | 13.579 | 1 | -0.038 | 25 | 22.821 | 0.606 | -22.821 | 60.6
|  |  | 4.75 | 13.542 | 1 | -0.038 | 25 | 22.811 | 0.606 | -22.811 | 60.6
|  |  | 4.8 | 13.505 | 1 | -0.038 | 25 | 22.801 | 0.606 | -22.801 | 60.6
|  |  | 4.85 | 13.468 | 1 | -0.038 | 25 | 22.791 | 0.605 | -22.791 | 60.5
|  |  | 4.9 | 13.431 | 1 | -0.038 | 25 | 22.781 | 0.605 | -22.781 | 60.5
|  |  | 4.95 | 13.394 | 1 | -0.038 | 25 | 22.771 | 0.605 | -22.771 | 60.5
|  |  | 5 | 13.358 | 1 | -0.038 | 25 | 22.761 | 0.605 | -22.761 | 60.5
|  |  | 5.05 | 13.321 | 1 | -0.038 | 25 | 22.751 | 0.604 | -22.751 | 60.4
|  |  | 5.1 | 13.285 | 1 | -0.038 | 25 | 22.741 | 0.604 | -22.741 | 60.4
|  |  | 5.15 | 13.25 | 1 | -0.038 | 25 | 22.731 | 0.604 | -22.731 | 60.4
|  |  | 5.2 | 13.214 | 1 | -0.038 | 25 | 22.721 | 0.603 | -22.721 | 60.3
|  |  | 5.25 | 13.179 | 1 | -0.038 | 25 | 22.711 | 0.603 | -22.711 | 60.3
|  |  | 5.3 | 13.143 | 1 | -0.038 | 25 | 22.701 | 0.603 | -22.701 | 60.3
|  |  | 5.35 | 13.108 | 1 | -0.038 | 25 | 22.691 | 0.602 | -22.691 | 60.2
|  |  | 5.4 | 13.072 | 1 | -0.038 | 25 | 22.681 | 0.602 | -22.681 | 60.2
|  |  | 5.45 | 13.037 | 1 | -0.038 | 25 | 22.671 | 0.602 | -22.671 | 60.2
|  |  | 5.5 | 13.002 | 1 | -0.038 | 25 | 22.661 | 0.602 | -22.661 | 60.2
|  |  | 5.55 | 12.967 | 1 | -0.038 | 25 | 22.651 | 0.601 | -22.651 | 60.1
|  |  | 5.6 | 12.931 | 1 | -0.038 | 25 | 22.641 | 0.601 | -22.641 | 60.1
|  |  | 5.65 | 12.896 | 1 | -0.038 | 25 | 22.631 | 0.601 | -22.631 | 60.1
|  |  | 5.7 | 12.861 | 1 | -0.038 | 25 | 22.621 | 0.6 | -22.621 | 60
|  |  | 5.75 | 12.826 | 1 | -0.038 | 25 | 22.611 | 0.6 | -22.611 | 60
|  |  | 5.8 | 12.791 | 1 | -0.038 | 25 | 22.601 | 0.6 | -22.601 | 60
|  |  | 5.85 | 12.756 | 1 | -0.038 | 25 | 22.591 | 0.599 | -22.591 | 59.9
|  |  | 5.9 | 12.722 | 1 | -0.038 | 25 | 22.581 | 0.599 | -22.581 | 59.9
|  |  | 5.95 | 12.688 | 1 | -0.038 | 25 | 22.571 | 0.599 | -22.571 | 59.9
|  |  | 6 | 12.654 | 1 | -0.038 | 25 | 22.561 | 0.599 | -22.561 | 59.9
|  |  | 6.05 | 12.62 | 1 | -0.038 | 25 | 22.551 | 0.598 | -22.551 | 59.8
|  |  | 6.1 | 12.586 | 1 | -0.038 | 25 | 22.541 | 0.598 | -22.541 | 59.8
|  |  | 6.15 | 12.552 | 1 | -0.038 | 25 | 22.531 | 0.598 | -22.531 | 59.8
|  |  | 6.2 | 12.519 | 1 | -0.038 | 25 | 22.521 | 0.597 | -22.521 | 59.7
|  |  | 6.25 | 12.485 | 1 | -0.038 | 25 | 22.511 | 0.597 | -22.511 | 59.7
|  |  | 6.3 | 12.451 | 1 | -0.038 | 25 | 22.501 | 0.597 | -22.501 | 59.7
|  |  | 6.35 | 12.417 | 1 | -0.038 | 25 | 22.491 | 0.596 | -22.491 | 59.6
|  |  | 6.4 | 12.384 | 1 | -0.038 | 25 | 22.481 | 0.596 | -22.481 | 59.6
|  |  | 6.45 | 12.35 | 1 | -0.038 | 25 | 22.471 | 0.596 | -22.471 | 59.6
|  |  | 6.5 | 12.317 | 1 | -0.038 | 25 | 22.461 | 0.596 | -22.461 | 59.6
|  |  | 6.55 | 12.283 | 1 | -0.038 | 25 | 22.451 | 0.595 | -22.451 | 59.5
|  |  | 6.6 | 12.25 | 1 | -0.038 | 25 | 22.441 | 0.595 | -22.441 | 59.5
|  |  | 6.65 | 12.216 | 1 | -0.038 | 25 | 22.431 | 0.595 | -22.431 | 59.5
|  |  | 6.7 | 12.184 | 1 | -0.038 | 25 | 22.421 | 0.594 | -22.421 | 59.4
|  |  | 6.75 | 12.151 | 1 | -0.038 | 25 | 22.411 | 0.594 | -22.411 | 59.4
|  |  | 6.8 | 12.118 | 1 | -0.038 | 25 | 22.401 | 0.594 | -22.401 | 59.4
|  |  | 6.85 | 12.086 | 1 | -0.038 | 25 | 22.391 | 0.593 | -22.391 | 59.3
|  |  | 6.9 | 12.053 | 1 | -0.038 | 25 | 22.381 | 0.593 | -22.381 | 59.3
|  |  | 6.95 | 12.021 | 1 | -0.038 | 25 | 22.371 | 0.593 | -22.371 | 59.3
|  |  | 7 | 11.988 | 1 | -0.038 | 25 | 22.361 | 0.593 | -22.361 | 59.3
|  |  | 7.05 | 11.956 | 1 | -0.038 | 25 | 22.351 | 0.592 | -22.351 | 59.2
|  |  | 7.1 | 11.924 | 1 | -0.038 | 25 | 22.341 | 0.592 | -22.341 | 59.2
|  |  | 7.15 | 11.891 | 1 | -0.039 | 25 | 22.331 | 0.592 | -22.331 | 59.2
|  |  | 7.2 | 11.859 | 1 | -0.039 | 25 | 22.321 | 0.591 | -22.321 | 59.1
|  |  | 7.25 | 11.827 | 1 | -0.039 | 25 | 22.311 | 0.591 | -22.311 | 59.1
|  |  | 7.3 | 11.795 | 1 | -0.039 | 25 | 22.301 | 0.591 | -22.301 | 59.1
|  |  | 7.35 | 11.763 | 1 | -0.039 | 25 | 22.291 | 0.59 | -22.291 | 59
|  |  | 7.4 | 11.731 | 1 | -0.039 | 25 | 22.281 | 0.59 | -22.281 | 59
|  |  | 7.45 | 11.699 | 1 | -0.039 | 25 | 22.271 | 0.59 | -22.271 | 59
|  |  | 7.5 | 11.668 | 1 | -0.039 | 25 | 22.261 | 0.59 | -22.261 | 59
|  |  | 7.55 | 11.637 | 1 | -0.039 | 25 | 22.251 | 0.589 | -22.251 | 58.9
|  |  | 7.6 | 11.606 | 1 | -0.039 | 25 | 22.241 | 0.589 | -22.241 | 58.9
|  |  | 7.65 | 11.575 | 1 | -0.039 | 25 | 22.231 | 0.589 | -22.231 | 58.9
|  |  | 7.7 | 11.545 | 1 | -0.039 | 25 | 22.221 | 0.588 | -22.221 | 58.8
|  |  | 7.75 | 11.514 | 1 | -0.039 | 25 | 22.211 | 0.588 | -22.211 | 58.8
|  |  | 7.8 | 11.484 | 1 | -0.039 | 25 | 22.201 | 0.588 | -22.201 | 58.8
|  |  | 7.85 | 11.453 | 1 | -0.039 | 25 | 22.191 | 0.588 | -22.191 | 58.8
|  |  | 7.9 | 11.423 | 1 | -0.039 | 25 | 22.181 | 0.587 | -22.181 | 58.7
|  |  | 7.95 | 11.392 | 1 | -0.039 | 25 | 22.171 | 0.587 | -22.171 | 58.7
|  |  | 8 | 11.362 | 1 | -0.039 | 25 | 22.161 | 0.587 | -22.161 | 58.7
|  |  | 8.05 | 11.331 | 1 | -0.039 | 25 | 22.151 | 0.586 | -22.151 | 58.6
|  |  | 8.1 | 11.301 | 1 | -0.039 | 25 | 22.141 | 0.586 | -22.141 | 58.6
|  |  | 8.15 | 11.271 | 1 | -0.039 | 25 | 22.131 | 0.586 | -22.131 | 58.6
|  |  | 8.2 | 11.24 | 1 | -0.039 | 25 | 22.121 | 0.585 | -22.121 | 58.5
|  |  | 8.25 | 11.21 | 1 | -0.039 | 25 | 22.111 | 0.585 | -22.111 | 58.5
|  |  | 8.3 | 11.18 | 1 | -0.039 | 25 | 22.101 | 0.585 | -22.101 | 58.5
|  |  | 8.35 | 11.151 | 1 | -0.039 | 25 | 22.091 | 0.585 | -22.091 | 58.5
|  |  | 8.4 | 11.121 | 1 | -0.039 | 25 | 22.081 | 0.584 | -22.081 | 58.4
|  |  | 8.45 | 11.091 | 1 | -0.039 | 25 | 22.071 | 0.584 | -22.071 | 58.4
|  |  | 8.5 | 11.062 | 1 | -0.039 | 25 | 22.061 | 0.584 | -22.061 | 58.4
|  |  | 8.55 | 11.032 | 1 | -0.039 | 25 | 22.051 | 0.583 | -22.051 | 58.3
|  |  | 8.6 | 11.003 | 1 | -0.039 | 25 | 22.041 | 0.583 | -22.041 | 58.3
|  |  | 8.65 | 10.973 | 1 | -0.039 | 25 | 22.031 | 0.583 | -22.031 | 58.3
|  |  | 8.7 | 10.944 | 1 | -0.039 | 25 | 22.021 | 0.582 | -22.021 | 58.2
|  |  | 8.75 | 10.914 | 1 | -0.039 | 25 | 22.011 | 0.582 | -22.011 | 58.2
|  |  | 8.8 | 10.885 | 1 | -0.039 | 25 | 22.001 | 0.582 | -22.001 | 58.2
|  |  | 8.85 | 10.855 | 1 | -0.039 | 25 | 21.991 | 0.582 | -21.991 | 58.2
|  |  | 8.9 | 10.826 | 1 | -0.039 | 25 | 21.981 | 0.581 | -21.981 | 58.1
|  |  | 8.95 | 10.797 | 1 | -0.039 | 25 | 21.971 | 0.581 | -21.971 | 58.1
|  |  | 9 | 10.768 | 1 | -0.039 | 25 | 21.961 | 0.581 | -21.961 | 58.1
|  |  | 9.05 | 10.738 | 1 | -0.039 | 25 | 21.951 | 0.58 | -21.951 | 58
|  |  | 9.1 | 10.71 | 1 | -0.039 | 25 | 21.941 | 0.58 | -21.941 | 58
|  |  | 9.15 | 10.682 | 1 | -0.039 | 25 | 21.931 | 0.58 | -21.931 | 58
|  |  | 9.2 | 10.654 | 1 | -0.039 | 25 | 21.921 | 0.58 | -21.921 | 58
|  |  | 9.25 | 10.626 | 1 | -0.039 | 25 | 21.911 | 0.579 | -21.911 | 57.9
|  |  | 9.3 | 10.598 | 1 | -0.039 | 25 | 21.901 | 0.579 | -21.901 | 57.9
|  |  | 9.35 | 10.569 | 1 | -0.039 | 25 | 21.891 | 0.579 | -21.891 | 57.9
|  |  | 9.4 | 10.541 | 1 | -0.039 | 25 | 21.881 | 0.578 | -21.881 | 57.8
|  |  | 9.45 | 10.513 | 1 | -0.039 | 25 | 21.871 | 0.578 | -21.871 | 57.8
|  |  | 9.5 | 10.485 | 1 | -0.039 | 25 | 21.861 | 0.578 | -21.861 | 57.8
|  |  | 9.55 | 10.457 | 1 | -0.039 | 25 | 21.851 | 0.577 | -21.851 | 57.7
|  |  | 9.6 | 10.43 | 1 | -0.039 | 25 | 21.841 | 0.577 | -21.841 | 57.7
|  |  | 9.65 | 10.402 | 1 | -0.039 | 25 | 21.831 | 0.577 | -21.831 | 57.7
|  |  | 9.7 | 10.374 | 1 | -0.039 | 25 | 21.821 | 0.577 | -21.821 | 57.7
|  |  | 9.75 | 10.346 | 1 | -0.039 | 25 | 21.811 | 0.576 | -21.811 | 57.6
|  |  | 9.8 | 10.318 | 1 | -0.039 | 25 | 21.801 | 0.576 | -21.801 | 57.6
|  |  | 9.85 | 10.291 | 1 | -0.039 | 25 | 21.791 | 0.576 | -21.791 | 57.6
|  |  | 9.9 | 10.264 | 1 | -0.039 | 25 | 21.781 | 0.575 | -21.781 | 57.5
|  |  | 9.95 | 10.237 | 1 | -0.039 | 25 | 21.771 | 0.575 | -21.771 | 57.5
|  |  | 10 | 10.21 | 1 | -0.039 | 25 | 21.761 | 0.575 | -21.761 | 57.5
|  |  | 10.05 | 10.183 | 1 | -0.039 | 25 | 21.751 | 0.575 | -21.751 | 57.5
|  |  | 10.1 | 10.156 | 1 | -0.039 | 25 | 21.741 | 0.574 | -21.741 | 57.4
|  |  | 10.15 | 10.129 | 1 | -0.039 | 25 | 21.731 | 0.574 | -21.731 | 57.4
|  |  | 10.2 | 10.103 | 1 | -0.039 | 25 | 21.721 | 0.574 | -21.721 | 57.4
|  |  | 10.25 | 10.076 | 1 | -0.039 | 25 | 21.711 | 0.573 | -21.711 | 57.3
|  |  | 10.3 | 10.049 | 1 | -0.039 | 25 | 21.701 | 0.573 | -21.701 | 57.3
|  |  | 10.35 | 10.023 | 1 | -0.039 | 25 | 21.691 | 0.573 | -21.691 | 57.3
|  |  | 10.4 | 9.996 | 1 | -0.039 | 25 | 21.681 | 0.572 | -21.681 | 57.2
|  |  | 10.45 | 9.9693 | 1 | -0.039 | 25 | 21.671 | 0.572 | -21.671 | 57.2
|  |  | 10.5 | 9.9429 | 1 | -0.039 | 25 | 21.661 | 0.572 | -21.661 | 57.2
|  |  | 10.55 | 9.9164 | 1 | -0.039 | 25 | 21.651 | 0.572 | -21.651 | 57.2
|  |  | 10.6 | 9.8899 | 1 | -0.04 | 25 | 21.641 | 0.571 | -21.641 | 57.1
|  |  | 10.65 | 9.8634 | 1 | -0.04 | 25 | 21.631 | 0.571 | -21.631 | 57.1
|  |  | 10.7 | 9.8375 | 1 | -0.04 | 25 | 21.621 | 0.571 | -21.621 | 57.1
|  |  | 10.75 | 9.8116 | 1 | -0.04 | 25 | 21.611 | 0.57 | -21.611 | 57
|  |  | 10.8 | 9.7857 | 1 | -0.04 | 25 | 21.601 | 0.57 | -21.601 | 57
|  |  | 10.85 | 9.7598 | 1 | -0.04 | 25 | 21.591 | 0.57 | -21.591 | 57
|  |  | 10.9 | 9.7341 | 1 | -0.04 | 25 | 21.581 | 0.57 | -21.581 | 57
|  |  | 10.95 | 9.7083 | 1 | -0.04 | 25 | 21.571 | 0.569 | -21.571 | 56.9
|  |  | 11 | 9.6826 | 1 | -0.04 | 25 | 21.561 | 0.569 | -21.561 | 56.9
|  |  | 11.05 | 9.6568 | 1 | -0.04 | 25 | 21.551 | 0.569 | -21.551 | 56.9
|  |  | 11.1 | 9.6312 | 1 | -0.04 | 25 | 21.541 | 0.568 | -21.541 | 56.8
|  |  | 11.15 | 9.6056 | 1 | -0.04 | 25 | 21.531 | 0.568 | -21.531 | 56.8
|  |  | 11.2 | 9.58 | 1 | -0.04 | 25 | 21.521 | 0.568 | -21.521 | 56.8
|  |  | 11.25 | 9.5543 | 1 | -0.04 | 25 | 21.511 | 0.568 | -21.511 | 56.8
|  |  | 11.3 | 9.5288 | 1 | -0.04 | 25 | 21.501 | 0.567 | -21.501 | 56.7
|  |  | 11.35 | 9.5034 | 1 | -0.04 | 25 | 21.491 | 0.567 | -21.491 | 56.7
|  |  | 11.4 | 9.4779 | 1 | -0.04 | 25 | 21.481 | 0.567 | -21.481 | 56.7
|  |  | 11.45 | 9.4524 | 1 | -0.04 | 25 | 21.471 | 0.566 | -21.471 | 56.6
|  |  | 11.5 | 9.4278 | 1 | -0.04 | 25 | 21.461 | 0.566 | -21.461 | 56.6
|  |  | 11.55 | 9.4031 | 1 | -0.04 | 25 | 21.451 | 0.566 | -21.451 | 56.6
|  |  | 11.6 | 9.3785 | 1 | -0.04 | 25 | 21.441 | 0.565 | -21.441 | 56.5
|  |  | 11.65 | 9.3539 | 1 | -0.04 | 25 | 21.431 | 0.565 | -21.431 | 56.5
|  |  | 11.7 | 9.3293 | 1 | -0.04 | 25 | 21.421 | 0.565 | -21.421 | 56.5
|  |  | 11.75 | 9.3048 | 1 | -0.04 | 25 | 21.411 | 0.565 | -21.411 | 56.5
|  |  | 11.8 | 9.2803 | 1 | -0.04 | 25 | 21.401 | 0.564 | -21.401 | 56.4
|  |  | 11.85 | 9.2558 | 1 | -0.04 | 25 | 21.391 | 0.564 | -21.391 | 56.4
|  |  | 11.9 | 9.2315 | 1 | -0.04 | 25 | 21.381 | 0.564 | -21.381 | 56.4
|  |  | 11.95 | 9.2071 | 1 | -0.04 | 25 | 21.371 | 0.563 | -21.371 | 56.3
|  |  | 12 | 9.1827 | 1 | -0.04 | 25 | 21.361 | 0.563 | -21.361 | 56.3
|  |  | 12.05 | 9.1583 | 1 | -0.04 | 25 | 21.351 | 0.563 | -21.351 | 56.3
|  |  | 12.1 | 9.1341 | 1 | -0.04 | 25 | 21.341 | 0.563 | -21.341 | 56.3
|  |  | 12.15 | 9.1098 | 1 | -0.04 | 25 | 21.331 | 0.562 | -21.331 | 56.2
|  |  | 12.2 | 9.0856 | 1 | -0.04 | 25 | 21.321 | 0.562 | -21.321 | 56.2
|  |  | 12.25 | 9.0613 | 1 | -0.04 | 25 | 21.311 | 0.562 | -21.311 | 56.2
|  |  | 12.3 | 9.0375 | 1 | -0.04 | 25 | 21.301 | 0.561 | -21.301 | 56.1
|  |  | 12.35 | 9.0137 | 1 | -0.04 | 25 | 21.291 | 0.561 | -21.291 | 56.1
|  |  | 12.4 | 8.9898 | 1 | -0.04 | 25 | 21.281 | 0.561 | -21.281 | 56.1
|  |  | 12.45 | 8.966 | 1 | -0.04 | 25 | 21.271 | 0.56 | -21.271 | 56
|  |  | 12.5 | 8.9423 | 1 | -0.04 | 25 | 21.261 | 0.56 | -21.261 | 56
|  |  | 12.55 | 8.9186 | 1 | -0.04 | 25 | 21.251 | 0.56 | -21.251 | 56
|  |  | 12.6 | 8.8948 | 1 | -0.04 | 25 | 21.241 | 0.56 | -21.241 | 56
|  |  | 12.65 | 8.8711 | 1 | -0.04 | 25 | 21.231 | 0.559 | -21.231 | 55.9
|  |  | 12.7 | 8.8475 | 1 | -0.04 | 25 | 21.221 | 0.559 | -21.221 | 55.9
|  |  | 12.75 | 8.824 | 1 | -0.04 | 25 | 21.211 | 0.559 | -21.211 | 55.9
|  |  | 12.8 | 8.8004 | 1 | -0.04 | 25 | 21.201 | 0.558 | -21.201 | 55.8
|  |  | 12.85 | 8.7768 | 1 | -0.04 | 25 | 21.191 | 0.558 | -21.191 | 55.8
|  |  | 12.9 | 8.7533 | 1 | -0.04 | 25 | 21.181 | 0.558 | -21.181 | 55.8
|  |  | 12.95 | 8.7299 | 1 | -0.04 | 25 | 21.171 | 0.558 | -21.171 | 55.8
|  |  | 13 | 8.7064 | 1 | -0.04 | 25 | 21.161 | 0.557 | -21.161 | 55.7
|  |  | 13.05 | 8.6829 | 1 | -0.04 | 25 | 21.152 | 0.557 | -21.152 | 55.7
|  |  | 13.1 | 8.6603 | 1 | -0.04 | 25 | 21.142 | 0.557 | -21.142 | 55.7
|  |  | 13.15 | 8.6376 | 1 | -0.04 | 25 | 21.132 | 0.556 | -21.132 | 55.6
|  |  | 13.2 | 8.6149 | 1 | -0.04 | 25 | 21.122 | 0.556 | -21.122 | 55.6
|  |  | 13.25 | 8.5923 | 1 | -0.04 | 25 | 21.112 | 0.556 | -21.112 | 55.6
|  |  | 13.3 | 8.5697 | 1 | -0.04 | 25 | 21.102 | 0.555 | -21.102 | 55.5
|  |  | 13.35 | 8.5472 | 1 | -0.04 | 25 | 21.092 | 0.555 | -21.092 | 55.5
|  |  | 13.4 | 8.5246 | 1 | -0.04 | 25 | 21.082 | 0.555 | -21.082 | 55.5
|  |  | 13.45 | 8.5021 | 1 | -0.04 | 25 | 21.072 | 0.555 | -21.072 | 55.5
|  |  | 13.5 | 8.4797 | 1 | -0.04 | 25 | 21.062 | 0.554 | -21.062 | 55.4
|  |  | 13.55 | 8.4572 | 1 | -0.04 | 25 | 21.052 | 0.554 | -21.052 | 55.4
|  |  | 13.6 | 8.4348 | 1 | -0.04 | 25 | 21.042 | 0.554 | -21.042 | 55.4
|  |  | 13.65 | 8.4124 | 1 | -0.04 | 25 | 21.032 | 0.553 | -21.032 | 55.3
|  |  | 13.7 | 8.3901 | 1 | -0.04 | 25 | 21.022 | 0.553 | -21.022 | 55.3
|  |  | 13.75 | 8.3678 | 1 | -0.04 | 25 | 21.012 | 0.553 | -21.012 | 55.3
|  |  | 13.8 | 8.3455 | 1 | -0.04 | 25 | 21.002 | 0.553 | -21.002 | 55.3
|  |  | 13.85 | 8.3231 | 1 | -0.04 | 25 | 20.992 | 0.552 | -20.992 | 55.2
|  |  | 13.9 | 8.3015 | 1 | -0.04 | 25 | 20.982 | 0.552 | -20.982 | 55.2
|  |  | 13.95 | 8.2799 | 1 | -0.04 | 25 | 20.972 | 0.552 | -20.972 | 55.2
|  |  | 14 | 8.2583 | 1 | -0.04 | 25 | 20.962 | 0.551 | -20.962 | 55.1
|  |  | 14.05 | 8.2367 | 1 | -0.04 | 25 | 20.952 | 0.551 | -20.952 | 55.1
|  |  | 14.1 | 8.2152 | 1 | -0.041 | 25 | 20.942 | 0.551 | -20.942 | 55.1
|  |  | 14.15 | 8.1937 | 1 | -0.041 | 25 | 20.932 | 0.551 | -20.932 | 55.1
|  |  | 14.2 | 8.1723 | 1 | -0.041 | 25 | 20.922 | 0.55 | -20.922 | 55
|  |  | 14.25 | 8.1508 | 1 | -0.041 | 25 | 20.912 | 0.55 | -20.912 | 55
|  |  | 14.3 | 8.1294 | 1 | -0.041 | 25 | 20.902 | 0.55 | -20.902 | 55
|  |  | 14.35 | 8.108 | 1 | -0.041 | 25 | 20.892 | 0.549 | -20.892 | 54.9
|  |  | 14.4 | 8.0866 | 1 | -0.041 | 25 | 20.882 | 0.549 | -20.882 | 54.9
|  |  | 14.45 | 8.0652 | 1 | -0.041 | 25 | 20.872 | 0.549 | -20.872 | 54.9
|  |  | 14.5 | 8.044 | 1 | -0.041 | 25 | 20.862 | 0.548 | -20.862 | 54.8
|  |  | 14.55 | 8.0227 | 1 | -0.041 | 25 | 20.852 | 0.548 | -20.852 | 54.8
|  |  | 14.6 | 8.0014 | 1 | -0.041 | 25 | 20.842 | 0.548 | -20.842 | 54.8
|  |  | 14.65 | 7.9802 | 1 | -0.041 | 25 | 20.832 | 0.548 | -20.832 | 54.8
|  |  | 14.7 | 7.9594 | 1 | -0.041 | 25 | 20.822 | 0.547 | -20.822 | 54.7
|  |  | 14.75 | 7.9386 | 1 | -0.041 | 25 | 20.812 | 0.547 | -20.812 | 54.7
|  |  | 14.8 | 7.9178 | 1 | -0.041 | 25 | 20.802 | 0.547 | -20.802 | 54.7
|  |  | 14.85 | 7.897 | 1 | -0.041 | 25 | 20.792 | 0.546 | -20.792 | 54.6
|  |  | 14.9 | 7.8763 | 1 | -0.041 | 25 | 20.782 | 0.546 | -20.782 | 54.6
|  |  | 14.95 | 7.8556 | 1 | -0.041 | 25 | 20.772 | 0.546 | -20.772 | 54.6
|  |  | 15 | 7.835 | 1 | -0.041 | 25 | 20.762 | 0.546 | -20.762 | 54.6
|  |  | 15.05 | 7.8143 | 1 | -0.041 | 25 | 20.752 | 0.545 | -20.752 | 54.5
|  |  | 15.1 | 7.7937 | 1 | -0.041 | 25 | 20.742 | 0.545 | -20.742 | 54.5
|  |  | 15.15 | 7.7731 | 1 | -0.041 | 25 | 20.732 | 0.545 | -20.732 | 54.5
|  |  | 15.2 | 7.7526 | 1 | -0.041 | 25 | 20.722 | 0.544 | -20.722 | 54.4
|  |  | 15.25 | 7.732 | 1 | -0.041 | 25 | 20.712 | 0.544 | -20.712 | 54.4
|  |  | 15.3 | 7.7115 | 1 | -0.041 | 25 | 20.702 | 0.544 | -20.702 | 54.4
|  |  | 15.35 | 7.691 | 1 | -0.041 | 25 | 20.692 | 0.544 | -20.692 | 54.4
|  |  | 15.4 | 7.6706 | 1 | -0.041 | 25 | 20.682 | 0.543 | -20.682 | 54.3
|  |  | 15.45 | 7.6501 | 1 | -0.041 | 25 | 20.672 | 0.543 | -20.672 | 54.3
|  |  | 15.5 | 7.6304 | 1 | -0.041 | 25 | 20.662 | 0.543 | -20.662 | 54.3
|  |  | 15.55 | 7.6107 | 1 | -0.041 | 25 | 20.652 | 0.542 | -20.652 | 54.2
|  |  | 15.6 | 7.591 | 1 | -0.041 | 25 | 20.642 | 0.542 | -20.642 | 54.2
|  |  | 15.65 | 7.5713 | 1 | -0.041 | 25 | 20.632 | 0.542 | -20.632 | 54.2
|  |  | 15.7 | 7.5517 | 1 | -0.041 | 25 | 20.622 | 0.541 | -20.622 | 54.1
|  |  | 15.75 | 7.5321 | 1 | -0.041 | 25 | 20.612 | 0.541 | -20.612 | 54.1
|  |  | 15.8 | 7.5124 | 1 | -0.041 | 25 | 20.602 | 0.541 | -20.602 | 54.1
|  |  | 15.85 | 7.4928 | 1 | -0.041 | 25 | 20.592 | 0.541 | -20.592 | 54.1
|  |  | 15.9 | 7.4733 | 1 | -0.041 | 25 | 20.582 | 0.54 | -20.582 | 54
|  |  | 15.95 | 7.4538 | 1 | -0.041 | 25 | 20.572 | 0.54 | -20.572 | 54
|  |  | 16 | 7.4343 | 1 | -0.041 | 25 | 20.562 | 0.54 | -20.562 | 54
|  |  | 16.05 | 7.4148 | 1 | -0.041 | 25 | 20.552 | 0.539 | -20.552 | 53.9
|  |  | 16.1 | 7.3954 | 1 | -0.041 | 25 | 20.542 | 0.539 | -20.542 | 53.9
|  |  | 16.15 | 7.376 | 1 | -0.041 | 25 | 20.532 | 0.539 | -20.532 | 53.9
|  |  | 16.2 | 7.3566 | 1 | -0.041 | 25 | 20.522 | 0.539 | -20.522 | 53.9
|  |  | 16.25 | 7.3372 | 1 | -0.041 | 25 | 20.512 | 0.538 | -20.512 | 53.8
|  |  | 16.3 | 7.318 | 1 | -0.041 | 25 | 20.502 | 0.538 | -20.502 | 53.8
|  |  | 16.35 | 7.2988 | 1 | -0.041 | 25 | 20.492 | 0.538 | -20.492 | 53.8
|  |  | 16.4 | 7.2796 | 1 | -0.041 | 25 | 20.482 | 0.538 | -20.482 | 53.8
|  |  | 16.45 | 7.2605 | 1 | -0.041 | 25 | 20.472 | 0.537 | -20.472 | 53.7
|  |  | 16.5 | 7.2414 | 1 | -0.041 | 25 | 20.462 | 0.537 | -20.462 | 53.7
|  |  | 16.55 | 7.2223 | 1 | -0.041 | 25 | 20.452 | 0.537 | -20.452 | 53.7
|  |  | 16.6 | 7.2032 | 1 | -0.041 | 25 | 20.442 | 0.536 | -20.442 | 53.6
|  |  | 16.65 | 7.1841 | 1 | -0.041 | 25 | 20.432 | 0.536 | -20.432 | 53.6
|  |  | 16.7 | 7.1652 | 1 | -0.041 | 25 | 20.422 | 0.536 | -20.422 | 53.6
|  |  | 16.75 | 7.1462 | 1 | -0.041 | 25 | 20.412 | 0.536 | -20.412 | 53.6
|  |  | 16.8 | 7.1272 | 1 | -0.041 | 25 | 20.402 | 0.535 | -20.402 | 53.5
|  |  | 16.85 | 7.1082 | 1 | -0.041 | 25 | 20.392 | 0.535 | -20.392 | 53.5
|  |  | 16.9 | 7.0893 | 1 | -0.041 | 25 | 20.382 | 0.535 | -20.382 | 53.5
|  |  | 16.95 | 7.0704 | 1 | -0.041 | 25 | 20.372 | 0.535 | -20.372 | 53.5
|  |  | 17 | 7.0516 | 1 | -0.041 | 25 | 20.362 | 0.534 | -20.362 | 53.4
|  |  | 17.05 | 7.0327 | 1 | -0.041 | 25 | 20.352 | 0.534 | -20.352 | 53.4
|  |  | 17.1 | 7.0144 | 1 | -0.041 | 25 | 20.342 | 0.534 | -20.342 | 53.4
|  |  | 17.15 | 6.996 | 1 | -0.041 | 25 | 20.332 | 0.533 | -20.332 | 53.3
|  |  | 17.2 | 6.9777 | 1 | -0.041 | 25 | 20.322 | 0.533 | -20.322 | 53.3
|  |  | 17.25 | 6.9594 | 1 | -0.041 | 25 | 20.312 | 0.533 | -20.312 | 53.3
|  |  | 17.3 | 6.9412 | 1 | -0.041 | 25 | 20.302 | 0.533 | -20.302 | 53.3
|  |  | 17.35 | 6.923 | 1 | -0.041 | 25 | 20.292 | 0.532 | -20.292 | 53.2
|  |  | 17.4 | 6.9047 | 1 | -0.041 | 25 | 20.282 | 0.532 | -20.282 | 53.2
|  |  | 17.45 | 6.8865 | 1 | -0.041 | 25 | 20.272 | 0.532 | -20.272 | 53.2
|  |  | 17.5 | 6.8684 | 1 | -0.041 | 25 | 20.262 | 0.532 | -20.262 | 53.2
|  |  | 17.55 | 6.8502 | 1 | -0.041 | 25 | 20.252 | 0.531 | -20.252 | 53.1
|  |  | 17.6 | 6.8321 | 1 | -0.042 | 25 | 20.242 | 0.531 | -20.242 | 53.1
|  |  | 17.65 | 6.814 | 1 | -0.042 | 25 | 20.232 | 0.531 | -20.232 | 53.1
|  |  | 17.7 | 6.7959 | 1 | -0.042 | 25 | 20.222 | 0.53 | -20.222 | 53
|  |  | 17.75 | 6.7779 | 1 | -0.042 | 25 | 20.212 | 0.53 | -20.212 | 53
|  |  | 17.8 | 6.7599 | 1 | -0.042 | 25 | 20.202 | 0.53 | -20.202 | 53
|  |  | 17.85 | 6.7418 | 1 | -0.042 | 25 | 20.192 | 0.53 | -20.192 | 53
|  |  | 17.9 | 6.7243 | 1 | -0.042 | 25 | 20.182 | 0.529 | -20.182 | 52.9
|  |  | 17.95 | 6.7069 | 1 | -0.042 | 25 | 20.172 | 0.529 | -20.172 | 52.9
|  |  | 18 | 6.6894 | 1 | -0.042 | 25 | 20.162 | 0.529 | -20.162 | 52.9
|  |  | 18.05 | 6.6719 | 1 | -0.042 | 25 | 20.152 | 0.529 | -20.152 | 52.9
|  |  | 18.1 | 6.6545 | 1 | -0.042 | 25 | 20.142 | 0.528 | -20.142 | 52.8
|  |  | 18.15 | 6.6372 | 1 | -0.042 | 25 | 20.132 | 0.528 | -20.132 | 52.8
|  |  | 18.2 | 6.6198 | 1 | -0.042 | 25 | 20.122 | 0.528 | -20.122 | 52.8
|  |  | 18.25 | 6.6024 | 1 | -0.042 | 25 | 20.112 | 0.527 | -20.112 | 52.7
|  |  | 18.3 | 6.5851 | 1 | -0.042 | 25 | 20.102 | 0.527 | -20.102 | 52.7
|  |  | 18.35 | 6.5678 | 1 | -0.042 | 25 | 20.092 | 0.527 | -20.092 | 52.7
|  |  | 18.4 | 6.5505 | 1 | -0.042 | 25 | 20.082 | 0.527 | -20.082 | 52.7
|  |  | 18.45 | 6.5332 | 1 | -0.042 | 25 | 20.072 | 0.526 | -20.072 | 52.6
|  |  | 18.5 | 6.516 | 1 | -0.042 | 25 | 20.062 | 0.526 | -20.062 | 52.6
|  |  | 18.55 | 6.4988 | 1 | -0.042 | 25 | 20.052 | 0.526 | -20.052 | 52.6
|  |  | 18.6 | 6.4816 | 1 | -0.042 | 25 | 20.042 | 0.526 | -20.042 | 52.6
|  |  | 18.65 | 6.4644 | 1 | -0.042 | 25 | 20.032 | 0.525 | -20.032 | 52.5
|  |  | 18.7 | 6.4476 | 1 | -0.042 | 25 | 20.022 | 0.525 | -20.022 | 52.5
|  |  | 18.75 | 6.4307 | 1 | -0.042 | 25 | 20.012 | 0.525 | -20.012 | 52.5
|  |  | 18.8 | 6.4139 | 1 | -0.042 | 25 | 20.002 | 0.524 | -20.002 | 52.4
|  |  | 18.85 | 6.3971 | 1 | -0.042 | 25 | 19.992 | 0.524 | -19.992 | 52.4
|  |  | 18.9 | 6.3803 | 1 | -0.042 | 25 | 19.982 | 0.524 | -19.982 | 52.4
|  |  | 18.95 | 6.3636 | 1 | -0.042 | 25 | 19.972 | 0.524 | -19.972 | 52.4
|  |  | 19 | 6.3469 | 1 | -0.042 | 25 | 19.962 | 0.523 | -19.962 | 52.3
|  |  | 19.05 | 6.3301 | 1 | -0.042 | 25 | 19.952 | 0.523 | -19.952 | 52.3
|  |  | 19.1 | 6.3135 | 1 | -0.042 | 25 | 19.942 | 0.523 | -19.942 | 52.3
|  |  | 19.15 | 6.2968 | 1 | -0.042 | 25 | 19.932 | 0.523 | -19.932 | 52.3
|  |  | 19.2 | 6.2801 | 1 | -0.042 | 25 | 19.922 | 0.522 | -19.922 | 52.2
|  |  | 19.25 | 6.2635 | 1 | -0.042 | 25 | 19.912 | 0.522 | -19.912 | 52.2
|  |  | 19.3 | 6.2469 | 1 | -0.042 | 25 | 19.902 | 0.522 | -19.902 | 52.2
|  |  | 19.35 | 6.2303 | 1 | -0.042 | 25 | 19.892 | 0.521 | -19.892 | 52.1
|  |  | 19.4 | 6.2138 | 1 | -0.042 | 25 | 19.882 | 0.521 | -19.882 | 52.1
|  |  | 19.45 | 6.1972 | 1 | -0.042 | 25 | 19.872 | 0.521 | -19.872 | 52.1
|  |  | 19.5 | 6.1812 | 1 | -0.042 | 25 | 19.862 | 0.521 | -19.862 | 52.1
|  |  | 19.55 | 6.1653 | 1 | -0.042 | 25 | 19.852 | 0.52 | -19.852 | 52
|  |  | 19.6 | 6.1493 | 1 | -0.042 | 25 | 19.842 | 0.52 | -19.842 | 52
|  |  | 19.65 | 6.1334 | 1 | -0.042 | 25 | 19.832 | 0.52 | -19.832 | 52
|  |  | 19.7 | 6.1175 | 1 | -0.042 | 25 | 19.822 | 0.52 | -19.822 | 52
|  |  | 19.75 | 6.1016 | 1 | -0.042 | 25 | 19.812 | 0.519 | -19.812 | 51.9
|  |  | 19.8 | 6.0857 | 1 | -0.042 | 25 | 19.802 | 0.519 | -19.802 | 51.9
|  |  | 19.85 | 6.0699 | 1 | -0.042 | 25 | 19.792 | 0.519 | -19.792 | 51.9
|  |  | 19.9 | 6.0541 | 1 | -0.042 | 25 | 19.782 | 0.518 | -19.782 | 51.8
|  | END_DATA |  |  |  | -0.042 | 25 | 19.772 | 0.518 | -19.772 | 51.8
| END_SCAN 1 |  |  |  |  | -0.042 | 25 | 19.762 | 0.518 | -19.762 | 51.8
END_SCAN_DATA | [5 empty] | -0.042 | 25 | 19.752 | 0.518 | -19.752 | 51.8
[6 empty] | -0.042 | 25 | 19.742 | 0.517 | -19.742 | 51.7
[6 empty] | -0.042 | 25 | 19.732 | 0.517 | -19.732 | 51.7
[6 empty] | -0.042 | 25 | 19.722 | 0.517 | -19.722 | 51.7
[6 empty] | -0.042 | 25 | 19.712 | 0.517 | -19.712 | 51.7
[6 empty] | -0.042 | 25 | 19.702 | 0.516 | -19.702 | 51.6
[6 empty] | -0.042 | 25 | 19.692 | 0.516 | -19.692 | 51.6
[6 empty] | -0.042 | 25 | 19.682 | 0.516 | -19.682 | 51.6
[6 empty] | -0.042 | 25 | 19.672 | 0.516 | -19.672 | 51.6
[6 empty] | -0.042 | 25 | 19.662 | 0.515 | -19.662 | 51.5
[6 empty] | -0.042 | 25 | 19.652 | 0.515 | -19.652 | 51.5
[6 empty] | -0.042 | 25 | 19.642 | 0.515 | -19.642 | 51.5
[6 empty] | -0.042 | 25 | 19.632 | 0.514 | -19.632 | 51.4
[6 empty] | -0.042 | 25 | 19.622 | 0.514 | -19.622 | 51.4
[6 empty] | -0.042 | 25 | 19.612 | 0.514 | -19.612 | 51.4
[6 empty] | -0.042 | 25 | 19.602 | 0.514 | -19.602 | 51.4
[6 empty] | -0.042 | 25 | 19.592 | 0.513 | -19.592 | 51.3
[6 empty] | -0.042 | 25 | 19.582 | 0.513 | -19.582 | 51.3
[6 empty] | -0.042 | 25 | 19.572 | 0.513 | -19.572 | 51.3
[6 empty] | -0.042 | 25 | 19.562 | 0.513 | -19.562 | 51.3
[6 empty] | -0.042 | 25 | 19.552 | 0.512 | -19.552 | 51.2
[6 empty] | -0.043 | 25 | 19.542 | 0.512 | -19.542 | 51.2
[6 empty] | -0.043 | 25 | 19.532 | 0.512 | -19.532 | 51.2
[6 empty] | -0.043 | 25 | 19.522 | 0.511 | -19.522 | 51.1
[6 empty] | -0.043 | 25 | 19.512 | 0.511 | -19.512 | 51.1
[6 empty] | -0.043 | 25 | 19.502 | 0.511 | -19.502 | 51.1
[6 empty] | -0.043 | 25 | 19.492 | 0.511 | -19.492 | 51.1
[6 empty] | -0.043 | 25 | 19.482 | 0.51 | -19.482 | 51
[6 empty] | -0.043 | 25 | 19.472 | 0.51 | -19.472 | 51
[6 empty] | -0.043 | 25 | 19.462 | 0.51 | -19.462 | 51
[6 empty] | -0.043 | 25 | 19.452 | 0.51 | -19.452 | 51
[6 empty] | -0.043 | 25 | 19.442 | 0.509 | -19.442 | 50.9
[6 empty] | -0.043 | 25 | 19.432 | 0.509 | -19.432 | 50.9
[6 empty] | -0.043 | 25 | 19.422 | 0.509 | -19.422 | 50.9
[6 empty] | -0.043 | 25 | 19.412 | 0.508 | -19.412 | 50.8
[6 empty] | -0.043 | 25 | 19.402 | 0.508 | -19.402 | 50.8
[6 empty] | -0.043 | 25 | 19.392 | 0.508 | -19.392 | 50.8
[6 empty] | -0.043 | 25 | 19.382 | 0.508 | -19.382 | 50.8
[6 empty] | -0.043 | 25 | 19.372 | 0.507 | -19.372 | 50.7
[6 empty] | -0.043 | 25 | 19.363 | 0.507 | -19.363 | 50.7
[6 empty] | -0.043 | 25 | 19.353 | 0.507 | -19.353 | 50.7
[6 empty] | -0.043 | 25 | 19.343 | 0.507 | -19.343 | 50.7
[6 empty] | -0.043 | 25 | 19.333 | 0.506 | -19.333 | 50.6
[6 empty] | -0.043 | 25 | 19.323 | 0.506 | -19.323 | 50.6
[6 empty] | -0.043 | 25 | 19.313 | 0.506 | -19.313 | 50.6
[6 empty] | -0.043 | 25 | 19.303 | 0.506 | -19.303 | 50.6
[6 empty] | -0.043 | 25 | 19.293 | 0.505 | -19.293 | 50.5
[6 empty] | -0.043 | 25 | 19.283 | 0.505 | -19.283 | 50.5
[6 empty] | -0.043 | 25 | 19.273 | 0.505 | -19.273 | 50.5
[6 empty] | -0.043 | 25 | 19.263 | 0.504 | -19.263 | 50.4
[6 empty] | -0.043 | 25 | 19.253 | 0.504 | -19.253 | 50.4
[6 empty] | -0.043 | 25 | 19.243 | 0.504 | -19.243 | 50.4
[6 empty] | -0.043 | 25 | 19.233 | 0.504 | -19.233 | 50.4
[6 empty] | -0.043 | 25 | 19.223 | 0.503 | -19.223 | 50.3
[6 empty] | -0.043 | 25 | 19.213 | 0.503 | -19.213 | 50.3
[6 empty] | -0.043 | 25 | 19.203 | 0.503 | -19.203 | 50.3
[6 empty] | -0.043 | 25 | 19.193 | 0.503 | -19.193 | 50.3
[6 empty] | -0.043 | 25 | 19.183 | 0.502 | -19.183 | 50.2
[6 empty] | -0.043 | 25 | 19.173 | 0.502 | -19.173 | 50.2
[6 empty] | -0.043 | 25 | 19.163 | 0.502 | -19.163 | 50.2
[6 empty] | -0.043 | 25 | 19.153 | 0.502 | -19.153 | 50.2
[6 empty] | -0.043 | 25 | 19.143 | 0.501 | -19.143 | 50.1
[6 empty] | -0.043 | 25 | 19.133 | 0.501 | -19.133 | 50.1
[6 empty] | -0.043 | 25 | 19.123 | 0.501 | -19.123 | 50.1
[6 empty] | -0.043 | 25 | 19.113 | 0.5 | -19.113 | 50
[6 empty] | -0.043 | 25 | 19.103 | 0.5 | -19.103 | 50
[6 empty] | -0.043 | 25 | 19.093 | 0.5 | -19.093 | 50
[6 empty] | -0.043 | 25 | 19.083 | 0.5 | -19.083 | 50
[6 empty] | -0.043 | 25 | 19.073 | 0.499 | -19.073 | 49.9
[6 empty] | -0.043 | 25 | 19.063 | 0.499 | -19.063 | 49.9
[6 empty] | -0.043 | 25 | 19.053 | 0.499 | -19.053 | 49.9
[6 empty] | -0.043 | 25 | 19.043 | 0.499 | -19.043 | 49.9
[6 empty] | -0.043 | 25 | 19.033 | 0.498 | -19.033 | 49.8
[6 empty] | -0.043 | 25 | 19.023 | 0.498 | -19.023 | 49.8
[6 empty] | -0.043 | 25 | 19.013 | 0.498 | -19.013 | 49.8
[6 empty] | -0.043 | 25 | 19.003 | 0.498 | -19.003 | 49.8
[6 empty] | -0.043 | 25 | 18.993 | 0.497 | -18.993 | 49.7
[6 empty] | -0.043 | 25 | 18.983 | 0.497 | -18.983 | 49.7
[6 empty] | -0.043 | 25 | 18.973 | 0.497 | -18.973 | 49.7
[6 empty] | -0.043 | 25 | 18.963 | 0.496 | -18.963 | 49.6
[6 empty] | -0.043 | 25 | 18.953 | 0.496 | -18.953 | 49.6
[6 empty] | -0.043 | 25 | 18.943 | 0.496 | -18.943 | 49.6
[6 empty] | -0.043 | 25 | 18.933 | 0.496 | -18.933 | 49.6
[6 empty] | -0.043 | 25 | 18.923 | 0.495 | -18.923 | 49.5
[6 empty] | -0.043 | 25 | 18.913 | 0.495 | -18.913 | 49.5
[6 empty] | -0.043 | 25 | 18.903 | 0.495 | -18.903 | 49.5
[6 empty] | -0.043 | 25 | 18.893 | 0.495 | -18.893 | 49.5
[6 empty] | -0.043 | 25 | 18.883 | 0.494 | -18.883 | 49.4
[6 empty] | -0.043 | 25 | 18.873 | 0.494 | -18.873 | 49.4
[6 empty] | -0.043 | 25 | 18.863 | 0.494 | -18.863 | 49.4
[6 empty] | -0.043 | 25 | 18.853 | 0.494 | -18.853 | 49.4
[6 empty] | -0.044 | 25 | 18.843 | 0.493 | -18.843 | 49.3
[6 empty] | -0.044 | 25 | 18.833 | 0.493 | -18.833 | 49.3
[6 empty] | -0.044 | 25 | 18.823 | 0.493 | -18.823 | 49.3
[6 empty] | -0.044 | 25 | 18.813 | 0.493 | -18.813 | 49.3
[6 empty] | -0.044 | 25 | 18.803 | 0.492 | -18.803 | 49.2
[6 empty] | -0.044 | 25 | 18.793 | 0.492 | -18.793 | 49.2
[6 empty] | -0.044 | 25 | 18.783 | 0.492 | -18.783 | 49.2
[6 empty] | -0.044 | 25 | 18.773 | 0.492 | -18.773 | 49.2
[6 empty] | -0.044 | 25 | 18.763 | 0.491 | -18.763 | 49.1
[6 empty] | -0.044 | 25 | 18.753 | 0.491 | -18.753 | 49.1
[6 empty] | -0.044 | 25 | 18.743 | 0.491 | -18.743 | 49.1
[6 empty] | -0.044 | 25 | 18.733 | 0.491 | -18.733 | 49.1
[6 empty] | -0.044 | 25 | 18.723 | 0.49 | -18.723 | 49
[6 empty] | -0.044 | 25 | 18.713 | 0.49 | -18.713 | 49
[6 empty] | -0.044 | 25 | 18.703 | 0.49 | -18.703 | 49
[6 empty] | -0.044 | 25 | 18.693 | 0.489 | -18.693 | 48.9
[6 empty] | -0.044 | 25 | 18.683 | 0.489 | -18.683 | 48.9
[6 empty] | -0.044 | 25 | 18.673 | 0.489 | -18.673 | 48.9
[6 empty] | -0.044 | 25 | 18.663 | 0.489 | -18.663 | 48.9
[6 empty] | -0.044 | 25 | 18.653 | 0.488 | -18.653 | 48.8
[6 empty] | -0.044 | 25 | 18.643 | 0.488 | -18.643 | 48.8
[6 empty] | -0.044 | 25 | 18.633 | 0.488 | -18.633 | 48.8
[6 empty] | -0.044 | 25 | 18.623 | 0.488 | -18.623 | 48.8
[6 empty] | -0.044 | 25 | 18.613 | 0.487 | -18.613 | 48.7
[6 empty] | -0.044 | 25 | 18.603 | 0.487 | -18.603 | 48.7
[6 empty] | -0.044 | 25 | 18.593 | 0.487 | -18.593 | 48.7
[6 empty] | -0.044 | 25 | 18.583 | 0.487 | -18.583 | 48.7
[6 empty] | -0.044 | 25 | 18.573 | 0.486 | -18.573 | 48.6
[6 empty] | -0.044 | 25 | 18.563 | 0.486 | -18.563 | 48.6
[6 empty] | -0.044 | 25 | 18.553 | 0.486 | -18.553 | 48.6
[6 empty] | -0.044 | 25 | 18.543 | 0.486 | -18.543 | 48.6
[6 empty] | -0.044 | 25 | 18.533 | 0.485 | -18.533 | 48.5
[6 empty] | -0.044 | 25 | 18.523 | 0.485 | -18.523 | 48.5
[6 empty] | -0.044 | 25 | 18.513 | 0.485 | -18.513 | 48.5
[6 empty] | -0.044 | 25 | 18.503 | 0.485 | -18.503 | 48.5
[6 empty] | -0.044 | 25 | 18.493 | 0.484 | -18.493 | 48.4
[6 empty] | -0.044 | 25 | 18.483 | 0.484 | -18.483 | 48.4
[6 empty] | -0.044 | 25 | 18.473 | 0.484 | -18.473 | 48.4
[6 empty] | -0.044 | 25 | 18.463 | 0.484 | -18.463 | 48.4
[6 empty] | -0.044 | 25 | 18.453 | 0.483 | -18.453 | 48.3
[6 empty] | -0.044 | 25 | 18.443 | 0.483 | -18.443 | 48.3
[6 empty] | -0.044 | 25 | 18.433 | 0.483 | -18.433 | 48.3
[6 empty] | -0.044 | 25 | 18.423 | 0.483 | -18.423 | 48.3
[6 empty] | -0.044 | 25 | 18.413 | 0.482 | -18.413 | 48.2
[6 empty] | -0.044 | 25 | 18.403 | 0.482 | -18.403 | 48.2
[6 empty] | -0.044 | 25 | 18.393 | 0.482 | -18.393 | 48.2
[6 empty] | -0.044 | 25 | 18.383 | 0.482 | -18.383 | 48.2
[6 empty] | -0.044 | 25 | 18.373 | 0.481 | -18.373 | 48.1
[6 empty] | -0.044 | 25 | 18.363 | 0.481 | -18.363 | 48.1
[6 empty] | -0.044 | 25 | 18.353 | 0.481 | -18.353 | 48.1
[6 empty] | -0.044 | 25 | 18.343 | 0.481 | -18.343 | 48.1
[6 empty] | -0.044 | 25 | 18.333 | 0.48 | -18.333 | 48
[6 empty] | -0.044 | 25 | 18.323 | 0.48 | -18.323 | 48
[6 empty] | -0.044 | 25 | 18.313 | 0.48 | -18.313 | 48
[6 empty] | -0.044 | 25 | 18.303 | 0.48 | -18.303 | 48
[6 empty] | -0.044 | 25 | 18.293 | 0.479 | -18.293 | 47.9
[6 empty] | -0.044 | 25 | 18.283 | 0.479 | -18.283 | 47.9
[6 empty] | -0.044 | 25 | 18.273 | 0.479 | -18.273 | 47.9
[6 empty] | -0.044 | 25 | 18.263 | 0.479 | -18.263 | 47.9
[6 empty] | -0.044 | 25 | 18.253 | 0.478 | -18.253 | 47.8
[6 empty] | -0.044 | 25 | 18.243 | 0.478 | -18.243 | 47.8
[6 empty] | -0.044 | 25 | 18.233 | 0.478 | -18.233 | 47.8
[6 empty] | -0.044 | 25 | 18.223 | 0.478 | -18.223 | 47.8
[6 empty] | -0.044 | 25 | 18.213 | 0.477 | -18.213 | 47.7
[6 empty] | -0.044 | 25 | 18.203 | 0.477 | -18.203 | 47.7
[6 empty] | -0.044 | 25 | 18.193 | 0.477 | -18.193 | 47.7
[6 empty] | -0.044 | 25 | 18.183 | 0.476 | -18.183 | 47.6
[6 empty] | -0.044 | 25 | 18.173 | 0.476 | -18.173 | 47.6
[6 empty] | -0.044 | 25 | 18.163 | 0.476 | -18.163 | 47.6
[6 empty] | -0.044 | 25 | 18.153 | 0.476 | -18.153 | 47.6
[6 empty] | -0.045 | 25 | 18.143 | 0.475 | -18.143 | 47.5
[6 empty] | -0.045 | 25 | 18.133 | 0.475 | -18.133 | 47.5
[6 empty] | -0.045 | 25 | 18.123 | 0.475 | -18.123 | 47.5
[6 empty] | -0.045 | 25 | 18.113 | 0.475 | -18.113 | 47.5
[6 empty] | -0.045 | 25 | 18.103 | 0.474 | -18.103 | 47.4
[6 empty] | -0.045 | 25 | 18.093 | 0.474 | -18.093 | 47.4
[6 empty] | -0.045 | 25 | 18.083 | 0.474 | -18.083 | 47.4
[6 empty] | -0.045 | 25 | 18.073 | 0.474 | -18.073 | 47.4
[6 empty] | -0.045 | 25 | 18.063 | 0.473 | -18.063 | 47.3
[6 empty] | -0.045 | 25 | 18.053 | 0.473 | -18.053 | 47.3
[6 empty] | -0.045 | 25 | 18.043 | 0.473 | -18.043 | 47.3
[6 empty] | -0.045 | 25 | 18.033 | 0.473 | -18.033 | 47.3
[6 empty] | -0.045 | 25 | 18.023 | 0.472 | -18.023 | 47.2
[6 empty] | -0.045 | 25 | 18.013 | 0.472 | -18.013 | 47.2
[6 empty] | -0.045 | 25 | 18.003 | 0.472 | -18.003 | 47.2
[6 empty] | -0.045 | 25 | 17.993 | 0.472 | -17.993 | 47.2
[6 empty] | -0.045 | 25 | 17.983 | 0.471 | -17.983 | 47.1
[6 empty] | -0.045 | 25 | 17.973 | 0.471 | -17.973 | 47.1
[6 empty] | -0.045 | 25 | 17.963 | 0.471 | -17.963 | 47.1
[6 empty] | -0.045 | 25 | 17.953 | 0.471 | -17.953 | 47.1
[6 empty] | -0.045 | 25 | 17.943 | 0.471 | -17.943 | 47.1
[6 empty] | -0.045 | 25 | 17.933 | 0.47 | -17.933 | 47
[6 empty] | -0.045 | 25 | 17.923 | 0.47 | -17.923 | 47
[6 empty] | -0.045 | 25 | 17.913 | 0.47 | -17.913 | 47
[6 empty] | -0.045 | 25 | 17.903 | 0.47 | -17.903 | 47
[6 empty] | -0.045 | 25 | 17.893 | 0.469 | -17.893 | 46.9
[6 empty] | -0.045 | 25 | 17.883 | 0.469 | -17.883 | 46.9
[6 empty] | -0.045 | 25 | 17.873 | 0.469 | -17.873 | 46.9
[6 empty] | -0.045 | 25 | 17.863 | 0.469 | -17.863 | 46.9
[6 empty] | -0.045 | 25 | 17.853 | 0.468 | -17.853 | 46.8
[6 empty] | -0.045 | 25 | 17.843 | 0.468 | -17.843 | 46.8
[6 empty] | -0.045 | 25 | 17.833 | 0.468 | -17.833 | 46.8
[6 empty] | -0.045 | 25 | 17.823 | 0.468 | -17.823 | 46.8
[6 empty] | -0.045 | 25 | 17.813 | 0.467 | -17.813 | 46.7
[6 empty] | -0.045 | 25 | 17.803 | 0.467 | -17.803 | 46.7
[6 empty] | -0.045 | 25 | 17.793 | 0.467 | -17.793 | 46.7
[6 empty] | -0.045 | 25 | 17.783 | 0.467 | -17.783 | 46.7
[6 empty] | -0.045 | 25 | 17.773 | 0.466 | -17.773 | 46.6
[6 empty] | -0.045 | 25 | 17.763 | 0.466 | -17.763 | 46.6
[6 empty] | -0.045 | 25 | 17.753 | 0.466 | -17.753 | 46.6
[6 empty] | -0.045 | 25 | 17.743 | 0.466 | -17.743 | 46.6
[6 empty] | -0.045 | 25 | 17.733 | 0.465 | -17.733 | 46.5
[6 empty] | -0.045 | 25 | 17.723 | 0.465 | -17.723 | 46.5
[6 empty] | -0.045 | 25 | 17.713 | 0.465 | -17.713 | 46.5
[6 empty] | -0.045 | 25 | 17.703 | 0.465 | -17.703 | 46.5
[6 empty] | -0.045 | 25 | 17.693 | 0.464 | -17.693 | 46.4
[6 empty] | -0.045 | 25 | 17.683 | 0.464 | -17.683 | 46.4
[6 empty] | -0.045 | 25 | 17.673 | 0.464 | -17.673 | 46.4
[6 empty] | -0.045 | 25 | 17.663 | 0.464 | -17.663 | 46.4
[6 empty] | -0.045 | 25 | 17.653 | 0.463 | -17.653 | 46.3
[6 empty] | -0.045 | 25 | 17.643 | 0.463 | -17.643 | 46.3
[6 empty] | -0.045 | 25 | 17.633 | 0.463 | -17.633 | 46.3
[6 empty] | -0.045 | 25 | 17.623 | 0.463 | -17.623 | 46.3
[6 empty] | -0.045 | 25 | 17.613 | 0.462 | -17.613 | 46.2
[6 empty] | -0.045 | 25 | 17.603 | 0.462 | -17.603 | 46.2
[6 empty] | -0.045 | 25 | 17.593 | 0.462 | -17.593 | 46.2
[6 empty] | -0.045 | 25 | 17.583 | 0.462 | -17.583 | 46.2
[6 empty] | -0.045 | 25 | 17.574 | 0.461 | -17.574 | 46.1
[6 empty] | -0.045 | 25 | 17.564 | 0.461 | -17.564 | 46.1
[6 empty] | -0.045 | 25 | 17.554 | 0.461 | -17.554 | 46.1
[6 empty] | -0.045 | 25 | 17.544 | 0.461 | -17.544 | 46.1
[6 empty] | -0.045 | 25 | 17.534 | 0.46 | -17.534 | 46
[6 empty] | -0.045 | 25 | 17.524 | 0.46 | -17.524 | 46
[6 empty] | -0.045 | 25 | 17.514 | 0.46 | -17.514 | 46
[6 empty] | -0.045 | 25 | 17.504 | 0.46 | -17.504 | 46
[6 empty] | -0.045 | 25 | 17.494 | 0.459 | -17.494 | 45.9
[6 empty] | -0.045 | 25 | 17.484 | 0.459 | -17.484 | 45.9
[6 empty] | -0.045 | 25 | 17.474 | 0.459 | -17.474 | 45.9
[6 empty] | -0.045 | 25 | 17.464 | 0.459 | -17.464 | 45.9
[6 empty] | -0.045 | 25 | 17.454 | 0.458 | -17.454 | 45.8
[6 empty] | -0.046 | 25 | 17.444 | 0.458 | -17.444 | 45.8
[6 empty] | -0.046 | 25 | 17.434 | 0.458 | -17.434 | 45.8
[6 empty] | -0.046 | 25 | 17.424 | 0.458 | -17.424 | 45.8
[6 empty] | -0.046 | 25 | 17.414 | 0.457 | -17.414 | 45.7
[6 empty] | -0.046 | 25 | 17.404 | 0.457 | -17.404 | 45.7
[6 empty] | -0.046 | 25 | 17.394 | 0.457 | -17.394 | 45.7
[6 empty] | -0.046 | 25 | 17.384 | 0.457 | -17.384 | 45.7
[6 empty] | -0.046 | 25 | 17.374 | 0.456 | -17.374 | 45.6
[6 empty] | -0.046 | 25 | 17.364 | 0.456 | -17.364 | 45.6
[6 empty] | -0.046 | 25 | 17.354 | 0.456 | -17.354 | 45.6
[6 empty] | -0.046 | 25 | 17.344 | 0.456 | -17.344 | 45.6
[6 empty] | -0.046 | 25 | 17.334 | 0.456 | -17.334 | 45.6
[6 empty] | -0.046 | 25 | 17.324 | 0.455 | -17.324 | 45.5
[6 empty] | -0.046 | 25 | 17.314 | 0.455 | -17.314 | 45.5
[6 empty] | -0.046 | 25 | 17.304 | 0.455 | -17.304 | 45.5
[6 empty] | -0.046 | 25 | 17.294 | 0.455 | -17.294 | 45.5
[6 empty] | -0.046 | 25 | 17.284 | 0.454 | -17.284 | 45.4
[6 empty] | -0.046 | 25 | 17.274 | 0.454 | -17.274 | 45.4
[6 empty] | -0.046 | 25 | 17.264 | 0.454 | -17.264 | 45.4
[6 empty] | -0.046 | 25 | 17.254 | 0.454 | -17.254 | 45.4
[6 empty] | -0.046 | 25 | 17.244 | 0.453 | -17.244 | 45.3
[6 empty] | -0.046 | 25 | 17.234 | 0.453 | -17.234 | 45.3
[6 empty] | -0.046 | 25 | 17.224 | 0.453 | -17.224 | 45.3
[6 empty] | -0.046 | 25 | 17.214 | 0.453 | -17.214 | 45.3
[6 empty] | -0.046 | 25 | 17.204 | 0.452 | -17.204 | 45.2
[6 empty] | -0.046 | 25 | 17.194 | 0.452 | -17.194 | 45.2
[6 empty] | -0.046 | 25 | 17.184 | 0.452 | -17.184 | 45.2
[6 empty] | -0.046 | 25 | 17.174 | 0.452 | -17.174 | 45.2
[6 empty] | -0.046 | 25 | 17.164 | 0.451 | -17.164 | 45.1
[6 empty] | -0.046 | 25 | 17.154 | 0.451 | -17.154 | 45.1
[6 empty] | -0.046 | 25 | 17.144 | 0.451 | -17.144 | 45.1
[6 empty] | -0.046 | 25 | 17.134 | 0.451 | -17.134 | 45.1
[6 empty] | -0.046 | 25 | 17.124 | 0.45 | -17.124 | 45
[6 empty] | -0.046 | 25 | 17.114 | 0.45 | -17.114 | 45
[6 empty] | -0.046 | 25 | 17.104 | 0.45 | -17.104 | 45
[6 empty] | -0.046 | 25 | 17.094 | 0.45 | -17.094 | 45
[6 empty] | -0.046 | 25 | 17.084 | 0.449 | -17.084 | 44.9
[6 empty] | -0.046 | 25 | 17.074 | 0.449 | -17.074 | 44.9
[6 empty] | -0.046 | 25 | 17.064 | 0.449 | -17.064 | 44.9
[6 empty] | -0.046 | 25 | 17.054 | 0.449 | -17.054 | 44.9
[6 empty] | -0.046 | 25 | 17.044 | 0.448 | -17.044 | 44.8
[6 empty] | -0.046 | 25 | 17.034 | 0.448 | -17.034 | 44.8
[6 empty] | -0.046 | 25 | 17.024 | 0.448 | -17.024 | 44.8
[6 empty] | -0.046 | 25 | 17.014 | 0.448 | -17.014 | 44.8
[6 empty] | -0.046 | 25 | 17.004 | 0.447 | -17.004 | 44.7
[6 empty] | -0.046 | 25 | 16.994 | 0.447 | -16.994 | 44.7
[6 empty] | -0.046 | 25 | 16.984 | 0.447 | -16.984 | 44.7
[6 empty] | -0.046 | 25 | 16.974 | 0.447 | -16.974 | 44.7
[6 empty] | -0.046 | 25 | 16.964 | 0.446 | -16.964 | 44.6
[6 empty] | -0.046 | 25 | 16.954 | 0.446 | -16.954 | 44.6
[6 empty] | -0.046 | 25 | 16.944 | 0.446 | -16.944 | 44.6
[6 empty] | -0.046 | 25 | 16.934 | 0.446 | -16.934 | 44.6
[6 empty] | -0.046 | 25 | 16.924 | 0.445 | -16.924 | 44.5
[6 empty] | -0.046 | 25 | 16.914 | 0.445 | -16.914 | 44.5
[6 empty] | -0.046 | 25 | 16.904 | 0.445 | -16.904 | 44.5
[6 empty] | -0.046 | 25 | 16.894 | 0.445 | -16.894 | 44.5
[6 empty] | -0.046 | 25 | 16.884 | 0.444 | -16.884 | 44.4
[6 empty] | -0.046 | 25 | 16.874 | 0.444 | -16.874 | 44.4
[6 empty] | -0.046 | 25 | 16.864 | 0.444 | -16.864 | 44.4
[6 empty] | -0.046 | 25 | 16.854 | 0.444 | -16.854 | 44.4
[6 empty] | -0.046 | 25 | 16.844 | 0.443 | -16.844 | 44.3
[6 empty] | -0.046 | 25 | 16.834 | 0.443 | -16.834 | 44.3
[6 empty] | -0.046 | 25 | 16.824 | 0.443 | -16.824 | 44.3
[6 empty] | -0.046 | 25 | 16.814 | 0.443 | -16.814 | 44.3
[6 empty] | -0.046 | 25 | 16.804 | 0.442 | -16.804 | 44.2
[6 empty] | -0.046 | 25 | 16.794 | 0.442 | -16.794 | 44.2
[6 empty] | -0.046 | 25 | 16.784 | 0.442 | -16.784 | 44.2
[6 empty] | -0.046 | 25 | 16.774 | 0.442 | -16.774 | 44.2
[6 empty] | -0.046 | 25 | 16.764 | 0.441 | -16.764 | 44.1
[6 empty] | -0.046 | 25 | 16.754 | 0.441 | -16.754 | 44.1
[6 empty] | -0.047 | 25 | 16.744 | 0.441 | -16.744 | 44.1
[6 empty] | -0.047 | 25 | 16.734 | 0.441 | -16.734 | 44.1
[6 empty] | -0.047 | 25 | 16.724 | 0.44 | -16.724 | 44
[6 empty] | -0.047 | 25 | 16.714 | 0.44 | -16.714 | 44
[6 empty] | -0.047 | 25 | 16.704 | 0.44 | -16.704 | 44
[6 empty] | -0.047 | 25 | 16.694 | 0.44 | -16.694 | 44
[6 empty] | -0.047 | 25 | 16.684 | 0.44 | -16.684 | 44
[6 empty] | -0.047 | 25 | 16.674 | 0.439 | -16.674 | 43.9
[6 empty] | -0.047 | 25 | 16.664 | 0.439 | -16.664 | 43.9
[6 empty] | -0.047 | 25 | 16.654 | 0.439 | -16.654 | 43.9
[6 empty] | -0.047 | 25 | 16.644 | 0.439 | -16.644 | 43.9
[6 empty] | -0.047 | 25 | 16.634 | 0.438 | -16.634 | 43.8
[6 empty] | -0.047 | 25 | 16.624 | 0.438 | -16.624 | 43.8
[6 empty] | -0.047 | 25 | 16.614 | 0.438 | -16.614 | 43.8
[6 empty] | -0.047 | 25 | 16.604 | 0.438 | -16.604 | 43.8
[6 empty] | -0.047 | 25 | 16.594 | 0.437 | -16.594 | 43.7
[6 empty] | -0.047 | 25 | 16.584 | 0.437 | -16.584 | 43.7
[6 empty] | -0.047 | 25 | 16.574 | 0.437 | -16.574 | 43.7
[6 empty] | -0.047 | 25 | 16.564 | 0.437 | -16.564 | 43.7
[6 empty] | -0.047 | 25 | 16.554 | 0.436 | -16.554 | 43.6
[6 empty] | -0.047 | 25 | 16.544 | 0.436 | -16.544 | 43.6
[6 empty] | -0.047 | 25 | 16.534 | 0.436 | -16.534 | 43.6
[6 empty] | -0.047 | 25 | 16.524 | 0.436 | -16.524 | 43.6
[6 empty] | -0.047 | 25 | 16.514 | 0.435 | -16.514 | 43.5
[6 empty] | -0.047 | 25 | 16.504 | 0.435 | -16.504 | 43.5
[6 empty] | -0.047 | 25 | 16.494 | 0.435 | -16.494 | 43.5
[6 empty] | -0.047 | 25 | 16.484 | 0.435 | -16.484 | 43.5
[6 empty] | -0.047 | 25 | 16.474 | 0.434 | -16.474 | 43.4
[6 empty] | -0.047 | 25 | 16.464 | 0.434 | -16.464 | 43.4
[6 empty] | -0.047 | 25 | 16.454 | 0.434 | -16.454 | 43.4
[6 empty] | -0.047 | 25 | 16.444 | 0.434 | -16.444 | 43.4
[6 empty] | -0.047 | 25 | 16.434 | 0.434 | -16.434 | 43.4
[6 empty] | -0.047 | 25 | 16.424 | 0.433 | -16.424 | 43.3
[6 empty] | -0.047 | 25 | 16.414 | 0.433 | -16.414 | 43.3
[6 empty] | -0.047 | 25 | 16.404 | 0.433 | -16.404 | 43.3
[6 empty] | -0.047 | 25 | 16.394 | 0.433 | -16.394 | 43.3
[6 empty] | -0.047 | 25 | 16.384 | 0.432 | -16.384 | 43.2
[6 empty] | -0.047 | 25 | 16.374 | 0.432 | -16.374 | 43.2
[6 empty] | -0.047 | 25 | 16.364 | 0.432 | -16.364 | 43.2
[6 empty] | -0.047 | 25 | 16.354 | 0.432 | -16.354 | 43.2
[6 empty] | -0.047 | 25 | 16.344 | 0.431 | -16.344 | 43.1
[6 empty] | -0.047 | 25 | 16.334 | 0.431 | -16.334 | 43.1
[6 empty] | -0.047 | 25 | 16.324 | 0.431 | -16.324 | 43.1
[6 empty] | -0.047 | 25 | 16.314 | 0.431 | -16.314 | 43.1
[6 empty] | -0.047 | 25 | 16.304 | 0.431 | -16.304 | 43.1
[6 empty] | -0.047 | 25 | 16.294 | 0.43 | -16.294 | 43
[6 empty] | -0.047 | 25 | 16.284 | 0.43 | -16.284 | 43
[6 empty] | -0.047 | 25 | 16.274 | 0.43 | -16.274 | 43
[6 empty] | -0.047 | 25 | 16.264 | 0.43 | -16.264 | 43
[6 empty] | -0.047 | 25 | 16.254 | 0.429 | -16.254 | 42.9
[6 empty] | -0.047 | 25 | 16.244 | 0.429 | -16.244 | 42.9
[6 empty] | -0.047 | 25 | 16.234 | 0.429 | -16.234 | 42.9
[6 empty] | -0.047 | 25 | 16.224 | 0.429 | -16.224 | 42.9
[6 empty] | -0.047 | 25 | 16.214 | 0.428 | -16.214 | 42.8
[6 empty] | -0.047 | 25 | 16.204 | 0.428 | -16.204 | 42.8
[6 empty] | -0.047 | 25 | 16.194 | 0.428 | -16.194 | 42.8
[6 empty] | -0.047 | 25 | 16.184 | 0.428 | -16.184 | 42.8
[6 empty] | -0.047 | 25 | 16.174 | 0.427 | -16.174 | 42.7
[6 empty] | -0.047 | 25 | 16.164 | 0.427 | -16.164 | 42.7
[6 empty] | -0.047 | 25 | 16.154 | 0.427 | -16.154 | 42.7
[6 empty] | -0.047 | 25 | 16.144 | 0.427 | -16.144 | 42.7
[6 empty] | -0.047 | 25 | 16.134 | 0.427 | -16.134 | 42.7
[6 empty] | -0.047 | 25 | 16.124 | 0.426 | -16.124 | 42.6
[6 empty] | -0.047 | 25 | 16.114 | 0.426 | -16.114 | 42.6
[6 empty] | -0.047 | 25 | 16.104 | 0.426 | -16.104 | 42.6
[6 empty] | -0.047 | 25 | 16.094 | 0.426 | -16.094 | 42.6
[6 empty] | -0.047 | 25 | 16.084 | 0.425 | -16.084 | 42.5
[6 empty] | -0.047 | 25 | 16.074 | 0.425 | -16.074 | 42.5
[6 empty] | -0.047 | 25 | 16.064 | 0.425 | -16.064 | 42.5
[6 empty] | -0.047 | 25 | 16.054 | 0.425 | -16.054 | 42.5
[6 empty] | -0.048 | 25 | 16.044 | 0.424 | -16.044 | 42.4
[6 empty] | -0.048 | 25 | 16.034 | 0.424 | -16.034 | 42.4
[6 empty] | -0.048 | 25 | 16.024 | 0.424 | -16.024 | 42.4
[6 empty] | -0.048 | 25 | 16.014 | 0.424 | -16.014 | 42.4
[6 empty] | -0.048 | 25 | 16.004 | 0.424 | -16.004 | 42.4
[6 empty] | -0.048 | 25 | 15.994 | 0.423 | -15.994 | 42.3
[6 empty] | -0.048 | 25 | 15.984 | 0.423 | -15.984 | 42.3
[6 empty] | -0.048 | 25 | 15.974 | 0.423 | -15.974 | 42.3
[6 empty] | -0.048 | 25 | 15.964 | 0.423 | -15.964 | 42.3
[6 empty] | -0.048 | 25 | 15.954 | 0.422 | -15.954 | 42.2
[6 empty] | -0.048 | 25 | 15.944 | 0.422 | -15.944 | 42.2
[6 empty] | -0.048 | 25 | 15.934 | 0.422 | -15.934 | 42.2
[6 empty] | -0.048 | 25 | 15.924 | 0.422 | -15.924 | 42.2
[6 empty] | -0.048 | 25 | 15.914 | 0.421 | -15.914 | 42.1
[6 empty] | -0.048 | 25 | 15.904 | 0.421 | -15.904 | 42.1
[6 empty] | -0.048 | 25 | 15.894 | 0.421 | -15.894 | 42.1
[6 empty] | -0.048 | 25 | 15.884 | 0.421 | -15.884 | 42.1
[6 empty] | -0.048 | 25 | 15.874 | 0.421 | -15.874 | 42.1
[6 empty] | -0.048 | 25 | 15.864 | 0.42 | -15.864 | 42
[6 empty] | -0.048 | 25 | 15.854 | 0.42 | -15.854 | 42
[6 empty] | -0.048 | 25 | 15.844 | 0.42 | -15.844 | 42
[6 empty] | -0.048 | 25 | 15.834 | 0.42 | -15.834 | 42
[6 empty] | -0.048 | 25 | 15.824 | 0.419 | -15.824 | 41.9
[6 empty] | -0.048 | 25 | 15.814 | 0.419 | -15.814 | 41.9
[6 empty] | -0.048 | 25 | 15.804 | 0.419 | -15.804 | 41.9
[6 empty] | -0.048 | 25 | 15.794 | 0.419 | -15.794 | 41.9
[6 empty] | -0.048 | 25 | 15.785 | 0.418 | -15.785 | 41.8
[6 empty] | -0.048 | 25 | 15.775 | 0.418 | -15.775 | 41.8
[6 empty] | -0.048 | 25 | 15.765 | 0.418 | -15.765 | 41.8
[6 empty] | -0.048 | 25 | 15.755 | 0.418 | -15.755 | 41.8
[6 empty] | -0.048 | 25 | 15.745 | 0.418 | -15.745 | 41.8
[6 empty] | -0.048 | 25 | 15.735 | 0.417 | -15.735 | 41.7
[6 empty] | -0.048 | 25 | 15.725 | 0.417 | -15.725 | 41.7
[6 empty] | -0.048 | 25 | 15.715 | 0.417 | -15.715 | 41.7
[6 empty] | -0.048 | 25 | 15.705 | 0.417 | -15.705 | 41.7
[6 empty] | -0.048 | 25 | 15.695 | 0.416 | -15.695 | 41.6
[6 empty] | -0.048 | 25 | 15.685 | 0.416 | -15.685 | 41.6
[6 empty] | -0.048 | 25 | 15.675 | 0.416 | -15.675 | 41.6
[6 empty] | -0.048 | 25 | 15.665 | 0.416 | -15.665 | 41.6
[6 empty] | -0.048 | 25 | 15.655 | 0.416 | -15.655 | 41.6
[6 empty] | -0.048 | 25 | 15.645 | 0.415 | -15.645 | 41.5
[6 empty] | -0.048 | 25 | 15.635 | 0.415 | -15.635 | 41.5
[6 empty] | -0.048 | 25 | 15.625 | 0.415 | -15.625 | 41.5
[6 empty] | -0.048 | 25 | 15.615 | 0.415 | -15.615 | 41.5
[6 empty] | -0.048 | 25 | 15.605 | 0.414 | -15.605 | 41.4
[6 empty] | -0.048 | 25 | 15.595 | 0.414 | -15.595 | 41.4
[6 empty] | -0.048 | 25 | 15.585 | 0.414 | -15.585 | 41.4
[6 empty] | -0.048 | 25 | 15.575 | 0.414 | -15.575 | 41.4
[6 empty] | -0.048 | 25 | 15.565 | 0.413 | -15.565 | 41.3
[6 empty] | -0.048 | 25 | 15.555 | 0.413 | -15.555 | 41.3
[6 empty] | -0.048 | 25 | 15.545 | 0.413 | -15.545 | 41.3
[6 empty] | -0.048 | 25 | 15.535 | 0.413 | -15.535 | 41.3
[6 empty] | -0.048 | 25 | 15.525 | 0.413 | -15.525 | 41.3
[6 empty] | -0.048 | 25 | 15.515 | 0.412 | -15.515 | 41.2
[6 empty] | -0.048 | 25 | 15.505 | 0.412 | -15.505 | 41.2
[6 empty] | -0.048 | 25 | 15.495 | 0.412 | -15.495 | 41.2
[6 empty] | -0.048 | 25 | 15.485 | 0.412 | -15.485 | 41.2
[6 empty] | -0.048 | 25 | 15.475 | 0.411 | -15.475 | 41.1
[6 empty] | -0.048 | 25 | 15.465 | 0.411 | -15.465 | 41.1
[6 empty] | -0.048 | 25 | 15.455 | 0.411 | -15.455 | 41.1
[6 empty] | -0.048 | 25 | 15.445 | 0.411 | -15.445 | 41.1
[6 empty] | -0.048 | 25 | 15.435 | 0.41 | -15.435 | 41
[6 empty] | -0.048 | 25 | 15.425 | 0.41 | -15.425 | 41
[6 empty] | -0.048 | 25 | 15.415 | 0.41 | -15.415 | 41
[6 empty] | -0.048 | 25 | 15.405 | 0.41 | -15.405 | 41
[6 empty] | -0.048 | 25 | 15.395 | 0.41 | -15.395 | 41
[6 empty] | -0.048 | 25 | 15.385 | 0.409 | -15.385 | 40.9
[6 empty] | -0.048 | 25 | 15.375 | 0.409 | -15.375 | 40.9
[6 empty] | -0.048 | 25 | 15.365 | 0.409 | -15.365 | 40.9
[6 empty] | -0.048 | 25 | 15.355 | 0.409 | -15.355 | 40.9
[6 empty] | -0.049 | 25 | 15.345 | 0.408 | -15.345 | 40.8
[6 empty] | -0.049 | 25 | 15.335 | 0.408 | -15.335 | 40.8
[6 empty] | -0.049 | 25 | 15.325 | 0.408 | -15.325 | 40.8
[6 empty] | -0.049 | 25 | 15.315 | 0.408 | -15.315 | 40.8
[6 empty] | -0.049 | 25 | 15.305 | 0.408 | -15.305 | 40.8
[6 empty] | -0.049 | 25 | 15.295 | 0.407 | -15.295 | 40.7
[6 empty] | -0.049 | 25 | 15.285 | 0.407 | -15.285 | 40.7
[6 empty] | -0.049 | 25 | 15.275 | 0.407 | -15.275 | 40.7
[6 empty] | -0.049 | 25 | 15.265 | 0.407 | -15.265 | 40.7
[6 empty] | -0.049 | 25 | 15.255 | 0.406 | -15.255 | 40.6
[6 empty] | -0.049 | 25 | 15.245 | 0.406 | -15.245 | 40.6
[6 empty] | -0.049 | 25 | 15.235 | 0.406 | -15.235 | 40.6
[6 empty] | -0.049 | 25 | 15.225 | 0.406 | -15.225 | 40.6
[6 empty] | -0.049 | 25 | 15.215 | 0.405 | -15.215 | 40.5
[6 empty] | -0.049 | 25 | 15.205 | 0.405 | -15.205 | 40.5
[6 empty] | -0.049 | 25 | 15.195 | 0.405 | -15.195 | 40.5
[6 empty] | -0.049 | 25 | 15.185 | 0.405 | -15.185 | 40.5
[6 empty] | -0.049 | 25 | 15.175 | 0.405 | -15.175 | 40.5
[6 empty] | -0.049 | 25 | 15.165 | 0.404 | -15.165 | 40.4
[6 empty] | -0.049 | 25 | 15.155 | 0.404 | -15.155 | 40.4
[6 empty] | -0.049 | 25 | 15.145 | 0.404 | -15.145 | 40.4
[6 empty] | -0.049 | 25 | 15.135 | 0.404 | -15.135 | 40.4
[6 empty] | -0.049 | 25 | 15.125 | 0.403 | -15.125 | 40.3
[6 empty] | -0.049 | 25 | 15.115 | 0.403 | -15.115 | 40.3
[6 empty] | -0.049 | 25 | 15.105 | 0.403 | -15.105 | 40.3
[6 empty] | -0.049 | 25 | 15.095 | 0.403 | -15.095 | 40.3
[6 empty] | -0.049 | 25 | 15.085 | 0.403 | -15.085 | 40.3
[6 empty] | -0.049 | 25 | 15.075 | 0.402 | -15.075 | 40.2
[6 empty] | -0.049 | 25 | 15.065 | 0.402 | -15.065 | 40.2
[6 empty] | -0.049 | 25 | 15.055 | 0.402 | -15.055 | 40.2
[6 empty] | -0.049 | 25 | 15.045 | 0.402 | -15.045 | 40.2
[6 empty] | -0.049 | 25 | 15.035 | 0.401 | -15.035 | 40.1
[6 empty] | -0.049 | 25 | 15.025 | 0.401 | -15.025 | 40.1
[6 empty] | -0.049 | 25 | 15.015 | 0.401 | -15.015 | 40.1
[6 empty] | -0.049 | 25 | 15.005 | 0.401 | -15.005 | 40.1
[6 empty] | -0.049 | 25 | 14.995 | 0.401 | -14.995 | 40.1
[6 empty] | -0.049 | 25 | 14.985 | 0.4 | -14.985 | 40
[6 empty] | -0.049 | 25 | 14.975 | 0.4 | -14.975 | 40
[6 empty] | -0.049 | 25 | 14.965 | 0.4 | -14.965 | 40
[6 empty] | -0.049 | 25 | 14.955 | 0.4 | -14.955 | 40
[6 empty] | -0.049 | 25 | 14.945 | 0.399 | -14.945 | 39.9
[6 empty] | -0.049 | 25 | 14.935 | 0.399 | -14.935 | 39.9
[6 empty] | -0.049 | 25 | 14.925 | 0.399 | -14.925 | 39.9
[6 empty] | -0.049 | 25 | 14.915 | 0.399 | -14.915 | 39.9
[6 empty] | -0.049 | 25 | 14.905 | 0.399 | -14.905 | 39.9
[6 empty] | -0.049 | 25 | 14.895 | 0.398 | -14.895 | 39.8
[6 empty] | -0.049 | 25 | 14.885 | 0.398 | -14.885 | 39.8
[6 empty] | -0.049 | 25 | 14.875 | 0.398 | -14.875 | 39.8
[6 empty] | -0.049 | 25 | 14.865 | 0.398 | -14.865 | 39.8
[6 empty] | -0.049 | 25 | 14.855 | 0.397 | -14.855 | 39.7
[6 empty] | -0.049 | 25 | 14.845 | 0.397 | -14.845 | 39.7
[6 empty] | -0.049 | 25 | 14.835 | 0.397 | -14.835 | 39.7
[6 empty] | -0.049 | 25 | 14.825 | 0.397 | -14.825 | 39.7
[6 empty] | -0.049 | 25 | 14.815 | 0.397 | -14.815 | 39.7
[6 empty] | -0.049 | 25 | 14.805 | 0.396 | -14.805 | 39.6
[6 empty] | -0.049 | 25 | 14.795 | 0.396 | -14.795 | 39.6
[6 empty] | -0.049 | 25 | 14.785 | 0.396 | -14.785 | 39.6
[6 empty] | -0.049 | 25 | 14.775 | 0.396 | -14.775 | 39.6
[6 empty] | -0.049 | 25 | 14.765 | 0.395 | -14.765 | 39.5
[6 empty] | -0.049 | 25 | 14.755 | 0.395 | -14.755 | 39.5
[6 empty] | -0.049 | 25 | 14.745 | 0.395 | -14.745 | 39.5
[6 empty] | -0.049 | 25 | 14.735 | 0.395 | -14.735 | 39.5
[6 empty] | -0.049 | 25 | 14.725 | 0.395 | -14.725 | 39.5
[6 empty] | -0.049 | 25 | 14.715 | 0.394 | -14.715 | 39.4
[6 empty] | -0.049 | 25 | 14.705 | 0.394 | -14.705 | 39.4
[6 empty] | -0.049 | 25 | 14.695 | 0.394 | -14.695 | 39.4
[6 empty] | -0.049 | 25 | 14.685 | 0.394 | -14.685 | 39.4
[6 empty] | -0.049 | 25 | 14.675 | 0.393 | -14.675 | 39.3
[6 empty] | -0.049 | 25 | 14.665 | 0.393 | -14.665 | 39.3
[6 empty] | -0.049 | 25 | 14.655 | 0.393 | -14.655 | 39.3
[6 empty] | -0.05 | 25 | 14.645 | 0.393 | -14.645 | 39.3
[6 empty] | -0.05 | 25 | 14.635 | 0.393 | -14.635 | 39.3
[6 empty] | -0.05 | 25 | 14.625 | 0.392 | -14.625 | 39.2
[6 empty] | -0.05 | 25 | 14.615 | 0.392 | -14.615 | 39.2
[6 empty] | -0.05 | 25 | 14.605 | 0.392 | -14.605 | 39.2
[6 empty] | -0.05 | 25 | 14.595 | 0.392 | -14.595 | 39.2
[6 empty] | -0.05 | 25 | 14.585 | 0.392 | -14.585 | 39.2
[6 empty] | -0.05 | 25 | 14.575 | 0.391 | -14.575 | 39.1
[6 empty] | -0.05 | 25 | 14.565 | 0.391 | -14.565 | 39.1
[6 empty] | -0.05 | 25 | 14.555 | 0.391 | -14.555 | 39.1
[6 empty] | -0.05 | 25 | 14.545 | 0.391 | -14.545 | 39.1
[6 empty] | -0.05 | 25 | 14.535 | 0.39 | -14.535 | 39
[6 empty] | -0.05 | 25 | 14.525 | 0.39 | -14.525 | 39
[6 empty] | -0.05 | 25 | 14.515 | 0.39 | -14.515 | 39
[6 empty] | -0.05 | 25 | 14.505 | 0.39 | -14.505 | 39
[6 empty] | -0.05 | 25 | 14.495 | 0.39 | -14.495 | 39
[6 empty] | -0.05 | 25 | 14.485 | 0.389 | -14.485 | 38.9
[6 empty] | -0.05 | 25 | 14.475 | 0.389 | -14.475 | 38.9
[6 empty] | -0.05 | 25 | 14.465 | 0.389 | -14.465 | 38.9
[6 empty] | -0.05 | 25 | 14.455 | 0.389 | -14.455 | 38.9
[6 empty] | -0.05 | 25 | 14.445 | 0.388 | -14.445 | 38.8
[6 empty] | -0.05 | 25 | 14.435 | 0.388 | -14.435 | 38.8
[6 empty] | -0.05 | 25 | 14.425 | 0.388 | -14.425 | 38.8
[6 empty] | -0.05 | 25 | 14.415 | 0.388 | -14.415 | 38.8
[6 empty] | -0.05 | 25 | 14.405 | 0.388 | -14.405 | 38.8
[6 empty] | -0.05 | 25 | 14.395 | 0.387 | -14.395 | 38.7
[6 empty] | -0.05 | 25 | 14.385 | 0.387 | -14.385 | 38.7
[6 empty] | -0.05 | 25 | 14.375 | 0.387 | -14.375 | 38.7
[6 empty] | -0.05 | 25 | 14.365 | 0.387 | -14.365 | 38.7
[6 empty] | -0.05 | 25 | 14.355 | 0.387 | -14.355 | 38.7
[6 empty] | -0.05 | 25 | 14.345 | 0.386 | -14.345 | 38.6
[6 empty] | -0.05 | 25 | 14.335 | 0.386 | -14.335 | 38.6
[6 empty] | -0.05 | 25 | 14.325 | 0.386 | -14.325 | 38.6
[6 empty] | -0.05 | 25 | 14.315 | 0.386 | -14.315 | 38.6
[6 empty] | -0.05 | 25 | 14.305 | 0.385 | -14.305 | 38.5
[6 empty] | -0.05 | 25 | 14.295 | 0.385 | -14.295 | 38.5
[6 empty] | -0.05 | 25 | 14.285 | 0.385 | -14.285 | 38.5
[6 empty] | -0.05 | 25 | 14.275 | 0.385 | -14.275 | 38.5
[6 empty] | -0.05 | 25 | 14.265 | 0.385 | -14.265 | 38.5
[6 empty] | -0.05 | 25 | 14.255 | 0.384 | -14.255 | 38.4
[6 empty] | -0.05 | 25 | 14.245 | 0.384 | -14.245 | 38.4
[6 empty] | -0.05 | 25 | 14.235 | 0.384 | -14.235 | 38.4
[6 empty] | -0.05 | 25 | 14.225 | 0.384 | -14.225 | 38.4
[6 empty] | -0.05 | 25 | 14.215 | 0.384 | -14.215 | 38.4
[6 empty] | -0.05 | 25 | 14.205 | 0.383 | -14.205 | 38.3
[6 empty] | -0.05 | 25 | 14.195 | 0.383 | -14.195 | 38.3
[6 empty] | -0.05 | 25 | 14.185 | 0.383 | -14.185 | 38.3
[6 empty] | -0.05 | 25 | 14.175 | 0.383 | -14.175 | 38.3
[6 empty] | -0.05 | 25 | 14.165 | 0.382 | -14.165 | 38.2
[6 empty] | -0.05 | 25 | 14.155 | 0.382 | -14.155 | 38.2
[6 empty] | -0.05 | 25 | 14.145 | 0.382 | -14.145 | 38.2
[6 empty] | -0.05 | 25 | 14.135 | 0.382 | -14.135 | 38.2
[6 empty] | -0.05 | 25 | 14.125 | 0.382 | -14.125 | 38.2
[6 empty] | -0.05 | 25 | 14.115 | 0.381 | -14.115 | 38.1
[6 empty] | -0.05 | 25 | 14.105 | 0.381 | -14.105 | 38.1
[6 empty] | -0.05 | 25 | 14.095 | 0.381 | -14.095 | 38.1
[6 empty] | -0.05 | 25 | 14.085 | 0.381 | -14.085 | 38.1
[6 empty] | -0.05 | 25 | 14.075 | 0.381 | -14.075 | 38.1
[6 empty] | -0.05 | 25 | 14.065 | 0.38 | -14.065 | 38
[6 empty] | -0.05 | 25 | 14.055 | 0.38 | -14.055 | 38
[6 empty] | -0.05 | 25 | 14.045 | 0.38 | -14.045 | 38
[6 empty] | -0.05 | 25 | 14.035 | 0.38 | -14.035 | 38
[6 empty] | -0.05 | 25 | 14.025 | 0.38 | -14.025 | 38
[6 empty] | -0.05 | 25 | 14.015 | 0.379 | -14.015 | 37.9
[6 empty] | -0.05 | 25 | 14.005 | 0.379 | -14.005 | 37.9
[6 empty] | -0.05 | 25 | 13.996 | 0.379 | -13.996 | 37.9
[6 empty] | -0.05 | 25 | 13.986 | 0.379 | -13.986 | 37.9
[6 empty] | -0.05 | 25 | 13.976 | 0.378 | -13.976 | 37.8
[6 empty] | -0.05 | 25 | 13.966 | 0.378 | -13.966 | 37.8
[6 empty] | -0.05 | 25 | 13.956 | 0.378 | -13.956 | 37.8
[6 empty] | -0.051 | 25 | 13.946 | 0.378 | -13.946 | 37.8
[6 empty] | -0.051 | 25 | 13.936 | 0.378 | -13.936 | 37.8
[6 empty] | -0.051 | 25 | 13.926 | 0.377 | -13.926 | 37.7
[6 empty] | -0.051 | 25 | 13.916 | 0.377 | -13.916 | 37.7
[6 empty] | -0.051 | 25 | 13.906 | 0.377 | -13.906 | 37.7
[6 empty] | -0.051 | 25 | 13.896 | 0.377 | -13.896 | 37.7
[6 empty] | -0.051 | 25 | 13.886 | 0.377 | -13.886 | 37.7
[6 empty] | -0.051 | 25 | 13.876 | 0.376 | -13.876 | 37.6
[6 empty] | -0.051 | 25 | 13.866 | 0.376 | -13.866 | 37.6
[6 empty] | -0.051 | 25 | 13.856 | 0.376 | -13.856 | 37.6
[6 empty] | -0.051 | 25 | 13.846 | 0.376 | -13.846 | 37.6
[6 empty] | -0.051 | 25 | 13.836 | 0.376 | -13.836 | 37.6
[6 empty] | -0.051 | 25 | 13.826 | 0.375 | -13.826 | 37.5
[6 empty] | -0.051 | 25 | 13.816 | 0.375 | -13.816 | 37.5
[6 empty] | -0.051 | 25 | 13.806 | 0.375 | -13.806 | 37.5
[6 empty] | -0.051 | 25 | 13.796 | 0.375 | -13.796 | 37.5
[6 empty] | -0.051 | 25 | 13.786 | 0.375 | -13.786 | 37.5
[6 empty] | -0.051 | 25 | 13.776 | 0.374 | -13.776 | 37.4
[6 empty] | -0.051 | 25 | 13.766 | 0.374 | -13.766 | 37.4
[6 empty] | -0.051 | 25 | 13.756 | 0.374 | -13.756 | 37.4
[6 empty] | -0.051 | 25 | 13.746 | 0.374 | -13.746 | 37.4
[6 empty] | -0.051 | 25 | 13.736 | 0.373 | -13.736 | 37.3
[6 empty] | -0.051 | 25 | 13.726 | 0.373 | -13.726 | 37.3
[6 empty] | -0.051 | 25 | 13.716 | 0.373 | -13.716 | 37.3
[6 empty] | -0.051 | 25 | 13.706 | 0.373 | -13.706 | 37.3
[6 empty] | -0.051 | 25 | 13.696 | 0.373 | -13.696 | 37.3
[6 empty] | -0.051 | 25 | 13.686 | 0.372 | -13.686 | 37.2
[6 empty] | -0.051 | 25 | 13.676 | 0.372 | -13.676 | 37.2
[6 empty] | -0.051 | 25 | 13.666 | 0.372 | -13.666 | 37.2
[6 empty] | -0.051 | 25 | 13.656 | 0.372 | -13.656 | 37.2
[6 empty] | -0.051 | 25 | 13.646 | 0.372 | -13.646 | 37.2
[6 empty] | -0.051 | 25 | 13.636 | 0.371 | -13.636 | 37.1
[6 empty] | -0.051 | 25 | 13.626 | 0.371 | -13.626 | 37.1
[6 empty] | -0.051 | 25 | 13.616 | 0.371 | -13.616 | 37.1
[6 empty] | -0.051 | 25 | 13.606 | 0.371 | -13.606 | 37.1
[6 empty] | -0.051 | 25 | 13.596 | 0.371 | -13.596 | 37.1
[6 empty] | -0.051 | 25 | 13.586 | 0.37 | -13.586 | 37
[6 empty] | -0.051 | 25 | 13.576 | 0.37 | -13.576 | 37
[6 empty] | -0.051 | 25 | 13.566 | 0.37 | -13.566 | 37
[6 empty] | -0.051 | 25 | 13.556 | 0.37 | -13.556 | 37
[6 empty] | -0.051 | 25 | 13.546 | 0.37 | -13.546 | 37
[6 empty] | -0.051 | 25 | 13.536 | 0.369 | -13.536 | 36.9
[6 empty] | -0.051 | 25 | 13.526 | 0.369 | -13.526 | 36.9
[6 empty] | -0.051 | 25 | 13.516 | 0.369 | -13.516 | 36.9
[6 empty] | -0.051 | 25 | 13.506 | 0.369 | -13.506 | 36.9
[6 empty] | -0.051 | 25 | 13.496 | 0.369 | -13.496 | 36.9
[6 empty] | -0.051 | 25 | 13.486 | 0.368 | -13.486 | 36.8
[6 empty] | -0.051 | 25 | 13.476 | 0.368 | -13.476 | 36.8
[6 empty] | -0.051 | 25 | 13.466 | 0.368 | -13.466 | 36.8
[6 empty] | -0.051 | 25 | 13.456 | 0.368 | -13.456 | 36.8
[6 empty] | -0.051 | 25 | 13.446 | 0.368 | -13.446 | 36.8
[6 empty] | -0.051 | 25 | 13.436 | 0.367 | -13.436 | 36.7
[6 empty] | -0.051 | 25 | 13.426 | 0.367 | -13.426 | 36.7
[6 empty] | -0.051 | 25 | 13.416 | 0.367 | -13.416 | 36.7
[6 empty] | -0.051 | 25 | 13.406 | 0.367 | -13.406 | 36.7
[6 empty] | -0.051 | 25 | 13.396 | 0.367 | -13.396 | 36.7
[6 empty] | -0.051 | 25 | 13.386 | 0.366 | -13.386 | 36.6
[6 empty] | -0.051 | 25 | 13.376 | 0.366 | -13.376 | 36.6
[6 empty] | -0.051 | 25 | 13.366 | 0.366 | -13.366 | 36.6
[6 empty] | -0.051 | 25 | 13.356 | 0.366 | -13.356 | 36.6
[6 empty] | -0.051 | 25 | 13.346 | 0.365 | -13.346 | 36.5
[6 empty] | -0.051 | 25 | 13.336 | 0.365 | -13.336 | 36.5
[6 empty] | -0.051 | 25 | 13.326 | 0.365 | -13.326 | 36.5
[6 empty] | -0.051 | 25 | 13.316 | 0.365 | -13.316 | 36.5
[6 empty] | -0.051 | 25 | 13.306 | 0.365 | -13.306 | 36.5
[6 empty] | -0.051 | 25 | 13.296 | 0.364 | -13.296 | 36.4
[6 empty] | -0.051 | 25 | 13.286 | 0.364 | -13.286 | 36.4
[6 empty] | -0.051 | 25 | 13.276 | 0.364 | -13.276 | 36.4
[6 empty] | -0.051 | 25 | 13.266 | 0.364 | -13.266 | 36.4
[6 empty] | -0.051 | 25 | 13.256 | 0.364 | -13.256 | 36.4
[6 empty] | -0.052 | 25 | 13.246 | 0.363 | -13.246 | 36.3
[6 empty] | -0.052 | 25 | 13.236 | 0.363 | -13.236 | 36.3
[6 empty] | -0.052 | 25 | 13.226 | 0.363 | -13.226 | 36.3
[6 empty] | -0.052 | 25 | 13.216 | 0.363 | -13.216 | 36.3
[6 empty] | -0.052 | 25 | 13.206 | 0.363 | -13.206 | 36.3
[6 empty] | -0.052 | 25 | 13.196 | 0.362 | -13.196 | 36.2
[6 empty] | -0.052 | 25 | 13.186 | 0.362 | -13.186 | 36.2
[6 empty] | -0.052 | 25 | 13.176 | 0.362 | -13.176 | 36.2
[6 empty] | -0.052 | 25 | 13.166 | 0.362 | -13.166 | 36.2
[6 empty] | -0.052 | 25 | 13.156 | 0.362 | -13.156 | 36.2
[6 empty] | -0.052 | 25 | 13.146 | 0.361 | -13.146 | 36.1
[6 empty] | -0.052 | 25 | 13.136 | 0.361 | -13.136 | 36.1
[6 empty] | -0.052 | 25 | 13.126 | 0.361 | -13.126 | 36.1
[6 empty] | -0.052 | 25 | 13.116 | 0.361 | -13.116 | 36.1
[6 empty] | -0.052 | 25 | 13.106 | 0.361 | -13.106 | 36.1
[6 empty] | -0.052 | 25 | 13.096 | 0.36 | -13.096 | 36
[6 empty] | -0.052 | 25 | 13.086 | 0.36 | -13.086 | 36
[6 empty] | -0.052 | 25 | 13.076 | 0.36 | -13.076 | 36
[6 empty] | -0.052 | 25 | 13.066 | 0.36 | -13.066 | 36
[6 empty] | -0.052 | 25 | 13.056 | 0.36 | -13.056 | 36
[6 empty] | -0.052 | 25 | 13.046 | 0.359 | -13.046 | 35.9
[6 empty] | -0.052 | 25 | 13.036 | 0.359 | -13.036 | 35.9
[6 empty] | -0.052 | 25 | 13.026 | 0.359 | -13.026 | 35.9
[6 empty] | -0.052 | 25 | 13.016 | 0.359 | -13.016 | 35.9
[6 empty] | -0.052 | 25 | 13.006 | 0.359 | -13.006 | 35.9
[6 empty] | -0.052 | 25 | 12.996 | 0.358 | -12.996 | 35.8
[6 empty] | -0.052 | 25 | 12.986 | 0.358 | -12.986 | 35.8
[6 empty] | -0.052 | 25 | 12.976 | 0.358 | -12.976 | 35.8
[6 empty] | -0.052 | 25 | 12.966 | 0.358 | -12.966 | 35.8
[6 empty] | -0.052 | 25 | 12.956 | 0.358 | -12.956 | 35.8
[6 empty] | -0.052 | 25 | 12.946 | 0.357 | -12.946 | 35.7
[6 empty] | -0.052 | 25 | 12.936 | 0.357 | -12.936 | 35.7
[6 empty] | -0.052 | 25 | 12.926 | 0.357 | -12.926 | 35.7
[6 empty] | -0.052 | 25 | 12.916 | 0.357 | -12.916 | 35.7
[6 empty] | -0.052 | 25 | 12.906 | 0.357 | -12.906 | 35.7
[6 empty] | -0.052 | 25 | 12.896 | 0.356 | -12.896 | 35.6
[6 empty] | -0.052 | 25 | 12.886 | 0.356 | -12.886 | 35.6
[6 empty] | -0.052 | 25 | 12.876 | 0.356 | -12.876 | 35.6
[6 empty] | -0.052 | 25 | 12.866 | 0.356 | -12.866 | 35.6
[6 empty] | -0.052 | 25 | 12.856 | 0.356 | -12.856 | 35.6
[6 empty] | -0.052 | 25 | 12.846 | 0.355 | -12.846 | 35.5
[6 empty] | -0.052 | 25 | 12.836 | 0.355 | -12.836 | 35.5
[6 empty] | -0.052 | 25 | 12.826 | 0.355 | -12.826 | 35.5
[6 empty] | -0.052 | 25 | 12.816 | 0.355 | -12.816 | 35.5
[6 empty] | -0.052 | 25 | 12.806 | 0.355 | -12.806 | 35.5
[6 empty] | -0.052 | 25 | 12.796 | 0.354 | -12.796 | 35.4
[6 empty] | -0.052 | 25 | 12.786 | 0.354 | -12.786 | 35.4
[6 empty] | -0.052 | 25 | 12.776 | 0.354 | -12.776 | 35.4
[6 empty] | -0.052 | 25 | 12.766 | 0.354 | -12.766 | 35.4
[6 empty] | -0.052 | 25 | 12.756 | 0.354 | -12.756 | 35.4
[6 empty] | -0.052 | 25 | 12.746 | 0.353 | -12.746 | 35.3
[6 empty] | -0.052 | 25 | 12.736 | 0.353 | -12.736 | 35.3
[6 empty] | -0.052 | 25 | 12.726 | 0.353 | -12.726 | 35.3
[6 empty] | -0.052 | 25 | 12.716 | 0.353 | -12.716 | 35.3
[6 empty] | -0.052 | 25 | 12.706 | 0.353 | -12.706 | 35.3
[6 empty] | -0.052 | 25 | 12.696 | 0.352 | -12.696 | 35.2
[6 empty] | -0.052 | 25 | 12.686 | 0.352 | -12.686 | 35.2
[6 empty] | -0.052 | 25 | 12.676 | 0.352 | -12.676 | 35.2
[6 empty] | -0.052 | 25 | 12.666 | 0.352 | -12.666 | 35.2
[6 empty] | -0.052 | 25 | 12.656 | 0.352 | -12.656 | 35.2
[6 empty] | -0.052 | 25 | 12.646 | 0.352 | -12.646 | 35.2
[6 empty] | -0.052 | 25 | 12.636 | 0.351 | -12.636 | 35.1
[6 empty] | -0.052 | 25 | 12.626 | 0.351 | -12.626 | 35.1
[6 empty] | -0.052 | 25 | 12.616 | 0.351 | -12.616 | 35.1
[6 empty] | -0.052 | 25 | 12.606 | 0.351 | -12.606 | 35.1
[6 empty] | -0.052 | 25 | 12.596 | 0.351 | -12.596 | 35.1
[6 empty] | -0.052 | 25 | 12.586 | 0.35 | -12.586 | 35
[6 empty] | -0.052 | 25 | 12.576 | 0.35 | -12.576 | 35
[6 empty] | -0.052 | 25 | 12.566 | 0.35 | -12.566 | 35
[6 empty] | -0.052 | 25 | 12.556 | 0.35 | -12.556 | 35
[6 empty] | -0.053 | 25 | 12.546 | 0.35 | -12.546 | 35
[6 empty] | -0.053 | 25 | 12.536 | 0.349 | -12.536 | 34.9
[6 empty] | -0.053 | 25 | 12.526 | 0.349 | -12.526 | 34.9
[6 empty] | -0.053 | 25 | 12.516 | 0.349 | -12.516 | 34.9
[6 empty] | -0.053 | 25 | 12.506 | 0.349 | -12.506 | 34.9
[6 empty] | -0.053 | 25 | 12.496 | 0.349 | -12.496 | 34.9
[6 empty] | -0.053 | 25 | 12.486 | 0.348 | -12.486 | 34.8
[6 empty] | -0.053 | 25 | 12.476 | 0.348 | -12.476 | 34.8
[6 empty] | -0.053 | 25 | 12.466 | 0.348 | -12.466 | 34.8
[6 empty] | -0.053 | 25 | 12.456 | 0.348 | -12.456 | 34.8
[6 empty] | -0.053 | 25 | 12.446 | 0.348 | -12.446 | 34.8
[6 empty] | -0.053 | 25 | 12.436 | 0.347 | -12.436 | 34.7
[6 empty] | -0.053 | 25 | 12.426 | 0.347 | -12.426 | 34.7
[6 empty] | -0.053 | 25 | 12.416 | 0.347 | -12.416 | 34.7
[6 empty] | -0.053 | 25 | 12.406 | 0.347 | -12.406 | 34.7
[6 empty] | -0.053 | 25 | 12.396 | 0.347 | -12.396 | 34.7
[6 empty] | -0.053 | 25 | 12.386 | 0.347 | -12.386 | 34.7
[6 empty] | -0.053 | 25 | 12.376 | 0.346 | -12.376 | 34.6
[6 empty] | -0.053 | 25 | 12.366 | 0.346 | -12.366 | 34.6
[6 empty] | -0.053 | 25 | 12.356 | 0.346 | -12.356 | 34.6
[6 empty] | -0.053 | 25 | 12.346 | 0.346 | -12.346 | 34.6
[6 empty] | -0.053 | 25 | 12.336 | 0.346 | -12.336 | 34.6
[6 empty] | -0.053 | 25 | 12.326 | 0.345 | -12.326 | 34.5
[6 empty] | -0.053 | 25 | 12.316 | 0.345 | -12.316 | 34.5
[6 empty] | -0.053 | 25 | 12.306 | 0.345 | -12.306 | 34.5
[6 empty] | -0.053 | 25 | 12.296 | 0.345 | -12.296 | 34.5
[6 empty] | -0.053 | 25 | 12.286 | 0.345 | -12.286 | 34.5
[6 empty] | -0.053 | 25 | 12.276 | 0.344 | -12.276 | 34.4
[6 empty] | -0.053 | 25 | 12.266 | 0.344 | -12.266 | 34.4
[6 empty] | -0.053 | 25 | 12.256 | 0.344 | -12.256 | 34.4
[6 empty] | -0.053 | 25 | 12.246 | 0.344 | -12.246 | 34.4
[6 empty] | -0.053 | 25 | 12.236 | 0.344 | -12.236 | 34.4
[6 empty] | -0.053 | 25 | 12.226 | 0.344 | -12.226 | 34.4
[6 empty] | -0.053 | 25 | 12.216 | 0.343 | -12.216 | 34.3
[6 empty] | -0.053 | 25 | 12.207 | 0.343 | -12.207 | 34.3
[6 empty] | -0.053 | 25 | 12.197 | 0.343 | -12.197 | 34.3
[6 empty] | -0.053 | 25 | 12.187 | 0.343 | -12.187 | 34.3
[6 empty] | -0.053 | 25 | 12.177 | 0.343 | -12.177 | 34.3
[6 empty] | -0.053 | 25 | 12.167 | 0.342 | -12.167 | 34.2
[6 empty] | -0.053 | 25 | 12.157 | 0.342 | -12.157 | 34.2
[6 empty] | -0.053 | 25 | 12.147 | 0.342 | -12.147 | 34.2
[6 empty] | -0.053 | 25 | 12.137 | 0.342 | -12.137 | 34.2
[6 empty] | -0.053 | 25 | 12.127 | 0.342 | -12.127 | 34.2
[6 empty] | -0.053 | 25 | 12.117 | 0.341 | -12.117 | 34.1
[6 empty] | -0.053 | 25 | 12.107 | 0.341 | -12.107 | 34.1
[6 empty] | -0.053 | 25 | 12.097 | 0.341 | -12.097 | 34.1
[6 empty] | -0.053 | 25 | 12.087 | 0.341 | -12.087 | 34.1
[6 empty] | -0.053 | 25 | 12.077 | 0.341 | -12.077 | 34.1
[6 empty] | -0.053 | 25 | 12.067 | 0.341 | -12.067 | 34.1
[6 empty] | -0.053 | 25 | 12.057 | 0.34 | -12.057 | 34
[6 empty] | -0.053 | 25 | 12.047 | 0.34 | -12.047 | 34
[6 empty] | -0.053 | 25 | 12.037 | 0.34 | -12.037 | 34
[6 empty] | -0.053 | 25 | 12.027 | 0.34 | -12.027 | 34
[6 empty] | -0.053 | 25 | 12.017 | 0.34 | -12.017 | 34
[6 empty] | -0.053 | 25 | 12.007 | 0.339 | -12.007 | 33.9
[6 empty] | -0.053 | 25 | 11.997 | 0.339 | -11.997 | 33.9
[6 empty] | -0.053 | 25 | 11.987 | 0.339 | -11.987 | 33.9
[6 empty] | -0.053 | 25 | 11.977 | 0.339 | -11.977 | 33.9
[6 empty] | -0.053 | 25 | 11.967 | 0.339 | -11.967 | 33.9
[6 empty] | -0.053 | 25 | 11.957 | 0.339 | -11.957 | 33.9
[6 empty] | -0.053 | 25 | 11.947 | 0.338 | -11.947 | 33.8
[6 empty] | -0.053 | 25 | 11.937 | 0.338 | -11.937 | 33.8
[6 empty] | -0.053 | 25 | 11.927 | 0.338 | -11.927 | 33.8
[6 empty] | -0.053 | 25 | 11.917 | 0.338 | -11.917 | 33.8
[6 empty] | -0.053 | 25 | 11.907 | 0.338 | -11.907 | 33.8
[6 empty] | -0.053 | 25 | 11.897 | 0.337 | -11.897 | 33.7
[6 empty] | -0.053 | 25 | 11.887 | 0.337 | -11.887 | 33.7
[6 empty] | -0.053 | 25 | 11.877 | 0.337 | -11.877 | 33.7
[6 empty] | -0.053 | 25 | 11.867 | 0.337 | -11.867 | 33.7
[6 empty] | -0.053 | 25 | 11.857 | 0.337 | -11.857 | 33.7
[6 empty] | -0.054 | 25 | 11.847 | 0.336 | -11.847 | 33.6
[6 empty] | -0.054 | 25 | 11.837 | 0.336 | -11.837 | 33.6
[6 empty] | -0.054 | 25 | 11.827 | 0.336 | -11.827 | 33.6
[6 empty] | -0.054 | 25 | 11.817 | 0.336 | -11.817 | 33.6
[6 empty] | -0.054 | 25 | 11.807 | 0.336 | -11.807 | 33.6
[6 empty] | -0.054 | 25 | 11.797 | 0.336 | -11.797 | 33.6
[6 empty] | -0.054 | 25 | 11.787 | 0.335 | -11.787 | 33.5
[6 empty] | -0.054 | 25 | 11.777 | 0.335 | -11.777 | 33.5
[6 empty] | -0.054 | 25 | 11.767 | 0.335 | -11.767 | 33.5
[6 empty] | -0.054 | 25 | 11.757 | 0.335 | -11.757 | 33.5
[6 empty] | -0.054 | 25 | 11.747 | 0.335 | -11.747 | 33.5
[6 empty] | -0.054 | 25 | 11.737 | 0.334 | -11.737 | 33.4
[6 empty] | -0.054 | 25 | 11.727 | 0.334 | -11.727 | 33.4
[6 empty] | -0.054 | 25 | 11.717 | 0.334 | -11.717 | 33.4
[6 empty] | -0.054 | 25 | 11.707 | 0.334 | -11.707 | 33.4
[6 empty] | -0.054 | 25 | 11.697 | 0.334 | -11.697 | 33.4
[6 empty] | -0.054 | 25 | 11.687 | 0.334 | -11.687 | 33.4
[6 empty] | -0.054 | 25 | 11.677 | 0.333 | -11.677 | 33.3
[6 empty] | -0.054 | 25 | 11.667 | 0.333 | -11.667 | 33.3
[6 empty] | -0.054 | 25 | 11.657 | 0.333 | -11.657 | 33.3
[6 empty] | -0.054 | 25 | 11.647 | 0.333 | -11.647 | 33.3
[6 empty] | -0.054 | 25 | 11.637 | 0.333 | -11.637 | 33.3
[6 empty] | -0.054 | 25 | 11.627 | 0.332 | -11.627 | 33.2
[6 empty] | -0.054 | 25 | 11.617 | 0.332 | -11.617 | 33.2
[6 empty] | -0.054 | 25 | 11.607 | 0.332 | -11.607 | 33.2
[6 empty] | -0.054 | 25 | 11.597 | 0.332 | -11.597 | 33.2
[6 empty] | -0.054 | 25 | 11.587 | 0.332 | -11.587 | 33.2
[6 empty] | -0.054 | 25 | 11.577 | 0.332 | -11.577 | 33.2
[6 empty] | -0.054 | 25 | 11.567 | 0.331 | -11.567 | 33.1
[6 empty] | -0.054 | 25 | 11.557 | 0.331 | -11.557 | 33.1
[6 empty] | -0.054 | 25 | 11.547 | 0.331 | -11.547 | 33.1
[6 empty] | -0.054 | 25 | 11.537 | 0.331 | -11.537 | 33.1
[6 empty] | -0.054 | 25 | 11.527 | 0.331 | -11.527 | 33.1
[6 empty] | -0.054 | 25 | 11.517 | 0.33 | -11.517 | 33
[6 empty] | -0.054 | 25 | 11.507 | 0.33 | -11.507 | 33
[6 empty] | -0.054 | 25 | 11.497 | 0.33 | -11.497 | 33
[6 empty] | -0.054 | 25 | 11.487 | 0.33 | -11.487 | 33
[6 empty] | -0.054 | 25 | 11.477 | 0.33 | -11.477 | 33
[6 empty] | -0.054 | 25 | 11.467 | 0.33 | -11.467 | 33
[6 empty] | -0.054 | 25 | 11.457 | 0.329 | -11.457 | 32.9
[6 empty] | -0.054 | 25 | 11.447 | 0.329 | -11.447 | 32.9
[6 empty] | -0.054 | 25 | 11.437 | 0.329 | -11.437 | 32.9
[6 empty] | -0.054 | 25 | 11.427 | 0.329 | -11.427 | 32.9
[6 empty] | -0.054 | 25 | 11.417 | 0.329 | -11.417 | 32.9
[6 empty] | -0.054 | 25 | 11.407 | 0.328 | -11.407 | 32.8
[6 empty] | -0.054 | 25 | 11.397 | 0.328 | -11.397 | 32.8
[6 empty] | -0.054 | 25 | 11.387 | 0.328 | -11.387 | 32.8
[6 empty] | -0.054 | 25 | 11.377 | 0.328 | -11.377 | 32.8
[6 empty] | -0.054 | 25 | 11.367 | 0.328 | -11.367 | 32.8
[6 empty] | -0.054 | 25 | 11.357 | 0.327 | -11.357 | 32.7
[6 empty] | -0.054 | 25 | 11.347 | 0.327 | -11.347 | 32.7
[6 empty] | -0.054 | 25 | 11.337 | 0.327 | -11.337 | 32.7
[6 empty] | -0.054 | 25 | 11.327 | 0.327 | -11.327 | 32.7
[6 empty] | -0.054 | 25 | 11.317 | 0.327 | -11.317 | 32.7
[6 empty] | -0.054 | 25 | 11.307 | 0.327 | -11.307 | 32.7
[6 empty] | -0.054 | 25 | 11.297 | 0.326 | -11.297 | 32.6
[6 empty] | -0.054 | 25 | 11.287 | 0.326 | -11.287 | 32.6
[6 empty] | -0.054 | 25 | 11.277 | 0.326 | -11.277 | 32.6
[6 empty] | -0.054 | 25 | 11.267 | 0.326 | -11.267 | 32.6
[6 empty] | -0.054 | 25 | 11.257 | 0.326 | -11.257 | 32.6
[6 empty] | -0.054 | 25 | 11.247 | 0.325 | -11.247 | 32.5
[6 empty] | -0.054 | 25 | 11.237 | 0.325 | -11.237 | 32.5
[6 empty] | -0.054 | 25 | 11.227 | 0.325 | -11.227 | 32.5
[6 empty] | -0.054 | 25 | 11.217 | 0.325 | -11.217 | 32.5
[6 empty] | -0.054 | 25 | 11.207 | 0.325 | -11.207 | 32.5
[6 empty] | -0.054 | 25 | 11.197 | 0.325 | -11.197 | 32.5
[6 empty] | -0.054 | 25 | 11.187 | 0.324 | -11.187 | 32.4
[6 empty] | -0.054 | 25 | 11.177 | 0.324 | -11.177 | 32.4
[6 empty] | -0.054 | 25 | 11.167 | 0.324 | -11.167 | 32.4
[6 empty] | -0.054 | 25 | 11.157 | 0.324 | -11.157 | 32.4
[6 empty] | -0.055 | 25 | 11.147 | 0.324 | -11.147 | 32.4
[6 empty] | -0.055 | 25 | 11.137 | 0.323 | -11.137 | 32.3
[6 empty] | -0.055 | 25 | 11.127 | 0.323 | -11.127 | 32.3
[6 empty] | -0.055 | 25 | 11.117 | 0.323 | -11.117 | 32.3
[6 empty] | -0.055 | 25 | 11.107 | 0.323 | -11.107 | 32.3
[6 empty] | -0.055 | 25 | 11.097 | 0.323 | -11.097 | 32.3
[6 empty] | -0.055 | 25 | 11.087 | 0.323 | -11.087 | 32.3
[6 empty] | -0.055 | 25 | 11.077 | 0.322 | -11.077 | 32.2
[6 empty] | -0.055 | 25 | 11.067 | 0.322 | -11.067 | 32.2
[6 empty] | -0.055 | 25 | 11.057 | 0.322 | -11.057 | 32.2
[6 empty] | -0.055 | 25 | 11.047 | 0.322 | -11.047 | 32.2
[6 empty] | -0.055 | 25 | 11.037 | 0.322 | -11.037 | 32.2
[6 empty] | -0.055 | 25 | 11.027 | 0.322 | -11.027 | 32.2
[6 empty] | -0.055 | 25 | 11.017 | 0.321 | -11.017 | 32.1
[6 empty] | -0.055 | 25 | 11.007 | 0.321 | -11.007 | 32.1
[6 empty] | -0.055 | 25 | 10.997 | 0.321 | -10.997 | 32.1
[6 empty] | -0.055 | 25 | 10.987 | 0.321 | -10.987 | 32.1
[6 empty] | -0.055 | 25 | 10.977 | 0.321 | -10.977 | 32.1
[6 empty] | -0.055 | 25 | 10.967 | 0.32 | -10.967 | 32
[6 empty] | -0.055 | 25 | 10.957 | 0.32 | -10.957 | 32
[6 empty] | -0.055 | 25 | 10.947 | 0.32 | -10.947 | 32
[6 empty] | -0.055 | 25 | 10.937 | 0.32 | -10.937 | 32
[6 empty] | -0.055 | 25 | 10.927 | 0.32 | -10.927 | 32
[6 empty] | -0.055 | 25 | 10.917 | 0.32 | -10.917 | 32
[6 empty] | -0.055 | 25 | 10.907 | 0.319 | -10.907 | 31.9
[6 empty] | -0.055 | 25 | 10.897 | 0.319 | -10.897 | 31.9
[6 empty] | -0.055 | 25 | 10.887 | 0.319 | -10.887 | 31.9
[6 empty] | -0.055 | 25 | 10.877 | 0.319 | -10.877 | 31.9
[6 empty] | -0.055 | 25 | 10.867 | 0.319 | -10.867 | 31.9
[6 empty] | -0.055 | 25 | 10.857 | 0.319 | -10.857 | 31.9
[6 empty] | -0.055 | 25 | 10.847 | 0.318 | -10.847 | 31.8
[6 empty] | -0.055 | 25 | 10.837 | 0.318 | -10.837 | 31.8
[6 empty] | -0.055 | 25 | 10.827 | 0.318 | -10.827 | 31.8
[6 empty] | -0.055 | 25 | 10.817 | 0.318 | -10.817 | 31.8
[6 empty] | -0.055 | 25 | 10.807 | 0.318 | -10.807 | 31.8
[6 empty] | -0.055 | 25 | 10.797 | 0.317 | -10.797 | 31.7
[6 empty] | -0.055 | 25 | 10.787 | 0.317 | -10.787 | 31.7
[6 empty] | -0.055 | 25 | 10.777 | 0.317 | -10.777 | 31.7
[6 empty] | -0.055 | 25 | 10.767 | 0.317 | -10.767 | 31.7
[6 empty] | -0.055 | 25 | 10.757 | 0.317 | -10.757 | 31.7
[6 empty] | -0.055 | 25 | 10.747 | 0.317 | -10.747 | 31.7
[6 empty] | -0.055 | 25 | 10.737 | 0.316 | -10.737 | 31.6
[6 empty] | -0.055 | 25 | 10.727 | 0.316 | -10.727 | 31.6
[6 empty] | -0.055 | 25 | 10.717 | 0.316 | -10.717 | 31.6
[6 empty] | -0.055 | 25 | 10.707 | 0.316 | -10.707 | 31.6
[6 empty] | -0.055 | 25 | 10.697 | 0.316 | -10.697 | 31.6
[6 empty] | -0.055 | 25 | 10.687 | 0.316 | -10.687 | 31.6
[6 empty] | -0.055 | 25 | 10.677 | 0.315 | -10.677 | 31.5
[6 empty] | -0.055 | 25 | 10.667 | 0.315 | -10.667 | 31.5
[6 empty] | -0.055 | 25 | 10.657 | 0.315 | -10.657 | 31.5
[6 empty] | -0.055 | 25 | 10.647 | 0.315 | -10.647 | 31.5
[6 empty] | -0.055 | 25 | 10.637 | 0.315 | -10.637 | 31.5
[6 empty] | -0.055 | 25 | 10.627 | 0.315 | -10.627 | 31.5
[6 empty] | -0.055 | 25 | 10.617 | 0.314 | -10.617 | 31.4
[6 empty] | -0.055 | 25 | 10.607 | 0.314 | -10.607 | 31.4
[6 empty] | -0.055 | 25 | 10.597 | 0.314 | -10.597 | 31.4
[6 empty] | -0.055 | 25 | 10.587 | 0.314 | -10.587 | 31.4
[6 empty] | -0.055 | 25 | 10.577 | 0.314 | -10.577 | 31.4
[6 empty] | -0.055 | 25 | 10.567 | 0.313 | -10.567 | 31.3
[6 empty] | -0.055 | 25 | 10.557 | 0.313 | -10.557 | 31.3
[6 empty] | -0.055 | 25 | 10.547 | 0.313 | -10.547 | 31.3
[6 empty] | -0.055 | 25 | 10.537 | 0.313 | -10.537 | 31.3
[6 empty] | -0.055 | 25 | 10.527 | 0.313 | -10.527 | 31.3
[6 empty] | -0.055 | 25 | 10.517 | 0.313 | -10.517 | 31.3
[6 empty] | -0.055 | 25 | 10.507 | 0.312 | -10.507 | 31.2
[6 empty] | -0.055 | 25 | 10.497 | 0.312 | -10.497 | 31.2
[6 empty] | -0.055 | 25 | 10.487 | 0.312 | -10.487 | 31.2
[6 empty] | -0.055 | 25 | 10.477 | 0.312 | -10.477 | 31.2
[6 empty] | -0.055 | 25 | 10.467 | 0.312 | -10.467 | 31.2
[6 empty] | -0.055 | 25 | 10.457 | 0.312 | -10.457 | 31.2
[6 empty] | -0.056 | 25 | 10.447 | 0.311 | -10.447 | 31.1
[6 empty] | -0.056 | 25 | 10.437 | 0.311 | -10.437 | 31.1
[6 empty] | -0.056 | 25 | 10.427 | 0.311 | -10.427 | 31.1
[6 empty] | -0.056 | 25 | 10.418 | 0.311 | -10.418 | 31.1
[6 empty] | -0.056 | 25 | 10.408 | 0.311 | -10.408 | 31.1
[6 empty] | -0.056 | 25 | 10.398 | 0.311 | -10.398 | 31.1
[6 empty] | -0.056 | 25 | 10.388 | 0.31 | -10.388 | 31
[6 empty] | -0.056 | 25 | 10.378 | 0.31 | -10.378 | 31
[6 empty] | -0.056 | 25 | 10.368 | 0.31 | -10.368 | 31
[6 empty] | -0.056 | 25 | 10.358 | 0.31 | -10.358 | 31
[6 empty] | -0.056 | 25 | 10.348 | 0.31 | -10.348 | 31
[6 empty] | -0.056 | 25 | 10.338 | 0.31 | -10.338 | 31
[6 empty] | -0.056 | 25 | 10.328 | 0.309 | -10.328 | 30.9
[6 empty] | -0.056 | 25 | 10.318 | 0.309 | -10.318 | 30.9
[6 empty] | -0.056 | 25 | 10.308 | 0.309 | -10.308 | 30.9
[6 empty] | -0.056 | 25 | 10.298 | 0.309 | -10.298 | 30.9
[6 empty] | -0.056 | 25 | 10.288 | 0.309 | -10.288 | 30.9
[6 empty] | -0.056 | 25 | 10.278 | 0.308 | -10.278 | 30.8
[6 empty] | -0.056 | 25 | 10.268 | 0.308 | -10.268 | 30.8
[6 empty] | -0.056 | 25 | 10.258 | 0.308 | -10.258 | 30.8
[6 empty] | -0.056 | 25 | 10.248 | 0.308 | -10.248 | 30.8
[6 empty] | -0.056 | 25 | 10.238 | 0.308 | -10.238 | 30.8
[6 empty] | -0.056 | 25 | 10.228 | 0.308 | -10.228 | 30.8
[6 empty] | -0.056 | 25 | 10.218 | 0.307 | -10.218 | 30.7
[6 empty] | -0.056 | 25 | 10.208 | 0.307 | -10.208 | 30.7
[6 empty] | -0.056 | 25 | 10.198 | 0.307 | -10.198 | 30.7
[6 empty] | -0.056 | 25 | 10.188 | 0.307 | -10.188 | 30.7
[6 empty] | -0.056 | 25 | 10.178 | 0.307 | -10.178 | 30.7
[6 empty] | -0.056 | 25 | 10.168 | 0.307 | -10.168 | 30.7
[6 empty] | -0.056 | 25 | 10.158 | 0.306 | -10.158 | 30.6
[6 empty] | -0.056 | 25 | 10.148 | 0.306 | -10.148 | 30.6
[6 empty] | -0.056 | 25 | 10.138 | 0.306 | -10.138 | 30.6
[6 empty] | -0.056 | 25 | 10.128 | 0.306 | -10.128 | 30.6
[6 empty] | -0.056 | 25 | 10.118 | 0.306 | -10.118 | 30.6
[6 empty] | -0.056 | 25 | 10.108 | 0.306 | -10.108 | 30.6
[6 empty] | -0.056 | 25 | 10.098 | 0.305 | -10.098 | 30.5
[6 empty] | -0.056 | 25 | 10.088 | 0.305 | -10.088 | 30.5
[6 empty] | -0.056 | 25 | 10.078 | 0.305 | -10.078 | 30.5
[6 empty] | -0.056 | 25 | 10.068 | 0.305 | -10.068 | 30.5
[6 empty] | -0.056 | 25 | 10.058 | 0.305 | -10.058 | 30.5
[6 empty] | -0.056 | 25 | 10.048 | 0.305 | -10.048 | 30.5
[6 empty] | -0.056 | 25 | 10.038 | 0.304 | -10.038 | 30.4
[6 empty] | -0.056 | 25 | 10.028 | 0.304 | -10.028 | 30.4
[6 empty] | -0.056 | 25 | 10.018 | 0.304 | -10.018 | 30.4
[6 empty] | -0.056 | 25 | 10.008 | 0.304 | -10.008 | 30.4
[6 empty] | -0.056 | 25 | 9.998 | 0.304 | -9.998 | 30.4
[6 empty] | -0.056 | 25 | 9.988 | 0.304 | -9.988 | 30.4
[6 empty] | -0.056 | 25 | 9.978 | 0.303 | -9.978 | 30.3
[6 empty] | -0.056 | 25 | 9.968 | 0.303 | -9.968 | 30.3
[6 empty] | -0.056 | 25 | 9.958 | 0.303 | -9.958 | 30.3
[6 empty] | -0.056 | 25 | 9.948 | 0.303 | -9.948 | 30.3
[6 empty] | -0.056 | 25 | 9.938 | 0.303 | -9.938 | 30.3
[6 empty] | -0.056 | 25 | 9.928 | 0.303 | -9.928 | 30.3
[6 empty] | -0.056 | 25 | 9.918 | 0.302 | -9.918 | 30.2
[6 empty] | -0.056 | 25 | 9.908 | 0.302 | -9.908 | 30.2
[6 empty] | -0.056 | 25 | 9.898 | 0.302 | -9.898 | 30.2
[6 empty] | -0.056 | 25 | 9.888 | 0.302 | -9.888 | 30.2
[6 empty] | -0.056 | 25 | 9.878 | 0.302 | -9.878 | 30.2
[6 empty] | -0.056 | 25 | 9.868 | 0.302 | -9.868 | 30.2
[6 empty] | -0.056 | 25 | 9.858 | 0.301 | -9.858 | 30.1
[6 empty] | -0.056 | 25 | 9.848 | 0.301 | -9.848 | 30.1
[6 empty] | -0.056 | 25 | 9.838 | 0.301 | -9.838 | 30.1
[6 empty] | -0.056 | 25 | 9.828 | 0.301 | -9.828 | 30.1
[6 empty] | -0.056 | 25 | 9.818 | 0.301 | -9.818 | 30.1
[6 empty] | -0.056 | 25 | 9.808 | 0.301 | -9.808 | 30.1
[6 empty] | -0.056 | 25 | 9.798 | 0.3 | -9.798 | 30
[6 empty] | -0.056 | 25 | 9.788 | 0.3 | -9.788 | 30
[6 empty] | -0.056 | 25 | 9.778 | 0.3 | -9.778 | 30
[6 empty] | -0.056 | 25 | 9.768 | 0.3 | -9.768 | 30
[6 empty] | -0.057 | 25 | 9.758 | 0.3 | -9.758 | 30
[6 empty] | -0.057 | 25 | 9.748 | 0.3 | -9.748 | 30
[6 empty] | -0.057 | 25 | 9.738 | 0.299 | -9.738 | 29.9
[6 empty] | -0.057 | 25 | 9.728 | 0.299 | -9.728 | 29.9
[6 empty] | -0.057 | 25 | 9.718 | 0.299 | -9.718 | 29.9
[6 empty] | -0.057 | 25 | 9.708 | 0.299 | -9.708 | 29.9
[6 empty] | -0.057 | 25 | 9.698 | 0.299 | -9.698 | 29.9
[6 empty] | -0.057 | 25 | 9.688 | 0.299 | -9.688 | 29.9
[6 empty] | -0.057 | 25 | 9.678 | 0.298 | -9.678 | 29.8
[6 empty] | -0.057 | 25 | 9.668 | 0.298 | -9.668 | 29.8
[6 empty] | -0.057 | 25 | 9.658 | 0.298 | -9.658 | 29.8
[6 empty] | -0.057 | 25 | 9.648 | 0.298 | -9.648 | 29.8
[6 empty] | -0.057 | 25 | 9.638 | 0.298 | -9.638 | 29.8
[6 empty] | -0.057 | 25 | 9.628 | 0.298 | -9.628 | 29.8
[6 empty] | -0.057 | 25 | 9.618 | 0.297 | -9.618 | 29.7
[6 empty] | -0.057 | 25 | 9.608 | 0.297 | -9.608 | 29.7
[6 empty] | -0.057 | 25 | 9.598 | 0.297 | -9.598 | 29.7
[6 empty] | -0.057 | 25 | 9.588 | 0.297 | -9.588 | 29.7
[6 empty] | -0.057 | 25 | 9.578 | 0.297 | -9.578 | 29.7
[6 empty] | -0.057 | 25 | 9.568 | 0.297 | -9.568 | 29.7
[6 empty] | -0.057 | 25 | 9.558 | 0.297 | -9.558 | 29.7
[6 empty] | -0.057 | 25 | 9.548 | 0.296 | -9.548 | 29.6
[6 empty] | -0.057 | 25 | 9.538 | 0.296 | -9.538 | 29.6
[6 empty] | -0.057 | 25 | 9.528 | 0.296 | -9.528 | 29.6
[6 empty] | -0.057 | 25 | 9.518 | 0.296 | -9.518 | 29.6
[6 empty] | -0.057 | 25 | 9.508 | 0.296 | -9.508 | 29.6
[6 empty] | -0.057 | 25 | 9.498 | 0.296 | -9.498 | 29.6
[6 empty] | -0.057 | 25 | 9.488 | 0.295 | -9.488 | 29.5
[6 empty] | -0.057 | 25 | 9.478 | 0.295 | -9.478 | 29.5
[6 empty] | -0.057 | 25 | 9.468 | 0.295 | -9.468 | 29.5
[6 empty] | -0.057 | 25 | 9.458 | 0.295 | -9.458 | 29.5
[6 empty] | -0.057 | 25 | 9.448 | 0.295 | -9.448 | 29.5
[6 empty] | -0.057 | 25 | 9.438 | 0.295 | -9.438 | 29.5
[6 empty] | -0.057 | 25 | 9.428 | 0.294 | -9.428 | 29.4
[6 empty] | -0.057 | 25 | 9.418 | 0.294 | -9.418 | 29.4
[6 empty] | -0.057 | 25 | 9.408 | 0.294 | -9.408 | 29.4
[6 empty] | -0.057 | 25 | 9.398 | 0.294 | -9.398 | 29.4
[6 empty] | -0.057 | 25 | 9.388 | 0.294 | -9.388 | 29.4
[6 empty] | -0.057 | 25 | 9.378 | 0.294 | -9.378 | 29.4
[6 empty] | -0.057 | 25 | 9.368 | 0.293 | -9.368 | 29.3
[6 empty] | -0.057 | 25 | 9.358 | 0.293 | -9.358 | 29.3
[6 empty] | -0.057 | 25 | 9.348 | 0.293 | -9.348 | 29.3
[6 empty] | -0.057 | 25 | 9.338 | 0.293 | -9.338 | 29.3
[6 empty] | -0.057 | 25 | 9.328 | 0.293 | -9.328 | 29.3
[6 empty] | -0.057 | 25 | 9.318 | 0.293 | -9.318 | 29.3
[6 empty] | -0.057 | 25 | 9.308 | 0.292 | -9.308 | 29.2
[6 empty] | -0.057 | 25 | 9.298 | 0.292 | -9.298 | 29.2
[6 empty] | -0.057 | 25 | 9.288 | 0.292 | -9.288 | 29.2
[6 empty] | -0.057 | 25 | 9.278 | 0.292 | -9.278 | 29.2
[6 empty] | -0.057 | 25 | 9.268 | 0.292 | -9.268 | 29.2
[6 empty] | -0.057 | 25 | 9.258 | 0.292 | -9.258 | 29.2
[6 empty] | -0.057 | 25 | 9.248 | 0.291 | -9.248 | 29.1
[6 empty] | -0.057 | 25 | 9.238 | 0.291 | -9.238 | 29.1
[6 empty] | -0.057 | 25 | 9.228 | 0.291 | -9.228 | 29.1
[6 empty] | -0.057 | 25 | 9.218 | 0.291 | -9.218 | 29.1
[6 empty] | -0.057 | 25 | 9.208 | 0.291 | -9.208 | 29.1
[6 empty] | -0.057 | 25 | 9.198 | 0.291 | -9.198 | 29.1
[6 empty] | -0.057 | 25 | 9.188 | 0.291 | -9.188 | 29.1
[6 empty] | -0.057 | 25 | 9.178 | 0.29 | -9.178 | 29
[6 empty] | -0.057 | 25 | 9.168 | 0.29 | -9.168 | 29
[6 empty] | -0.057 | 25 | 9.158 | 0.29 | -9.158 | 29
[6 empty] | -0.057 | 25 | 9.148 | 0.29 | -9.148 | 29
[6 empty] | -0.057 | 25 | 9.138 | 0.29 | -9.138 | 29
[6 empty] | -0.057 | 25 | 9.128 | 0.29 | -9.128 | 29
[6 empty] | -0.057 | 25 | 9.118 | 0.289 | -9.118 | 28.9
[6 empty] | -0.057 | 25 | 9.108 | 0.289 | -9.108 | 28.9
[6 empty] | -0.057 | 25 | 9.098 | 0.289 | -9.098 | 28.9
[6 empty] | -0.057 | 25 | 9.088 | 0.289 | -9.088 | 28.9
[6 empty] | -0.057 | 25 | 9.078 | 0.289 | -9.078 | 28.9
[6 empty] | -0.057 | 25 | 9.068 | 0.289 | -9.068 | 28.9
[6 empty] | -0.058 | 25 | 9.058 | 0.288 | -9.058 | 28.8
[6 empty] | -0.058 | 25 | 9.048 | 0.288 | -9.048 | 28.8
[6 empty] | -0.058 | 25 | 9.038 | 0.288 | -9.038 | 28.8
[6 empty] | -0.058 | 25 | 9.028 | 0.288 | -9.028 | 28.8
[6 empty] | -0.058 | 25 | 9.018 | 0.288 | -9.018 | 28.8
[6 empty] | -0.058 | 25 | 9.008 | 0.288 | -9.008 | 28.8
[6 empty] | -0.058 | 25 | 8.998 | 0.287 | -8.998 | 28.7
[6 empty] | -0.058 | 25 | 8.988 | 0.287 | -8.988 | 28.7
[6 empty] | -0.058 | 25 | 8.978 | 0.287 | -8.978 | 28.7
[6 empty] | -0.058 | 25 | 8.968 | 0.287 | -8.968 | 28.7
[6 empty] | -0.058 | 25 | 8.958 | 0.287 | -8.958 | 28.7
[6 empty] | -0.058 | 25 | 8.948 | 0.287 | -8.948 | 28.7
[6 empty] | -0.058 | 25 | 8.938 | 0.287 | -8.938 | 28.7
[6 empty] | -0.058 | 25 | 8.928 | 0.286 | -8.928 | 28.6
[6 empty] | -0.058 | 25 | 8.918 | 0.286 | -8.918 | 28.6
[6 empty] | -0.058 | 25 | 8.908 | 0.286 | -8.908 | 28.6
[6 empty] | -0.058 | 25 | 8.898 | 0.286 | -8.898 | 28.6
[6 empty] | -0.058 | 25 | 8.888 | 0.286 | -8.888 | 28.6
[6 empty] | -0.058 | 25 | 8.878 | 0.286 | -8.878 | 28.6
[6 empty] | -0.058 | 25 | 8.868 | 0.285 | -8.868 | 28.5
[6 empty] | -0.058 | 25 | 8.858 | 0.285 | -8.858 | 28.5
[6 empty] | -0.058 | 25 | 8.848 | 0.285 | -8.848 | 28.5
[6 empty] | -0.058 | 25 | 8.838 | 0.285 | -8.838 | 28.5
[6 empty] | -0.058 | 25 | 8.828 | 0.285 | -8.828 | 28.5
[6 empty] | -0.058 | 25 | 8.818 | 0.285 | -8.818 | 28.5
[6 empty] | -0.058 | 25 | 8.808 | 0.284 | -8.808 | 28.4
[6 empty] | -0.058 | 25 | 8.798 | 0.284 | -8.798 | 28.4
[6 empty] | -0.058 | 25 | 8.788 | 0.284 | -8.788 | 28.4
[6 empty] | -0.058 | 25 | 8.778 | 0.284 | -8.778 | 28.4
[6 empty] | -0.058 | 25 | 8.768 | 0.284 | -8.768 | 28.4
[6 empty] | -0.058 | 25 | 8.758 | 0.284 | -8.758 | 28.4
[6 empty] | -0.058 | 25 | 8.748 | 0.283 | -8.748 | 28.3
[6 empty] | -0.058 | 25 | 8.738 | 0.283 | -8.738 | 28.3
[6 empty] | -0.058 | 25 | 8.728 | 0.283 | -8.728 | 28.3
[6 empty] | -0.058 | 25 | 8.718 | 0.283 | -8.718 | 28.3
[6 empty] | -0.058 | 25 | 8.708 | 0.283 | -8.708 | 28.3
[6 empty] | -0.058 | 25 | 8.698 | 0.283 | -8.698 | 28.3
[6 empty] | -0.058 | 25 | 8.688 | 0.283 | -8.688 | 28.3
[6 empty] | -0.058 | 25 | 8.678 | 0.282 | -8.678 | 28.2
[6 empty] | -0.058 | 25 | 8.668 | 0.282 | -8.668 | 28.2
[6 empty] | -0.058 | 25 | 8.658 | 0.282 | -8.658 | 28.2
[6 empty] | -0.058 | 25 | 8.648 | 0.282 | -8.648 | 28.2
[6 empty] | -0.058 | 25 | 8.638 | 0.282 | -8.638 | 28.2
[6 empty] | -0.058 | 25 | 8.629 | 0.282 | -8.629 | 28.2
[6 empty] | -0.058 | 25 | 8.619 | 0.281 | -8.619 | 28.1
[6 empty] | -0.058 | 25 | 8.609 | 0.281 | -8.609 | 28.1
[6 empty] | -0.058 | 25 | 8.599 | 0.281 | -8.599 | 28.1
[6 empty] | -0.058 | 25 | 8.589 | 0.281 | -8.589 | 28.1
[6 empty] | -0.058 | 25 | 8.579 | 0.281 | -8.579 | 28.1
[6 empty] | -0.058 | 25 | 8.569 | 0.281 | -8.569 | 28.1
[6 empty] | -0.058 | 25 | 8.559 | 0.28 | -8.559 | 28
[6 empty] | -0.058 | 25 | 8.549 | 0.28 | -8.549 | 28
[6 empty] | -0.058 | 25 | 8.539 | 0.28 | -8.539 | 28
[6 empty] | -0.058 | 25 | 8.529 | 0.28 | -8.529 | 28
[6 empty] | -0.058 | 25 | 8.519 | 0.28 | -8.519 | 28
[6 empty] | -0.058 | 25 | 8.509 | 0.28 | -8.509 | 28
[6 empty] | -0.058 | 25 | 8.499 | 0.28 | -8.499 | 28
[6 empty] | -0.058 | 25 | 8.489 | 0.279 | -8.489 | 27.9
[6 empty] | -0.058 | 25 | 8.479 | 0.279 | -8.479 | 27.9
[6 empty] | -0.058 | 25 | 8.469 | 0.279 | -8.469 | 27.9
[6 empty] | -0.058 | 25 | 8.459 | 0.279 | -8.459 | 27.9
[6 empty] | -0.058 | 25 | 8.449 | 0.279 | -8.449 | 27.9
[6 empty] | -0.058 | 25 | 8.439 | 0.279 | -8.439 | 27.9
[6 empty] | -0.058 | 25 | 8.429 | 0.279 | -8.429 | 27.9
[6 empty] | -0.058 | 25 | 8.419 | 0.278 | -8.419 | 27.8
[6 empty] | -0.058 | 25 | 8.409 | 0.278 | -8.409 | 27.8
[6 empty] | -0.058 | 25 | 8.399 | 0.278 | -8.399 | 27.8
[6 empty] | -0.058 | 25 | 8.389 | 0.278 | -8.389 | 27.8
[6 empty] | -0.058 | 25 | 8.379 | 0.278 | -8.379 | 27.8
[6 empty] | -0.058 | 25 | 8.369 | 0.278 | -8.369 | 27.8
[6 empty] | -0.059 | 25 | 8.359 | 0.277 | -8.359 | 27.7
[6 empty] | -0.059 | 25 | 8.349 | 0.277 | -8.349 | 27.7
[6 empty] | -0.059 | 25 | 8.339 | 0.277 | -8.339 | 27.7
[6 empty] | -0.059 | 25 | 8.329 | 0.277 | -8.329 | 27.7
[6 empty] | -0.059 | 25 | 8.319 | 0.277 | -8.319 | 27.7
[6 empty] | -0.059 | 25 | 8.309 | 0.277 | -8.309 | 27.7
[6 empty] | -0.059 | 25 | 8.299 | 0.277 | -8.299 | 27.7
[6 empty] | -0.059 | 25 | 8.289 | 0.276 | -8.289 | 27.6
[6 empty] | -0.059 | 25 | 8.279 | 0.276 | -8.279 | 27.6
[6 empty] | -0.059 | 25 | 8.269 | 0.276 | -8.269 | 27.6
[6 empty] | -0.059 | 25 | 8.259 | 0.276 | -8.259 | 27.6
[6 empty] | -0.059 | 25 | 8.249 | 0.276 | -8.249 | 27.6
[6 empty] | -0.059 | 25 | 8.239 | 0.276 | -8.239 | 27.6
[6 empty] | -0.059 | 25 | 8.229 | 0.275 | -8.229 | 27.5
[6 empty] | -0.059 | 25 | 8.219 | 0.275 | -8.219 | 27.5
[6 empty] | -0.059 | 25 | 8.209 | 0.275 | -8.209 | 27.5
[6 empty] | -0.059 | 25 | 8.199 | 0.275 | -8.199 | 27.5
[6 empty] | -0.059 | 25 | 8.189 | 0.275 | -8.189 | 27.5
[6 empty] | -0.059 | 25 | 8.179 | 0.275 | -8.179 | 27.5
[6 empty] | -0.059 | 25 | 8.169 | 0.275 | -8.169 | 27.5
[6 empty] | -0.059 | 25 | 8.159 | 0.274 | -8.159 | 27.4
[6 empty] | -0.059 | 25 | 8.149 | 0.274 | -8.149 | 27.4
[6 empty] | -0.059 | 25 | 8.139 | 0.274 | -8.139 | 27.4
[6 empty] | -0.059 | 25 | 8.129 | 0.274 | -8.129 | 27.4
[6 empty] | -0.059 | 25 | 8.119 | 0.274 | -8.119 | 27.4
[6 empty] | -0.059 | 25 | 8.109 | 0.274 | -8.109 | 27.4
[6 empty] | -0.059 | 25 | 8.099 | 0.274 | -8.099 | 27.4
[6 empty] | -0.059 | 25 | 8.089 | 0.273 | -8.089 | 27.3
[6 empty] | -0.059 | 25 | 8.079 | 0.273 | -8.079 | 27.3
[6 empty] | -0.059 | 25 | 8.069 | 0.273 | -8.069 | 27.3
[6 empty] | -0.059 | 25 | 8.059 | 0.273 | -8.059 | 27.3
[6 empty] | -0.059 | 25 | 8.049 | 0.273 | -8.049 | 27.3
[6 empty] | -0.059 | 25 | 8.039 | 0.273 | -8.039 | 27.3
[6 empty] | -0.059 | 25 | 8.029 | 0.272 | -8.029 | 27.2
[6 empty] | -0.059 | 25 | 8.019 | 0.272 | -8.019 | 27.2
[6 empty] | -0.059 | 25 | 8.009 | 0.272 | -8.009 | 27.2
[6 empty] | -0.059 | 25 | 7.999 | 0.272 | -7.999 | 27.2
[6 empty] | -0.059 | 25 | 7.989 | 0.272 | -7.989 | 27.2
[6 empty] | -0.059 | 25 | 7.979 | 0.272 | -7.979 | 27.2
[6 empty] | -0.059 | 25 | 7.969 | 0.272 | -7.969 | 27.2
[6 empty] | -0.059 | 25 | 7.959 | 0.271 | -7.959 | 27.1
[6 empty] | -0.059 | 25 | 7.949 | 0.271 | -7.949 | 27.1
[6 empty] | -0.059 | 25 | 7.939 | 0.271 | -7.939 | 27.1
[6 empty] | -0.059 | 25 | 7.929 | 0.271 | -7.929 | 27.1
[6 empty] | -0.059 | 25 | 7.919 | 0.271 | -7.919 | 27.1
[6 empty] | -0.059 | 25 | 7.909 | 0.271 | -7.909 | 27.1
[6 empty] | -0.059 | 25 | 7.899 | 0.271 | -7.899 | 27.1
[6 empty] | -0.059 | 25 | 7.889 | 0.27 | -7.889 | 27
[6 empty] | -0.059 | 25 | 7.879 | 0.27 | -7.879 | 27
[6 empty] | -0.059 | 25 | 7.869 | 0.27 | -7.869 | 27
[6 empty] | -0.059 | 25 | 7.859 | 0.27 | -7.859 | 27
[6 empty] | -0.059 | 25 | 7.849 | 0.27 | -7.849 | 27
[6 empty] | -0.059 | 25 | 7.839 | 0.27 | -7.839 | 27
[6 empty] | -0.059 | 25 | 7.829 | 0.269 | -7.829 | 26.9
[6 empty] | -0.059 | 25 | 7.819 | 0.269 | -7.819 | 26.9
[6 empty] | -0.059 | 25 | 7.809 | 0.269 | -7.809 | 26.9
[6 empty] | -0.059 | 25 | 7.799 | 0.269 | -7.799 | 26.9
[6 empty] | -0.059 | 25 | 7.789 | 0.269 | -7.789 | 26.9
[6 empty] | -0.059 | 25 | 7.779 | 0.269 | -7.779 | 26.9
[6 empty] | -0.059 | 25 | 7.769 | 0.269 | -7.769 | 26.9
[6 empty] | -0.059 | 25 | 7.759 | 0.268 | -7.759 | 26.8
[6 empty] | -0.059 | 25 | 7.749 | 0.268 | -7.749 | 26.8
[6 empty] | -0.059 | 25 | 7.739 | 0.268 | -7.739 | 26.8
[6 empty] | -0.059 | 25 | 7.729 | 0.268 | -7.729 | 26.8
[6 empty] | -0.059 | 25 | 7.719 | 0.268 | -7.719 | 26.8
[6 empty] | -0.059 | 25 | 7.709 | 0.268 | -7.709 | 26.8
[6 empty] | -0.059 | 25 | 7.699 | 0.268 | -7.699 | 26.8
[6 empty] | -0.059 | 25 | 7.689 | 0.267 | -7.689 | 26.7
[6 empty] | -0.059 | 25 | 7.679 | 0.267 | -7.679 | 26.7
[6 empty] | -0.059 | 25 | 7.669 | 0.267 | -7.669 | 26.7
[6 empty] | -0.06 | 25 | 7.659 | 0.267 | -7.659 | 26.7
[6 empty] | -0.06 | 25 | 7.649 | 0.267 | -7.649 | 26.7
[6 empty] | -0.06 | 25 | 7.639 | 0.267 | -7.639 | 26.7
[6 empty] | -0.06 | 25 | 7.629 | 0.267 | -7.629 | 26.7
[6 empty] | -0.06 | 25 | 7.619 | 0.266 | -7.619 | 26.6
[6 empty] | -0.06 | 25 | 7.609 | 0.266 | -7.609 | 26.6
[6 empty] | -0.06 | 25 | 7.599 | 0.266 | -7.599 | 26.6
[6 empty] | -0.06 | 25 | 7.589 | 0.266 | -7.589 | 26.6
[6 empty] | -0.06 | 25 | 7.579 | 0.266 | -7.579 | 26.6
[6 empty] | -0.06 | 25 | 7.569 | 0.266 | -7.569 | 26.6
[6 empty] | -0.06 | 25 | 7.559 | 0.265 | -7.559 | 26.5
[6 empty] | -0.06 | 25 | 7.549 | 0.265 | -7.549 | 26.5
[6 empty] | -0.06 | 25 | 7.539 | 0.265 | -7.539 | 26.5
[6 empty] | -0.06 | 25 | 7.529 | 0.265 | -7.529 | 26.5
[6 empty] | -0.06 | 25 | 7.519 | 0.265 | -7.519 | 26.5
[6 empty] | -0.06 | 25 | 7.509 | 0.265 | -7.509 | 26.5
[6 empty] | -0.06 | 25 | 7.499 | 0.265 | -7.499 | 26.5
[6 empty] | -0.06 | 25 | 7.489 | 0.264 | -7.489 | 26.4
[6 empty] | -0.06 | 25 | 7.479 | 0.264 | -7.479 | 26.4
[6 empty] | -0.06 | 25 | 7.469 | 0.264 | -7.469 | 26.4
[6 empty] | -0.06 | 25 | 7.459 | 0.264 | -7.459 | 26.4
[6 empty] | -0.06 | 25 | 7.449 | 0.264 | -7.449 | 26.4
[6 empty] | -0.06 | 25 | 7.439 | 0.264 | -7.439 | 26.4
[6 empty] | -0.06 | 25 | 7.429 | 0.264 | -7.429 | 26.4
[6 empty] | -0.06 | 25 | 7.419 | 0.263 | -7.419 | 26.3
[6 empty] | -0.06 | 25 | 7.409 | 0.263 | -7.409 | 26.3
[6 empty] | -0.06 | 25 | 7.399 | 0.263 | -7.399 | 26.3
[6 empty] | -0.06 | 25 | 7.389 | 0.263 | -7.389 | 26.3
[6 empty] | -0.06 | 25 | 7.379 | 0.263 | -7.379 | 26.3
[6 empty] | -0.06 | 25 | 7.369 | 0.263 | -7.369 | 26.3
[6 empty] | -0.06 | 25 | 7.359 | 0.263 | -7.359 | 26.3
[6 empty] | -0.06 | 25 | 7.349 | 0.262 | -7.349 | 26.2
[6 empty] | -0.06 | 25 | 7.339 | 0.262 | -7.339 | 26.2
[6 empty] | -0.06 | 25 | 7.329 | 0.262 | -7.329 | 26.2
[6 empty] | -0.06 | 25 | 7.319 | 0.262 | -7.319 | 26.2
[6 empty] | -0.06 | 25 | 7.309 | 0.262 | -7.309 | 26.2
[6 empty] | -0.06 | 25 | 7.299 | 0.262 | -7.299 | 26.2
[6 empty] | -0.06 | 25 | 7.289 | 0.261 | -7.289 | 26.1
[6 empty] | -0.06 | 25 | 7.279 | 0.261 | -7.279 | 26.1
[6 empty] | -0.06 | 25 | 7.269 | 0.261 | -7.269 | 26.1
[6 empty] | -0.06 | 25 | 7.259 | 0.261 | -7.259 | 26.1
[6 empty] | -0.06 | 25 | 7.249 | 0.261 | -7.249 | 26.1
[6 empty] | -0.06 | 25 | 7.239 | 0.261 | -7.239 | 26.1
[6 empty] | -0.06 | 25 | 7.229 | 0.261 | -7.229 | 26.1
[6 empty] | -0.06 | 25 | 7.219 | 0.26 | -7.219 | 26
[6 empty] | -0.06 | 25 | 7.209 | 0.26 | -7.209 | 26
[6 empty] | -0.06 | 25 | 7.199 | 0.26 | -7.199 | 26
[6 empty] | -0.06 | 25 | 7.189 | 0.26 | -7.189 | 26
[6 empty] | -0.06 | 25 | 7.179 | 0.26 | -7.179 | 26
[6 empty] | -0.06 | 25 | 7.169 | 0.26 | -7.169 | 26
[6 empty] | -0.06 | 25 | 7.159 | 0.26 | -7.159 | 26
[6 empty] | -0.06 | 25 | 7.149 | 0.259 | -7.149 | 25.9
[6 empty] | -0.06 | 25 | 7.139 | 0.259 | -7.139 | 25.9
[6 empty] | -0.06 | 25 | 7.129 | 0.259 | -7.129 | 25.9
[6 empty] | -0.06 | 25 | 7.119 | 0.259 | -7.119 | 25.9
[6 empty] | -0.06 | 25 | 7.109 | 0.259 | -7.109 | 25.9
[6 empty] | -0.06 | 25 | 7.099 | 0.259 | -7.099 | 25.9
[6 empty] | -0.06 | 25 | 7.089 | 0.259 | -7.089 | 25.9
[6 empty] | -0.06 | 25 | 7.079 | 0.258 | -7.079 | 25.8
[6 empty] | -0.06 | 25 | 7.069 | 0.258 | -7.069 | 25.8
[6 empty] | -0.06 | 25 | 7.059 | 0.258 | -7.059 | 25.8
[6 empty] | -0.06 | 25 | 7.049 | 0.258 | -7.049 | 25.8
[6 empty] | -0.06 | 25 | 7.039 | 0.258 | -7.039 | 25.8
[6 empty] | -0.06 | 25 | 7.029 | 0.258 | -7.029 | 25.8
[6 empty] | -0.06 | 25 | 7.019 | 0.258 | -7.019 | 25.8
[6 empty] | -0.06 | 25 | 7.009 | 0.257 | -7.009 | 25.7
[6 empty] | -0.06 | 25 | 6.999 | 0.257 | -6.999 | 25.7
[6 empty] | -0.06 | 25 | 6.989 | 0.257 | -6.989 | 25.7
[6 empty] | -0.06 | 25 | 6.979 | 0.257 | -6.979 | 25.7
[6 empty] | -0.06 | 25 | 6.969 | 0.257 | -6.969 | 25.7
[6 empty] | -0.061 | 25 | 6.959 | 0.257 | -6.959 | 25.7
[6 empty] | -0.061 | 25 | 6.949 | 0.257 | -6.949 | 25.7
[6 empty] | -0.061 | 25 | 6.939 | 0.256 | -6.939 | 25.6
[6 empty] | -0.061 | 25 | 6.929 | 0.256 | -6.929 | 25.6
[6 empty] | -0.061 | 25 | 6.919 | 0.256 | -6.919 | 25.6
[6 empty] | -0.061 | 25 | 6.909 | 0.256 | -6.909 | 25.6
[6 empty] | -0.061 | 25 | 6.899 | 0.256 | -6.899 | 25.6
[6 empty] | -0.061 | 25 | 6.889 | 0.256 | -6.889 | 25.6
[6 empty] | -0.061 | 25 | 6.879 | 0.256 | -6.879 | 25.6
[6 empty] | -0.061 | 25 | 6.869 | 0.255 | -6.869 | 25.5
[6 empty] | -0.061 | 25 | 6.859 | 0.255 | -6.859 | 25.5
[6 empty] | -0.061 | 25 | 6.849 | 0.255 | -6.849 | 25.5
[6 empty] | -0.061 | 25 | 6.84 | 0.255 | -6.84 | 25.5
[6 empty] | -0.061 | 25 | 6.83 | 0.255 | -6.83 | 25.5
[6 empty] | -0.061 | 25 | 6.82 | 0.255 | -6.82 | 25.5
[6 empty] | -0.061 | 25 | 6.81 | 0.255 | -6.81 | 25.5
[6 empty] | -0.061 | 25 | 6.8 | 0.254 | -6.8 | 25.4
[6 empty] | -0.061 | 25 | 6.79 | 0.254 | -6.79 | 25.4
[6 empty] | -0.061 | 25 | 6.78 | 0.254 | -6.78 | 25.4
[6 empty] | -0.061 | 25 | 6.77 | 0.254 | -6.77 | 25.4
[6 empty] | -0.061 | 25 | 6.76 | 0.254 | -6.76 | 25.4
[6 empty] | -0.061 | 25 | 6.75 | 0.254 | -6.75 | 25.4
[6 empty] | -0.061 | 25 | 6.74 | 0.254 | -6.74 | 25.4
[6 empty] | -0.061 | 25 | 6.73 | 0.253 | -6.73 | 25.3
[6 empty] | -0.061 | 25 | 6.72 | 0.253 | -6.72 | 25.3
[6 empty] | -0.061 | 25 | 6.71 | 0.253 | -6.71 | 25.3
[6 empty] | -0.061 | 25 | 6.7 | 0.253 | -6.7 | 25.3
[6 empty] | -0.061 | 25 | 6.69 | 0.253 | -6.69 | 25.3
[6 empty] | -0.061 | 25 | 6.68 | 0.253 | -6.68 | 25.3
[6 empty] | -0.061 | 25 | 6.67 | 0.253 | -6.67 | 25.3
[6 empty] | -0.061 | 25 | 6.66 | 0.252 | -6.66 | 25.2
[6 empty] | -0.061 | 25 | 6.65 | 0.252 | -6.65 | 25.2
[6 empty] | -0.061 | 25 | 6.64 | 0.252 | -6.64 | 25.2
[6 empty] | -0.061 | 25 | 6.63 | 0.252 | -6.63 | 25.2
[6 empty] | -0.061 | 25 | 6.62 | 0.252 | -6.62 | 25.2
[6 empty] | -0.061 | 25 | 6.61 | 0.252 | -6.61 | 25.2
[6 empty] | -0.061 | 25 | 6.6 | 0.252 | -6.6 | 25.2
[6 empty] | -0.061 | 25 | 6.59 | 0.251 | -6.59 | 25.1
[6 empty] | -0.061 | 25 | 6.58 | 0.251 | -6.58 | 25.1
[6 empty] | -0.061 | 25 | 6.57 | 0.251 | -6.57 | 25.1
[6 empty] | -0.061 | 25 | 6.56 | 0.251 | -6.56 | 25.1
[6 empty] | -0.061 | 25 | 6.55 | 0.251 | -6.55 | 25.1
[6 empty] | -0.061 | 25 | 6.54 | 0.251 | -6.54 | 25.1
[6 empty] | -0.061 | 25 | 6.53 | 0.251 | -6.53 | 25.1
[6 empty] | -0.061 | 25 | 6.52 | 0.25 | -6.52 | 25
[6 empty] | -0.061 | 25 | 6.51 | 0.25 | -6.51 | 25
[6 empty] | -0.061 | 25 | 6.5 | 0.25 | -6.5 | 25
[6 empty] | -0.061 | 25 | 6.49 | 0.25 | -6.49 | 25
[6 empty] | -0.061 | 25 | 6.48 | 0.25 | -6.48 | 25
[6 empty] | -0.061 | 25 | 6.47 | 0.25 | -6.47 | 25
[6 empty] | -0.061 | 25 | 6.46 | 0.25 | -6.46 | 25
[6 empty] | -0.061 | 25 | 6.45 | 0.249 | -6.45 | 24.9
[6 empty] | -0.061 | 25 | 6.44 | 0.249 | -6.44 | 24.9
[6 empty] | -0.061 | 25 | 6.43 | 0.249 | -6.43 | 24.9
[6 empty] | -0.061 | 25 | 6.42 | 0.249 | -6.42 | 24.9
[6 empty] | -0.061 | 25 | 6.41 | 0.249 | -6.41 | 24.9
[6 empty] | -0.061 | 25 | 6.4 | 0.249 | -6.4 | 24.9
[6 empty] | -0.061 | 25 | 6.39 | 0.249 | -6.39 | 24.9
[6 empty] | -0.061 | 25 | 6.38 | 0.249 | -6.38 | 24.9
[6 empty] | -0.061 | 25 | 6.37 | 0.248 | -6.37 | 24.8
[6 empty] | -0.061 | 25 | 6.36 | 0.248 | -6.36 | 24.8
[6 empty] | -0.061 | 25 | 6.35 | 0.248 | -6.35 | 24.8
[6 empty] | -0.061 | 25 | 6.34 | 0.248 | -6.34 | 24.8
[6 empty] | -0.061 | 25 | 6.33 | 0.248 | -6.33 | 24.8
[6 empty] | -0.061 | 25 | 6.32 | 0.248 | -6.32 | 24.8
[6 empty] | -0.061 | 25 | 6.31 | 0.248 | -6.31 | 24.8
[6 empty] | -0.061 | 25 | 6.3 | 0.247 | -6.3 | 24.7
[6 empty] | -0.061 | 25 | 6.29 | 0.247 | -6.29 | 24.7
[6 empty] | -0.061 | 25 | 6.28 | 0.247 | -6.28 | 24.7
[6 empty] | -0.061 | 25 | 6.27 | 0.247 | -6.27 | 24.7
[6 empty] | -0.062 | 25 | 6.26 | 0.247 | -6.26 | 24.7
[6 empty] | -0.062 | 25 | 6.25 | 0.247 | -6.25 | 24.7
[6 empty] | -0.062 | 25 | 6.24 | 0.247 | -6.24 | 24.7
[6 empty] | -0.062 | 25 | 6.23 | 0.246 | -6.23 | 24.6
[6 empty] | -0.062 | 25 | 6.22 | 0.246 | -6.22 | 24.6
[6 empty] | -0.062 | 25 | 6.21 | 0.246 | -6.21 | 24.6
[6 empty] | -0.062 | 25 | 6.2 | 0.246 | -6.2 | 24.6
[6 empty] | -0.062 | 25 | 6.19 | 0.246 | -6.19 | 24.6
[6 empty] | -0.062 | 25 | 6.18 | 0.246 | -6.18 | 24.6
[6 empty] | -0.062 | 25 | 6.17 | 0.246 | -6.17 | 24.6
[6 empty] | -0.062 | 25 | 6.16 | 0.245 | -6.16 | 24.5
[6 empty] | -0.062 | 25 | 6.15 | 0.245 | -6.15 | 24.5
[6 empty] | -0.062 | 25 | 6.14 | 0.245 | -6.14 | 24.5
[6 empty] | -0.062 | 25 | 6.13 | 0.245 | -6.13 | 24.5
[6 empty] | -0.062 | 25 | 6.12 | 0.245 | -6.12 | 24.5
[6 empty] | -0.062 | 25 | 6.11 | 0.245 | -6.11 | 24.5
[6 empty] | -0.062 | 25 | 6.1 | 0.245 | -6.1 | 24.5
[6 empty] | -0.062 | 25 | 6.09 | 0.245 | -6.09 | 24.5
[6 empty] | -0.062 | 25 | 6.08 | 0.244 | -6.08 | 24.4
[6 empty] | -0.062 | 25 | 6.07 | 0.244 | -6.07 | 24.4
[6 empty] | -0.062 | 25 | 6.06 | 0.244 | -6.06 | 24.4
[6 empty] | -0.062 | 25 | 6.05 | 0.244 | -6.05 | 24.4
[6 empty] | -0.062 | 25 | 6.04 | 0.244 | -6.04 | 24.4
[6 empty] | -0.062 | 25 | 6.03 | 0.244 | -6.03 | 24.4
[6 empty] | -0.062 | 25 | 6.02 | 0.244 | -6.02 | 24.4
[6 empty] | -0.062 | 25 | 6.01 | 0.243 | -6.01 | 24.3
[6 empty] | -0.062 | 25 | 6 | 0.243 | -6 | 24.3
[6 empty] | -0.062 | 25 | 5.99 | 0.243 | -5.99 | 24.3
[6 empty] | -0.062 | 25 | 5.98 | 0.243 | -5.98 | 24.3
[6 empty] | -0.062 | 25 | 5.97 | 0.243 | -5.97 | 24.3
[6 empty] | -0.062 | 25 | 5.96 | 0.243 | -5.96 | 24.3
[6 empty] | -0.062 | 25 | 5.95 | 0.243 | -5.95 | 24.3
[6 empty] | -0.062 | 25 | 5.94 | 0.243 | -5.94 | 24.3
[6 empty] | -0.062 | 25 | 5.93 | 0.242 | -5.93 | 24.2
[6 empty] | -0.062 | 25 | 5.92 | 0.242 | -5.92 | 24.2
[6 empty] | -0.062 | 25 | 5.91 | 0.242 | -5.91 | 24.2
[6 empty] | -0.062 | 25 | 5.9 | 0.242 | -5.9 | 24.2
[6 empty] | -0.062 | 25 | 5.89 | 0.242 | -5.89 | 24.2
[6 empty] | -0.062 | 25 | 5.88 | 0.242 | -5.88 | 24.2
[6 empty] | -0.062 | 25 | 5.87 | 0.242 | -5.87 | 24.2
[6 empty] | -0.062 | 25 | 5.86 | 0.241 | -5.86 | 24.1
[6 empty] | -0.062 | 25 | 5.85 | 0.241 | -5.85 | 24.1
[6 empty] | -0.062 | 25 | 5.84 | 0.241 | -5.84 | 24.1
[6 empty] | -0.062 | 25 | 5.83 | 0.241 | -5.83 | 24.1
[6 empty] | -0.062 | 25 | 5.82 | 0.241 | -5.82 | 24.1
[6 empty] | -0.062 | 25 | 5.81 | 0.241 | -5.81 | 24.1
[6 empty] | -0.062 | 25 | 5.8 | 0.241 | -5.8 | 24.1
[6 empty] | -0.062 | 25 | 5.79 | 0.241 | -5.79 | 24.1
[6 empty] | -0.062 | 25 | 5.78 | 0.24 | -5.78 | 24
[6 empty] | -0.062 | 25 | 5.77 | 0.24 | -5.77 | 24
[6 empty] | -0.062 | 25 | 5.76 | 0.24 | -5.76 | 24
[6 empty] | -0.062 | 25 | 5.75 | 0.24 | -5.75 | 24
[6 empty] | -0.062 | 25 | 5.74 | 0.24 | -5.74 | 24
[6 empty] | -0.062 | 25 | 5.73 | 0.24 | -5.73 | 24
[6 empty] | -0.062 | 25 | 5.72 | 0.24 | -5.72 | 24
[6 empty] | -0.062 | 25 | 5.71 | 0.239 | -5.71 | 23.9
[6 empty] | -0.062 | 25 | 5.7 | 0.239 | -5.7 | 23.9
[6 empty] | -0.062 | 25 | 5.69 | 0.239 | -5.69 | 23.9
[6 empty] | -0.062 | 25 | 5.68 | 0.239 | -5.68 | 23.9
[6 empty] | -0.062 | 25 | 5.67 | 0.239 | -5.67 | 23.9
[6 empty] | -0.062 | 25 | 5.66 | 0.239 | -5.66 | 23.9
[6 empty] | -0.062 | 25 | 5.65 | 0.239 | -5.65 | 23.9
[6 empty] | -0.062 | 25 | 5.64 | 0.239 | -5.64 | 23.9
[6 empty] | -0.062 | 25 | 5.63 | 0.238 | -5.63 | 23.8
[6 empty] | -0.062 | 25 | 5.62 | 0.238 | -5.62 | 23.8
[6 empty] | -0.062 | 25 | 5.61 | 0.238 | -5.61 | 23.8
[6 empty] | -0.062 | 25 | 5.6 | 0.238 | -5.6 | 23.8
[6 empty] | -0.062 | 25 | 5.59 | 0.238 | -5.59 | 23.8
[6 empty] | -0.062 | 25 | 5.58 | 0.238 | -5.58 | 23.8
[6 empty] | -0.062 | 25 | 5.57 | 0.238 | -5.57 | 23.8
[6 empty] | -0.063 | 25 | 5.56 | 0.237 | -5.56 | 23.7
[6 empty] | -0.063 | 25 | 5.55 | 0.237 | -5.55 | 23.7
[6 empty] | -0.063 | 25 | 5.54 | 0.237 | -5.54 | 23.7
[6 empty] | -0.063 | 25 | 5.53 | 0.237 | -5.53 | 23.7
[6 empty] | -0.063 | 25 | 5.52 | 0.237 | -5.52 | 23.7
[6 empty] | -0.063 | 25 | 5.51 | 0.237 | -5.51 | 23.7
[6 empty] | -0.063 | 25 | 5.5 | 0.237 | -5.5 | 23.7
[6 empty] | -0.063 | 25 | 5.49 | 0.237 | -5.49 | 23.7
[6 empty] | -0.063 | 25 | 5.48 | 0.236 | -5.48 | 23.6
[6 empty] | -0.063 | 25 | 5.47 | 0.236 | -5.47 | 23.6
[6 empty] | -0.063 | 25 | 5.46 | 0.236 | -5.46 | 23.6
[6 empty] | -0.063 | 25 | 5.45 | 0.236 | -5.45 | 23.6
[6 empty] | -0.063 | 25 | 5.44 | 0.236 | -5.44 | 23.6
[6 empty] | -0.063 | 25 | 5.43 | 0.236 | -5.43 | 23.6
[6 empty] | -0.063 | 25 | 5.42 | 0.236 | -5.42 | 23.6
[6 empty] | -0.063 | 25 | 5.41 | 0.236 | -5.41 | 23.6
[6 empty] | -0.063 | 25 | 5.4 | 0.235 | -5.4 | 23.5
[6 empty] | -0.063 | 25 | 5.39 | 0.235 | -5.39 | 23.5
[6 empty] | -0.063 | 25 | 5.38 | 0.235 | -5.38 | 23.5
[6 empty] | -0.063 | 25 | 5.37 | 0.235 | -5.37 | 23.5
[6 empty] | -0.063 | 25 | 5.36 | 0.235 | -5.36 | 23.5
[6 empty] | -0.063 | 25 | 5.35 | 0.235 | -5.35 | 23.5
[6 empty] | -0.063 | 25 | 5.34 | 0.235 | -5.34 | 23.5
[6 empty] | -0.063 | 25 | 5.33 | 0.234 | -5.33 | 23.4
[6 empty] | -0.063 | 25 | 5.32 | 0.234 | -5.32 | 23.4
[6 empty] | -0.063 | 25 | 5.31 | 0.234 | -5.31 | 23.4
[6 empty] | -0.063 | 25 | 5.3 | 0.234 | -5.3 | 23.4
[6 empty] | -0.063 | 25 | 5.29 | 0.234 | -5.29 | 23.4
[6 empty] | -0.063 | 25 | 5.28 | 0.234 | -5.28 | 23.4
[6 empty] | -0.063 | 25 | 5.27 | 0.234 | -5.27 | 23.4
[6 empty] | -0.063 | 25 | 5.26 | 0.234 | -5.26 | 23.4
[6 empty] | -0.063 | 25 | 5.25 | 0.233 | -5.25 | 23.3
[6 empty] | -0.063 | 25 | 5.24 | 0.233 | -5.24 | 23.3
[6 empty] | -0.063 | 25 | 5.23 | 0.233 | -5.23 | 23.3
[6 empty] | -0.063 | 25 | 5.22 | 0.233 | -5.22 | 23.3
[6 empty] | -0.063 | 25 | 5.21 | 0.233 | -5.21 | 23.3
[6 empty] | -0.063 | 25 | 5.2 | 0.233 | -5.2 | 23.3
[6 empty] | -0.063 | 25 | 5.19 | 0.233 | -5.19 | 23.3
[6 empty] | -0.063 | 25 | 5.18 | 0.233 | -5.18 | 23.3
[6 empty] | -0.063 | 25 | 5.17 | 0.232 | -5.17 | 23.2
[6 empty] | -0.063 | 25 | 5.16 | 0.232 | -5.16 | 23.2
[6 empty] | -0.063 | 25 | 5.15 | 0.232 | -5.15 | 23.2
[6 empty] | -0.063 | 25 | 5.14 | 0.232 | -5.14 | 23.2
[6 empty] | -0.063 | 25 | 5.13 | 0.232 | -5.13 | 23.2
[6 empty] | -0.063 | 25 | 5.12 | 0.232 | -5.12 | 23.2
[6 empty] | -0.063 | 25 | 5.11 | 0.232 | -5.11 | 23.2
[6 empty] | -0.063 | 25 | 5.1 | 0.232 | -5.1 | 23.2
[6 empty] | -0.063 | 25 | 5.09 | 0.231 | -5.09 | 23.1
[6 empty] | -0.063 | 25 | 5.08 | 0.231 | -5.08 | 23.1
[6 empty] | -0.063 | 25 | 5.07 | 0.231 | -5.07 | 23.1
[6 empty] | -0.063 | 25 | 5.06 | 0.231 | -5.06 | 23.1
[6 empty] | -0.063 | 25 | 5.051 | 0.231 | -5.051 | 23.1
[6 empty] | -0.063 | 25 | 5.041 | 0.231 | -5.041 | 23.1
[6 empty] | -0.063 | 25 | 5.031 | 0.231 | -5.031 | 23.1
[6 empty] | -0.063 | 25 | 5.021 | 0.23 | -5.021 | 23
[6 empty] | -0.063 | 25 | 5.011 | 0.23 | -5.011 | 23
[6 empty] | -0.063 | 25 | 5.001 | 0.23 | -5.001 | 23
[6 empty] | -0.063 | 25 | 4.991 | 0.23 | -4.991 | 23
[6 empty] | -0.063 | 25 | 4.981 | 0.23 | -4.981 | 23
[6 empty] | -0.063 | 25 | 4.971 | 0.23 | -4.971 | 23
[6 empty] | -0.063 | 25 | 4.961 | 0.23 | -4.961 | 23
[6 empty] | -0.063 | 25 | 4.951 | 0.23 | -4.951 | 23
[6 empty] | -0.063 | 25 | 4.941 | 0.229 | -4.941 | 22.9
[6 empty] | -0.063 | 25 | 4.931 | 0.229 | -4.931 | 22.9
[6 empty] | -0.063 | 25 | 4.921 | 0.229 | -4.921 | 22.9
[6 empty] | -0.063 | 25 | 4.911 | 0.229 | -4.911 | 22.9
[6 empty] | -0.063 | 25 | 4.901 | 0.229 | -4.901 | 22.9
[6 empty] | -0.063 | 25 | 4.891 | 0.229 | -4.891 | 22.9
[6 empty] | -0.063 | 25 | 4.881 | 0.229 | -4.881 | 22.9
[6 empty] | -0.063 | 25 | 4.871 | 0.229 | -4.871 | 22.9
[6 empty] | -0.064 | 25 | 4.861 | 0.228 | -4.861 | 22.8
[6 empty] | -0.064 | 25 | 4.851 | 0.228 | -4.851 | 22.8
[6 empty] | -0.064 | 25 | 4.841 | 0.228 | -4.841 | 22.8
[6 empty] | -0.064 | 25 | 4.831 | 0.228 | -4.831 | 22.8
[6 empty] | -0.064 | 25 | 4.821 | 0.228 | -4.821 | 22.8
[6 empty] | -0.064 | 25 | 4.811 | 0.228 | -4.811 | 22.8
[6 empty] | -0.064 | 25 | 4.801 | 0.228 | -4.801 | 22.8
[6 empty] | -0.064 | 25 | 4.791 | 0.228 | -4.791 | 22.8
[6 empty] | -0.064 | 25 | 4.781 | 0.227 | -4.781 | 22.7
[6 empty] | -0.064 | 25 | 4.771 | 0.227 | -4.771 | 22.7
[6 empty] | -0.064 | 25 | 4.761 | 0.227 | -4.761 | 22.7
[6 empty] | -0.064 | 25 | 4.751 | 0.227 | -4.751 | 22.7
[6 empty] | -0.064 | 25 | 4.741 | 0.227 | -4.741 | 22.7
[6 empty] | -0.064 | 25 | 4.731 | 0.227 | -4.731 | 22.7
[6 empty] | -0.064 | 25 | 4.721 | 0.227 | -4.721 | 22.7
[6 empty] | -0.064 | 25 | 4.711 | 0.227 | -4.711 | 22.7
[6 empty] | -0.064 | 25 | 4.701 | 0.226 | -4.701 | 22.6
[6 empty] | -0.064 | 25 | 4.691 | 0.226 | -4.691 | 22.6
[6 empty] | -0.064 | 25 | 4.681 | 0.226 | -4.681 | 22.6
[6 empty] | -0.064 | 25 | 4.671 | 0.226 | -4.671 | 22.6
[6 empty] | -0.064 | 25 | 4.661 | 0.226 | -4.661 | 22.6
[6 empty] | -0.064 | 25 | 4.651 | 0.226 | -4.651 | 22.6
[6 empty] | -0.064 | 25 | 4.641 | 0.226 | -4.641 | 22.6
[6 empty] | -0.064 | 25 | 4.631 | 0.226 | -4.631 | 22.6
[6 empty] | -0.064 | 25 | 4.621 | 0.225 | -4.621 | 22.5
[6 empty] | -0.064 | 25 | 4.611 | 0.225 | -4.611 | 22.5
[6 empty] | -0.064 | 25 | 4.601 | 0.225 | -4.601 | 22.5
[6 empty] | -0.064 | 25 | 4.591 | 0.225 | -4.591 | 22.5
[6 empty] | -0.064 | 25 | 4.581 | 0.225 | -4.581 | 22.5
[6 empty] | -0.064 | 25 | 4.571 | 0.225 | -4.571 | 22.5
[6 empty] | -0.064 | 25 | 4.561 | 0.225 | -4.561 | 22.5
[6 empty] | -0.064 | 25 | 4.551 | 0.225 | -4.551 | 22.5
[6 empty] | -0.064 | 25 | 4.541 | 0.224 | -4.541 | 22.4
[6 empty] | -0.064 | 25 | 4.531 | 0.224 | -4.531 | 22.4
[6 empty] | -0.064 | 25 | 4.521 | 0.224 | -4.521 | 22.4
[6 empty] | -0.064 | 25 | 4.511 | 0.224 | -4.511 | 22.4
[6 empty] | -0.064 | 25 | 4.501 | 0.224 | -4.501 | 22.4
[6 empty] | -0.064 | 25 | 4.491 | 0.224 | -4.491 | 22.4
[6 empty] | -0.064 | 25 | 4.481 | 0.224 | -4.481 | 22.4
[6 empty] | -0.064 | 25 | 4.471 | 0.224 | -4.471 | 22.4
[6 empty] | -0.064 | 25 | 4.461 | 0.223 | -4.461 | 22.3
[6 empty] | -0.064 | 25 | 4.451 | 0.223 | -4.451 | 22.3
[6 empty] | -0.064 | 25 | 4.441 | 0.223 | -4.441 | 22.3
[6 empty] | -0.064 | 25 | 4.431 | 0.223 | -4.431 | 22.3
[6 empty] | -0.064 | 25 | 4.421 | 0.223 | -4.421 | 22.3
[6 empty] | -0.064 | 25 | 4.411 | 0.223 | -4.411 | 22.3
[6 empty] | -0.064 | 25 | 4.401 | 0.223 | -4.401 | 22.3
[6 empty] | -0.064 | 25 | 4.391 | 0.223 | -4.391 | 22.3
[6 empty] | -0.064 | 25 | 4.381 | 0.222 | -4.381 | 22.2
[6 empty] | -0.064 | 25 | 4.371 | 0.222 | -4.371 | 22.2
[6 empty] | -0.064 | 25 | 4.361 | 0.222 | -4.361 | 22.2
[6 empty] | -0.064 | 25 | 4.351 | 0.222 | -4.351 | 22.2
[6 empty] | -0.064 | 25 | 4.341 | 0.222 | -4.341 | 22.2
[6 empty] | -0.064 | 25 | 4.331 | 0.222 | -4.331 | 22.2
[6 empty] | -0.064 | 25 | 4.321 | 0.222 | -4.321 | 22.2
[6 empty] | -0.064 | 25 | 4.311 | 0.222 | -4.311 | 22.2
[6 empty] | -0.064 | 25 | 4.301 | 0.222 | -4.301 | 22.2
[6 empty] | -0.064 | 25 | 4.291 | 0.221 | -4.291 | 22.1
[6 empty] | -0.064 | 25 | 4.281 | 0.221 | -4.281 | 22.1
[6 empty] | -0.064 | 25 | 4.271 | 0.221 | -4.271 | 22.1
[6 empty] | -0.064 | 25 | 4.261 | 0.221 | -4.261 | 22.1
[6 empty] | -0.064 | 25 | 4.251 | 0.221 | -4.251 | 22.1
[6 empty] | -0.064 | 25 | 4.241 | 0.221 | -4.241 | 22.1
[6 empty] | -0.064 | 25 | 4.231 | 0.221 | -4.231 | 22.1
[6 empty] | -0.064 | 25 | 4.221 | 0.221 | -4.221 | 22.1
[6 empty] | -0.064 | 25 | 4.211 | 0.22 | -4.211 | 22
[6 empty] | -0.064 | 25 | 4.201 | 0.22 | -4.201 | 22
[6 empty] | -0.064 | 25 | 4.191 | 0.22 | -4.191 | 22
[6 empty] | -0.064 | 25 | 4.181 | 0.22 | -4.181 | 22
[6 empty] | -0.064 | 25 | 4.171 | 0.22 | -4.171 | 22
[6 empty] | -0.065 | 25 | 4.161 | 0.22 | -4.161 | 22
[6 empty] | -0.065 | 25 | 4.151 | 0.22 | -4.151 | 22
[6 empty] | -0.065 | 25 | 4.141 | 0.22 | -4.141 | 22
[6 empty] | -0.065 | 25 | 4.131 | 0.219 | -4.131 | 21.9
[6 empty] | -0.065 | 25 | 4.121 | 0.219 | -4.121 | 21.9
[6 empty] | -0.065 | 25 | 4.111 | 0.219 | -4.111 | 21.9
[6 empty] | -0.065 | 25 | 4.101 | 0.219 | -4.101 | 21.9
[6 empty] | -0.065 | 25 | 4.091 | 0.219 | -4.091 | 21.9
[6 empty] | -0.065 | 25 | 4.081 | 0.219 | -4.081 | 21.9
[6 empty] | -0.065 | 25 | 4.071 | 0.219 | -4.071 | 21.9
[6 empty] | -0.065 | 25 | 4.061 | 0.219 | -4.061 | 21.9
[6 empty] | -0.065 | 25 | 4.051 | 0.219 | -4.051 | 21.9
[6 empty] | -0.065 | 25 | 4.041 | 0.218 | -4.041 | 21.8
[6 empty] | -0.065 | 25 | 4.031 | 0.218 | -4.031 | 21.8
[6 empty] | -0.065 | 25 | 4.021 | 0.218 | -4.021 | 21.8
[6 empty] | -0.065 | 25 | 4.011 | 0.218 | -4.011 | 21.8
[6 empty] | -0.065 | 25 | 4.001 | 0.218 | -4.001 | 21.8
[6 empty] | -0.065 | 25 | 3.991 | 0.218 | -3.991 | 21.8
[6 empty] | -0.065 | 25 | 3.981 | 0.218 | -3.981 | 21.8
[6 empty] | -0.065 | 25 | 3.971 | 0.218 | -3.971 | 21.8
[6 empty] | -0.065 | 25 | 3.961 | 0.217 | -3.961 | 21.7
[6 empty] | -0.065 | 25 | 3.951 | 0.217 | -3.951 | 21.7
[6 empty] | -0.065 | 25 | 3.941 | 0.217 | -3.941 | 21.7
[6 empty] | -0.065 | 25 | 3.931 | 0.217 | -3.931 | 21.7
[6 empty] | -0.065 | 25 | 3.921 | 0.217 | -3.921 | 21.7
[6 empty] | -0.065 | 25 | 3.911 | 0.217 | -3.911 | 21.7
[6 empty] | -0.065 | 25 | 3.901 | 0.217 | -3.901 | 21.7
[6 empty] | -0.065 | 25 | 3.891 | 0.217 | -3.891 | 21.7
[6 empty] | -0.065 | 25 | 3.881 | 0.217 | -3.881 | 21.7
[6 empty] | -0.065 | 25 | 3.871 | 0.216 | -3.871 | 21.6
[6 empty] | -0.065 | 25 | 3.861 | 0.216 | -3.861 | 21.6
[6 empty] | -0.065 | 25 | 3.851 | 0.216 | -3.851 | 21.6
[6 empty] | -0.065 | 25 | 3.841 | 0.216 | -3.841 | 21.6
[6 empty] | -0.065 | 25 | 3.831 | 0.216 | -3.831 | 21.6
[6 empty] | -0.065 | 25 | 3.821 | 0.216 | -3.821 | 21.6
[6 empty] | -0.065 | 25 | 3.811 | 0.216 | -3.811 | 21.6
[6 empty] | -0.065 | 25 | 3.801 | 0.216 | -3.801 | 21.6
[6 empty] | -0.065 | 25 | 3.791 | 0.215 | -3.791 | 21.5
[6 empty] | -0.065 | 25 | 3.781 | 0.215 | -3.781 | 21.5
[6 empty] | -0.065 | 25 | 3.771 | 0.215 | -3.771 | 21.5
[6 empty] | -0.065 | 25 | 3.761 | 0.215 | -3.761 | 21.5
[6 empty] | -0.065 | 25 | 3.751 | 0.215 | -3.751 | 21.5
[6 empty] | -0.065 | 25 | 3.741 | 0.215 | -3.741 | 21.5
[6 empty] | -0.065 | 25 | 3.731 | 0.215 | -3.731 | 21.5
[6 empty] | -0.065 | 25 | 3.721 | 0.215 | -3.721 | 21.5
[6 empty] | -0.065 | 25 | 3.711 | 0.215 | -3.711 | 21.5
[6 empty] | -0.065 | 25 | 3.701 | 0.214 | -3.701 | 21.4
[6 empty] | -0.065 | 25 | 3.691 | 0.214 | -3.691 | 21.4
[6 empty] | -0.065 | 25 | 3.681 | 0.214 | -3.681 | 21.4
[6 empty] | -0.065 | 25 | 3.671 | 0.214 | -3.671 | 21.4
[6 empty] | -0.065 | 25 | 3.661 | 0.214 | -3.661 | 21.4
[6 empty] | -0.065 | 25 | 3.651 | 0.214 | -3.651 | 21.4
[6 empty] | -0.065 | 25 | 3.641 | 0.214 | -3.641 | 21.4
[6 empty] | -0.065 | 25 | 3.631 | 0.214 | -3.631 | 21.4
[6 empty] | -0.065 | 25 | 3.621 | 0.213 | -3.621 | 21.3
[6 empty] | -0.065 | 25 | 3.611 | 0.213 | -3.611 | 21.3
[6 empty] | -0.065 | 25 | 3.601 | 0.213 | -3.601 | 21.3
[6 empty] | -0.065 | 25 | 3.591 | 0.213 | -3.591 | 21.3
[6 empty] | -0.065 | 25 | 3.581 | 0.213 | -3.581 | 21.3
[6 empty] | -0.065 | 25 | 3.571 | 0.213 | -3.571 | 21.3
[6 empty] | -0.065 | 25 | 3.561 | 0.213 | -3.561 | 21.3
[6 empty] | -0.065 | 25 | 3.551 | 0.213 | -3.551 | 21.3
[6 empty] | -0.065 | 25 | 3.541 | 0.213 | -3.541 | 21.3
[6 empty] | -0.065 | 25 | 3.531 | 0.212 | -3.531 | 21.2
[6 empty] | -0.065 | 25 | 3.521 | 0.212 | -3.521 | 21.2
[6 empty] | -0.065 | 25 | 3.511 | 0.212 | -3.511 | 21.2
[6 empty] | -0.065 | 25 | 3.501 | 0.212 | -3.501 | 21.2
[6 empty] | -0.065 | 25 | 3.491 | 0.212 | -3.491 | 21.2
[6 empty] | -0.065 | 25 | 3.481 | 0.212 | -3.481 | 21.2
[6 empty] | -0.065 | 25 | 3.471 | 0.212 | -3.471 | 21.2
[6 empty] | -0.066 | 25 | 3.461 | 0.212 | -3.461 | 21.2
[6 empty] | -0.066 | 25 | 3.451 | 0.211 | -3.451 | 21.1
[6 empty] | -0.066 | 25 | 3.441 | 0.211 | -3.441 | 21.1
[6 empty] | -0.066 | 25 | 3.431 | 0.211 | -3.431 | 21.1
[6 empty] | -0.066 | 25 | 3.421 | 0.211 | -3.421 | 21.1
[6 empty] | -0.066 | 25 | 3.411 | 0.211 | -3.411 | 21.1
[6 empty] | -0.066 | 25 | 3.401 | 0.211 | -3.401 | 21.1
[6 empty] | -0.066 | 25 | 3.391 | 0.211 | -3.391 | 21.1
[6 empty] | -0.066 | 25 | 3.381 | 0.211 | -3.381 | 21.1
[6 empty] | -0.066 | 25 | 3.371 | 0.211 | -3.371 | 21.1
[6 empty] | -0.066 | 25 | 3.361 | 0.21 | -3.361 | 21
[6 empty] | -0.066 | 25 | 3.351 | 0.21 | -3.351 | 21
[6 empty] | -0.066 | 25 | 3.341 | 0.21 | -3.341 | 21
[6 empty] | -0.066 | 25 | 3.331 | 0.21 | -3.331 | 21
[6 empty] | -0.066 | 25 | 3.321 | 0.21 | -3.321 | 21
[6 empty] | -0.066 | 25 | 3.311 | 0.21 | -3.311 | 21
[6 empty] | -0.066 | 25 | 3.301 | 0.21 | -3.301 | 21
[6 empty] | -0.066 | 25 | 3.291 | 0.21 | -3.291 | 21
[6 empty] | -0.066 | 25 | 3.281 | 0.209 | -3.281 | 20.9
[6 empty] | -0.066 | 25 | 3.271 | 0.209 | -3.271 | 20.9
[6 empty] | -0.066 | 25 | 3.262 | 0.209 | -3.262 | 20.9
[6 empty] | -0.066 | 25 | 3.252 | 0.209 | -3.252 | 20.9
[6 empty] | -0.066 | 25 | 3.242 | 0.209 | -3.242 | 20.9
[6 empty] | -0.066 | 25 | 3.232 | 0.209 | -3.232 | 20.9
[6 empty] | -0.066 | 25 | 3.222 | 0.209 | -3.222 | 20.9
[6 empty] | -0.066 | 25 | 3.212 | 0.209 | -3.212 | 20.9
[6 empty] | -0.066 | 25 | 3.202 | 0.209 | -3.202 | 20.9
[6 empty] | -0.066 | 25 | 3.192 | 0.208 | -3.192 | 20.8
[6 empty] | -0.066 | 25 | 3.182 | 0.208 | -3.182 | 20.8
[6 empty] | -0.066 | 25 | 3.172 | 0.208 | -3.172 | 20.8
[6 empty] | -0.066 | 25 | 3.162 | 0.208 | -3.162 | 20.8
[6 empty] | -0.066 | 25 | 3.152 | 0.208 | -3.152 | 20.8
[6 empty] | -0.066 | 25 | 3.142 | 0.208 | -3.142 | 20.8
[6 empty] | -0.066 | 25 | 3.132 | 0.208 | -3.132 | 20.8
[6 empty] | -0.066 | 25 | 3.122 | 0.208 | -3.122 | 20.8
[6 empty] | -0.066 | 25 | 3.112 | 0.208 | -3.112 | 20.8
[6 empty] | -0.066 | 25 | 3.102 | 0.207 | -3.102 | 20.7
[6 empty] | -0.066 | 25 | 3.092 | 0.207 | -3.092 | 20.7
[6 empty] | -0.066 | 25 | 3.082 | 0.207 | -3.082 | 20.7
[6 empty] | -0.066 | 25 | 3.072 | 0.207 | -3.072 | 20.7
[6 empty] | -0.066 | 25 | 3.062 | 0.207 | -3.062 | 20.7
[6 empty] | -0.066 | 25 | 3.052 | 0.207 | -3.052 | 20.7
[6 empty] | -0.066 | 25 | 3.042 | 0.207 | -3.042 | 20.7
[6 empty] | -0.066 | 25 | 3.032 | 0.207 | -3.032 | 20.7
[6 empty] | -0.066 | 25 | 3.022 | 0.207 | -3.022 | 20.7
[6 empty] | -0.066 | 25 | 3.012 | 0.206 | -3.012 | 20.6
[6 empty] | -0.066 | 25 | 3.002 | 0.206 | -3.002 | 20.6
[6 empty] | -0.066 | 25 | 2.992 | 0.206 | -2.992 | 20.6
[6 empty] | -0.066 | 25 | 2.982 | 0.206 | -2.982 | 20.6
[6 empty] | -0.066 | 25 | 2.972 | 0.206 | -2.972 | 20.6
[6 empty] | -0.066 | 25 | 2.962 | 0.206 | -2.962 | 20.6
[6 empty] | -0.066 | 25 | 2.952 | 0.206 | -2.952 | 20.6
[6 empty] | -0.066 | 25 | 2.942 | 0.206 | -2.942 | 20.6
[6 empty] | -0.066 | 25 | 2.932 | 0.205 | -2.932 | 20.5
[6 empty] | -0.066 | 25 | 2.922 | 0.205 | -2.922 | 20.5
[6 empty] | -0.066 | 25 | 2.912 | 0.205 | -2.912 | 20.5
[6 empty] | -0.066 | 25 | 2.902 | 0.205 | -2.902 | 20.5
[6 empty] | -0.066 | 25 | 2.892 | 0.205 | -2.892 | 20.5
[6 empty] | -0.066 | 25 | 2.882 | 0.205 | -2.882 | 20.5
[6 empty] | -0.066 | 25 | 2.872 | 0.205 | -2.872 | 20.5
[6 empty] | -0.066 | 25 | 2.862 | 0.205 | -2.862 | 20.5
[6 empty] | -0.066 | 25 | 2.852 | 0.205 | -2.852 | 20.5
[6 empty] | -0.066 | 25 | 2.842 | 0.204 | -2.842 | 20.4
[6 empty] | -0.066 | 25 | 2.832 | 0.204 | -2.832 | 20.4
[6 empty] | -0.066 | 25 | 2.822 | 0.204 | -2.822 | 20.4
[6 empty] | -0.066 | 25 | 2.812 | 0.204 | -2.812 | 20.4
[6 empty] | -0.066 | 25 | 2.802 | 0.204 | -2.802 | 20.4
[6 empty] | -0.066 | 25 | 2.792 | 0.204 | -2.792 | 20.4
[6 empty] | -0.066 | 25 | 2.782 | 0.204 | -2.782 | 20.4
[6 empty] | -0.066 | 25 | 2.772 | 0.204 | -2.772 | 20.4
[6 empty] | -0.067 | 25 | 2.762 | 0.204 | -2.762 | 20.4
[6 empty] | -0.067 | 25 | 2.752 | 0.203 | -2.752 | 20.3
[6 empty] | -0.067 | 25 | 2.742 | 0.203 | -2.742 | 20.3
[6 empty] | -0.067 | 25 | 2.732 | 0.203 | -2.732 | 20.3
[6 empty] | -0.067 | 25 | 2.722 | 0.203 | -2.722 | 20.3
[6 empty] | -0.067 | 25 | 2.712 | 0.203 | -2.712 | 20.3
[6 empty] | -0.067 | 25 | 2.702 | 0.203 | -2.702 | 20.3
[6 empty] | -0.067 | 25 | 2.692 | 0.203 | -2.692 | 20.3
[6 empty] | -0.067 | 25 | 2.682 | 0.203 | -2.682 | 20.3
[6 empty] | -0.067 | 25 | 2.672 | 0.203 | -2.672 | 20.3
[6 empty] | -0.067 | 25 | 2.662 | 0.202 | -2.662 | 20.2
[6 empty] | -0.067 | 25 | 2.652 | 0.202 | -2.652 | 20.2
[6 empty] | -0.067 | 25 | 2.642 | 0.202 | -2.642 | 20.2
[6 empty] | -0.067 | 25 | 2.632 | 0.202 | -2.632 | 20.2
[6 empty] | -0.067 | 25 | 2.622 | 0.202 | -2.622 | 20.2
[6 empty] | -0.067 | 25 | 2.612 | 0.202 | -2.612 | 20.2
[6 empty] | -0.067 | 25 | 2.602 | 0.202 | -2.602 | 20.2
[6 empty] | -0.067 | 25 | 2.592 | 0.202 | -2.592 | 20.2
[6 empty] | -0.067 | 25 | 2.582 | 0.202 | -2.582 | 20.2
[6 empty] | -0.067 | 25 | 2.572 | 0.202 | -2.572 | 20.2
[6 empty] | -0.067 | 25 | 2.562 | 0.201 | -2.562 | 20.1
[6 empty] | -0.067 | 25 | 2.552 | 0.201 | -2.552 | 20.1
[6 empty] | -0.067 | 25 | 2.542 | 0.201 | -2.542 | 20.1
[6 empty] | -0.067 | 25 | 2.532 | 0.201 | -2.532 | 20.1
[6 empty] | -0.067 | 25 | 2.522 | 0.201 | -2.522 | 20.1
[6 empty] | -0.067 | 25 | 2.512 | 0.201 | -2.512 | 20.1
[6 empty] | -0.067 | 25 | 2.502 | 0.201 | -2.502 | 20.1
[6 empty] | -0.067 | 25 | 2.492 | 0.201 | -2.492 | 20.1
[6 empty] | -0.067 | 25 | 2.482 | 0.201 | -2.482 | 20.1
[6 empty] | -0.067 | 25 | 2.472 | 0.2 | -2.472 | 20
[6 empty] | -0.067 | 25 | 2.462 | 0.2 | -2.462 | 20
[6 empty] | -0.067 | 25 | 2.452 | 0.2 | -2.452 | 20
[6 empty] | -0.067 | 25 | 2.442 | 0.2 | -2.442 | 20
[6 empty] | -0.067 | 25 | 2.432 | 0.2 | -2.432 | 20
[6 empty] | -0.067 | 25 | 2.422 | 0.2 | -2.422 | 20
[6 empty] | -0.067 | 25 | 2.412 | 0.2 | -2.412 | 20
[6 empty] | -0.067 | 25 | 2.402 | 0.2 | -2.402 | 20
[6 empty] | -0.067 | 25 | 2.392 | 0.2 | -2.392 | 20
[6 empty] | -0.067 | 25 | 2.382 | 0.199 | -2.382 | 19.9
[6 empty] | -0.067 | 25 | 2.372 | 0.199 | -2.372 | 19.9
[6 empty] | -0.067 | 25 | 2.362 | 0.199 | -2.362 | 19.9
[6 empty] | -0.067 | 25 | 2.352 | 0.199 | -2.352 | 19.9
[6 empty] | -0.067 | 25 | 2.342 | 0.199 | -2.342 | 19.9
[6 empty] | -0.067 | 25 | 2.332 | 0.199 | -2.332 | 19.9
[6 empty] | -0.067 | 25 | 2.322 | 0.199 | -2.322 | 19.9
[6 empty] | -0.067 | 25 | 2.312 | 0.199 | -2.312 | 19.9
[6 empty] | -0.067 | 25 | 2.302 | 0.199 | -2.302 | 19.9
[6 empty] | -0.067 | 25 | 2.292 | 0.198 | -2.292 | 19.8
[6 empty] | -0.067 | 25 | 2.282 | 0.198 | -2.282 | 19.8
[6 empty] | -0.067 | 25 | 2.272 | 0.198 | -2.272 | 19.8
[6 empty] | -0.067 | 25 | 2.262 | 0.198 | -2.262 | 19.8
[6 empty] | -0.067 | 25 | 2.252 | 0.198 | -2.252 | 19.8
[6 empty] | -0.067 | 25 | 2.242 | 0.198 | -2.242 | 19.8
[6 empty] | -0.067 | 25 | 2.232 | 0.198 | -2.232 | 19.8
[6 empty] | -0.067 | 25 | 2.222 | 0.198 | -2.222 | 19.8
[6 empty] | -0.067 | 25 | 2.212 | 0.198 | -2.212 | 19.8
[6 empty] | -0.067 | 25 | 2.202 | 0.197 | -2.202 | 19.7
[6 empty] | -0.067 | 25 | 2.192 | 0.197 | -2.192 | 19.7
[6 empty] | -0.067 | 25 | 2.182 | 0.197 | -2.182 | 19.7
[6 empty] | -0.067 | 25 | 2.172 | 0.197 | -2.172 | 19.7
[6 empty] | -0.067 | 25 | 2.162 | 0.197 | -2.162 | 19.7
[6 empty] | -0.067 | 25 | 2.152 | 0.197 | -2.152 | 19.7
[6 empty] | -0.067 | 25 | 2.142 | 0.197 | -2.142 | 19.7
[6 empty] | -0.067 | 25 | 2.132 | 0.197 | -2.132 | 19.7
[6 empty] | -0.067 | 25 | 2.122 | 0.197 | -2.122 | 19.7
[6 empty] | -0.067 | 25 | 2.112 | 0.197 | -2.112 | 19.7
[6 empty] | -0.067 | 25 | 2.102 | 0.196 | -2.102 | 19.6
[6 empty] | -0.067 | 25 | 2.092 | 0.196 | -2.092 | 19.6
[6 empty] | -0.067 | 25 | 2.082 | 0.196 | -2.082 | 19.6
[6 empty] | -0.067 | 25 | 2.072 | 0.196 | -2.072 | 19.6
[6 empty] | -0.068 | 25 | 2.062 | 0.196 | -2.062 | 19.6
[6 empty] | -0.068 | 25 | 2.052 | 0.196 | -2.052 | 19.6
[6 empty] | -0.068 | 25 | 2.042 | 0.196 | -2.042 | 19.6
[6 empty] | -0.068 | 25 | 2.032 | 0.196 | -2.032 | 19.6
[6 empty] | -0.068 | 25 | 2.022 | 0.196 | -2.022 | 19.6
[6 empty] | -0.068 | 25 | 2.012 | 0.195 | -2.012 | 19.5
[6 empty] | -0.068 | 25 | 2.002 | 0.195 | -2.002 | 19.5
[6 empty] | -0.068 | 25 | 1.992 | 0.195 | -1.992 | 19.5
[6 empty] | -0.068 | 25 | 1.982 | 0.195 | -1.982 | 19.5
[6 empty] | -0.068 | 25 | 1.972 | 0.195 | -1.972 | 19.5
[6 empty] | -0.068 | 25 | 1.962 | 0.195 | -1.962 | 19.5
[6 empty] | -0.068 | 25 | 1.952 | 0.195 | -1.952 | 19.5
[6 empty] | -0.068 | 25 | 1.942 | 0.195 | -1.942 | 19.5
[6 empty] | -0.068 | 25 | 1.932 | 0.195 | -1.932 | 19.5
[6 empty] | -0.068 | 25 | 1.922 | 0.195 | -1.922 | 19.5
[6 empty] | -0.068 | 25 | 1.912 | 0.194 | -1.912 | 19.4
[6 empty] | -0.068 | 25 | 1.902 | 0.194 | -1.902 | 19.4
[6 empty] | -0.068 | 25 | 1.892 | 0.194 | -1.892 | 19.4
[6 empty] | -0.068 | 25 | 1.882 | 0.194 | -1.882 | 19.4
[6 empty] | -0.068 | 25 | 1.872 | 0.194 | -1.872 | 19.4
[6 empty] | -0.068 | 25 | 1.862 | 0.194 | -1.862 | 19.4
[6 empty] | -0.068 | 25 | 1.852 | 0.194 | -1.852 | 19.4
[6 empty] | -0.068 | 25 | 1.842 | 0.194 | -1.842 | 19.4
[6 empty] | -0.068 | 25 | 1.832 | 0.194 | -1.832 | 19.4
[6 empty] | -0.068 | 25 | 1.822 | 0.193 | -1.822 | 19.3
[6 empty] | -0.068 | 25 | 1.812 | 0.193 | -1.812 | 19.3
[6 empty] | -0.068 | 25 | 1.802 | 0.193 | -1.802 | 19.3
[6 empty] | -0.068 | 25 | 1.792 | 0.193 | -1.792 | 19.3
[6 empty] | -0.068 | 25 | 1.782 | 0.193 | -1.782 | 19.3
[6 empty] | -0.068 | 25 | 1.772 | 0.193 | -1.772 | 19.3
[6 empty] | -0.068 | 25 | 1.762 | 0.193 | -1.762 | 19.3
[6 empty] | -0.068 | 25 | 1.752 | 0.193 | -1.752 | 19.3
[6 empty] | -0.068 | 25 | 1.742 | 0.193 | -1.742 | 19.3
[6 empty] | -0.068 | 25 | 1.732 | 0.193 | -1.732 | 19.3
[6 empty] | -0.068 | 25 | 1.722 | 0.192 | -1.722 | 19.2
[6 empty] | -0.068 | 25 | 1.712 | 0.192 | -1.712 | 19.2
[6 empty] | -0.068 | 25 | 1.702 | 0.192 | -1.702 | 19.2
[6 empty] | -0.068 | 25 | 1.692 | 0.192 | -1.692 | 19.2
[6 empty] | -0.068 | 25 | 1.682 | 0.192 | -1.682 | 19.2
[6 empty] | -0.068 | 25 | 1.672 | 0.192 | -1.672 | 19.2
[6 empty] | -0.068 | 25 | 1.662 | 0.192 | -1.662 | 19.2
[6 empty] | -0.068 | 25 | 1.652 | 0.192 | -1.652 | 19.2
[6 empty] | -0.068 | 25 | 1.642 | 0.192 | -1.642 | 19.2
[6 empty] | -0.068 | 25 | 1.632 | 0.191 | -1.632 | 19.1
[6 empty] | -0.068 | 25 | 1.622 | 0.191 | -1.622 | 19.1
[6 empty] | -0.068 | 25 | 1.612 | 0.191 | -1.612 | 19.1
[6 empty] | -0.068 | 25 | 1.602 | 0.191 | -1.602 | 19.1
[6 empty] | -0.068 | 25 | 1.592 | 0.191 | -1.592 | 19.1
[6 empty] | -0.068 | 25 | 1.582 | 0.191 | -1.582 | 19.1
[6 empty] | -0.068 | 25 | 1.572 | 0.191 | -1.572 | 19.1
[6 empty] | -0.068 | 25 | 1.562 | 0.191 | -1.562 | 19.1
[6 empty] | -0.068 | 25 | 1.552 | 0.191 | -1.552 | 19.1
[6 empty] | -0.068 | 25 | 1.542 | 0.191 | -1.542 | 19.1
[6 empty] | -0.068 | 25 | 1.532 | 0.19 | -1.532 | 19
[6 empty] | -0.068 | 25 | 1.522 | 0.19 | -1.522 | 19
[6 empty] | -0.068 | 25 | 1.512 | 0.19 | -1.512 | 19
[6 empty] | -0.068 | 25 | 1.502 | 0.19 | -1.502 | 19
[6 empty] | -0.068 | 25 | 1.492 | 0.19 | -1.492 | 19
[6 empty] | -0.068 | 25 | 1.482 | 0.19 | -1.482 | 19
[6 empty] | -0.068 | 25 | 1.473 | 0.19 | -1.473 | 19
[6 empty] | -0.068 | 25 | 1.463 | 0.19 | -1.463 | 19
[6 empty] | -0.068 | 25 | 1.453 | 0.19 | -1.453 | 19
[6 empty] | -0.068 | 25 | 1.443 | 0.189 | -1.443 | 18.9
[6 empty] | -0.068 | 25 | 1.433 | 0.189 | -1.433 | 18.9
[6 empty] | -0.068 | 25 | 1.423 | 0.189 | -1.423 | 18.9
[6 empty] | -0.068 | 25 | 1.413 | 0.189 | -1.413 | 18.9
[6 empty] | -0.068 | 25 | 1.403 | 0.189 | -1.403 | 18.9
[6 empty] | -0.068 | 25 | 1.393 | 0.189 | -1.393 | 18.9
[6 empty] | -0.068 | 25 | 1.383 | 0.189 | -1.383 | 18.9
[6 empty] | -0.068 | 25 | 1.373 | 0.189 | -1.373 | 18.9
[6 empty] | -0.069 | 25 | 1.363 | 0.189 | -1.363 | 18.9
[6 empty] | -0.069 | 25 | 1.353 | 0.189 | -1.353 | 18.9
[6 empty] | -0.069 | 25 | 1.343 | 0.188 | -1.343 | 18.8
[6 empty] | -0.069 | 25 | 1.333 | 0.188 | -1.333 | 18.8
[6 empty] | -0.069 | 25 | 1.323 | 0.188 | -1.323 | 18.8
[6 empty] | -0.069 | 25 | 1.313 | 0.188 | -1.313 | 18.8
[6 empty] | -0.069 | 25 | 1.303 | 0.188 | -1.303 | 18.8
[6 empty] | -0.069 | 25 | 1.293 | 0.188 | -1.293 | 18.8
[6 empty] | -0.069 | 25 | 1.283 | 0.188 | -1.283 | 18.8
[6 empty] | -0.069 | 25 | 1.273 | 0.188 | -1.273 | 18.8
[6 empty] | -0.069 | 25 | 1.263 | 0.188 | -1.263 | 18.8
[6 empty] | -0.069 | 25 | 1.253 | 0.188 | -1.253 | 18.8
[6 empty] | -0.069 | 25 | 1.243 | 0.187 | -1.243 | 18.7
[6 empty] | -0.069 | 25 | 1.233 | 0.187 | -1.233 | 18.7
[6 empty] | -0.069 | 25 | 1.223 | 0.187 | -1.223 | 18.7
[6 empty] | -0.069 | 25 | 1.213 | 0.187 | -1.213 | 18.7
[6 empty] | -0.069 | 25 | 1.203 | 0.187 | -1.203 | 18.7
[6 empty] | -0.069 | 25 | 1.193 | 0.187 | -1.193 | 18.7
[6 empty] | -0.069 | 25 | 1.183 | 0.187 | -1.183 | 18.7
[6 empty] | -0.069 | 25 | 1.173 | 0.187 | -1.173 | 18.7
[6 empty] | -0.069 | 25 | 1.163 | 0.187 | -1.163 | 18.7
[6 empty] | -0.069 | 25 | 1.153 | 0.187 | -1.153 | 18.7
[6 empty] | -0.069 | 25 | 1.143 | 0.186 | -1.143 | 18.6
[6 empty] | -0.069 | 25 | 1.133 | 0.186 | -1.133 | 18.6
[6 empty] | -0.069 | 25 | 1.123 | 0.186 | -1.123 | 18.6
[6 empty] | -0.069 | 25 | 1.113 | 0.186 | -1.113 | 18.6
[6 empty] | -0.069 | 25 | 1.103 | 0.186 | -1.103 | 18.6
[6 empty] | -0.069 | 25 | 1.093 | 0.186 | -1.093 | 18.6
[6 empty] | -0.069 | 25 | 1.083 | 0.186 | -1.083 | 18.6
[6 empty] | -0.069 | 25 | 1.073 | 0.186 | -1.073 | 18.6
[6 empty] | -0.069 | 25 | 1.063 | 0.186 | -1.063 | 18.6
[6 empty] | -0.069 | 25 | 1.053 | 0.185 | -1.053 | 18.5
[6 empty] | -0.069 | 25 | 1.043 | 0.185 | -1.043 | 18.5
[6 empty] | -0.069 | 25 | 1.033 | 0.185 | -1.033 | 18.5
[6 empty] | -0.069 | 25 | 1.023 | 0.185 | -1.023 | 18.5
[6 empty] | -0.069 | 25 | 1.013 | 0.185 | -1.013 | 18.5
[6 empty] | -0.069 | 25 | 1.003 | 0.185 | -1.003 | 18.5
[6 empty] | -0.069 | 25 | 0.993 | 0.185 | -0.993 | 18.5
[6 empty] | -0.069 | 25 | 0.983 | 0.185 | -0.983 | 18.5
[6 empty] | -0.069 | 25 | 0.973 | 0.185 | -0.973 | 18.5
[6 empty] | -0.069 | 25 | 0.963 | 0.185 | -0.963 | 18.5
[6 empty] | -0.069 | 25 | 0.953 | 0.184 | -0.953 | 18.4
[6 empty] | -0.069 | 25 | 0.943 | 0.184 | -0.943 | 18.4
[6 empty] | -0.069 | 25 | 0.933 | 0.184 | -0.933 | 18.4
[6 empty] | -0.069 | 25 | 0.923 | 0.184 | -0.923 | 18.4
[6 empty] | -0.069 | 25 | 0.913 | 0.184 | -0.913 | 18.4
[6 empty] | -0.069 | 25 | 0.903 | 0.184 | -0.903 | 18.4
[6 empty] | -0.069 | 25 | 0.893 | 0.184 | -0.893 | 18.4
[6 empty] | -0.069 | 25 | 0.883 | 0.184 | -0.883 | 18.4
[6 empty] | -0.069 | 25 | 0.873 | 0.184 | -0.873 | 18.4
[6 empty] | -0.069 | 25 | 0.863 | 0.184 | -0.863 | 18.4
[6 empty] | -0.069 | 25 | 0.853 | 0.183 | -0.853 | 18.3
[6 empty] | -0.069 | 25 | 0.843 | 0.183 | -0.843 | 18.3
[6 empty] | -0.069 | 25 | 0.833 | 0.183 | -0.833 | 18.3
[6 empty] | -0.069 | 25 | 0.823 | 0.183 | -0.823 | 18.3
[6 empty] | -0.069 | 25 | 0.813 | 0.183 | -0.813 | 18.3
[6 empty] | -0.069 | 25 | 0.803 | 0.183 | -0.803 | 18.3
[6 empty] | -0.069 | 25 | 0.793 | 0.183 | -0.793 | 18.3
[6 empty] | -0.069 | 25 | 0.783 | 0.183 | -0.783 | 18.3
[6 empty] | -0.069 | 25 | 0.773 | 0.183 | -0.773 | 18.3
[6 empty] | -0.069 | 25 | 0.763 | 0.183 | -0.763 | 18.3
[6 empty] | -0.069 | 25 | 0.753 | 0.182 | -0.753 | 18.2
[6 empty] | -0.069 | 25 | 0.743 | 0.182 | -0.743 | 18.2
[6 empty] | -0.069 | 25 | 0.733 | 0.182 | -0.733 | 18.2
[6 empty] | -0.069 | 25 | 0.723 | 0.182 | -0.723 | 18.2
[6 empty] | -0.069 | 25 | 0.713 | 0.182 | -0.713 | 18.2
[6 empty] | -0.069 | 25 | 0.703 | 0.182 | -0.703 | 18.2
[6 empty] | -0.069 | 25 | 0.693 | 0.182 | -0.693 | 18.2
[6 empty] | -0.069 | 25 | 0.683 | 0.182 | -0.683 | 18.2
[6 empty] | -0.069 | 25 | 0.673 | 0.182 | -0.673 | 18.2
[6 empty] | -0.07 | 25 | 0.663 | 0.182 | -0.663 | 18.2
[6 empty] | -0.07 | 25 | 0.653 | 0.181 | -0.653 | 18.1
[6 empty] | -0.07 | 25 | 0.643 | 0.181 | -0.643 | 18.1
[6 empty] | -0.07 | 25 | 0.633 | 0.181 | -0.633 | 18.1
[6 empty] | -0.07 | 25 | 0.623 | 0.181 | -0.623 | 18.1
[6 empty] | -0.07 | 25 | 0.613 | 0.181 | -0.613 | 18.1
[6 empty] | -0.07 | 25 | 0.603 | 0.181 | -0.603 | 18.1
[6 empty] | -0.07 | 25 | 0.593 | 0.181 | -0.593 | 18.1
[6 empty] | -0.07 | 25 | 0.583 | 0.181 | -0.583 | 18.1
[6 empty] | -0.07 | 25 | 0.573 | 0.181 | -0.573 | 18.1
[6 empty] | -0.07 | 25 | 0.563 | 0.181 | -0.563 | 18.1
[6 empty] | -0.07 | 25 | 0.553 | 0.18 | -0.553 | 18
[6 empty] | -0.07 | 25 | 0.543 | 0.18 | -0.543 | 18
[6 empty] | -0.07 | 25 | 0.533 | 0.18 | -0.533 | 18
[6 empty] | -0.07 | 25 | 0.523 | 0.18 | -0.523 | 18
[6 empty] | -0.07 | 25 | 0.513 | 0.18 | -0.513 | 18
[6 empty] | -0.07 | 25 | 0.503 | 0.18 | -0.503 | 18
[6 empty] | -0.07 | 25 | 0.493 | 0.18 | -0.493 | 18
[6 empty] | -0.07 | 25 | 0.483 | 0.18 | -0.483 | 18
[6 empty] | -0.07 | 25 | 0.473 | 0.18 | -0.473 | 18
[6 empty] | -0.07 | 25 | 0.463 | 0.18 | -0.463 | 18
[6 empty] | -0.07 | 25 | 0.453 | 0.179 | -0.453 | 17.9
[6 empty] | -0.07 | 25 | 0.443 | 0.179 | -0.443 | 17.9
[6 empty] | -0.07 | 25 | 0.433 | 0.179 | -0.433 | 17.9
[6 empty] | -0.07 | 25 | 0.423 | 0.179 | -0.423 | 17.9
[6 empty] | -0.07 | 25 | 0.413 | 0.179 | -0.413 | 17.9
[6 empty] | -0.07 | 25 | 0.403 | 0.179 | -0.403 | 17.9
[6 empty] | -0.07 | 25 | 0.393 | 0.179 | -0.393 | 17.9
[6 empty] | -0.07 | 25 | 0.383 | 0.179 | -0.383 | 17.9
[6 empty] | -0.07 | 25 | 0.373 | 0.179 | -0.373 | 17.9
[6 empty] | -0.07 | 25 | 0.363 | 0.179 | -0.363 | 17.9
[6 empty] | -0.07 | 25 | 0.353 | 0.179 | -0.353 | 17.9
[6 empty] | -0.07 | 25 | 0.343 | 0.178 | -0.343 | 17.8
[6 empty] | -0.07 | 25 | 0.333 | 0.178 | -0.333 | 17.8
[6 empty] | -0.07 | 25 | 0.323 | 0.178 | -0.323 | 17.8
[6 empty] | -0.07 | 25 | 0.313 | 0.178 | -0.313 | 17.8
[6 empty] | -0.07 | 25 | 0.303 | 0.178 | -0.303 | 17.8
[6 empty] | -0.07 | 25 | 0.293 | 0.178 | -0.293 | 17.8
[6 empty] | -0.07 | 25 | 0.283 | 0.178 | -0.283 | 17.8
[6 empty] | -0.07 | 25 | 0.273 | 0.178 | -0.273 | 17.8
[6 empty] | -0.07 | 25 | 0.263 | 0.178 | -0.263 | 17.8
[6 empty] | -0.07 | 25 | 0.253 | 0.178 | -0.253 | 17.8
[6 empty] | -0.07 | 25 | 0.243 | 0.177 | -0.243 | 17.7
[6 empty] | -0.07 | 25 | 0.233 | 0.177 | -0.233 | 17.7
[6 empty] | -0.07 | 25 | 0.223 | 0.177 | -0.223 | 17.7
[6 empty] | -0.07 | 25 | 0.213 | 0.177 | -0.213 | 17.7
[6 empty] | -0.07 | 25 | 0.203 | 0.177 | -0.203 | 17.7
[6 empty] | -0.07 | 25 | 0.193 | 0.177 | -0.193 | 17.7
[6 empty] | -0.07 | 25 | 0.183 | 0.177 | -0.183 | 17.7
[6 empty] | -0.07 | 25 | 0.173 | 0.177 | -0.173 | 17.7
[6 empty] | -0.07 | 25 | 0.163 | 0.177 | -0.163 | 17.7
[6 empty] | -0.07 | 25 | 0.153 | 0.177 | -0.153 | 17.7
[6 empty] | -0.07 | 25 | 0.143 | 0.177 | -0.143 | 17.7
[6 empty] | -0.07 | 25 | 0.133 | 0.176 | -0.133 | 17.6
[6 empty] | -0.07 | 25 | 0.123 | 0.176 | -0.123 | 17.6
[6 empty] | ####################################################### |  |  |  | 0 | 0

## CROSSPLANE
SCRIPT | [5 empty] | RAYSTATION
BEGIN_SCAN_DATA | [5 empty] | #######################################################
| FORMAT=CC-Export V1.9 |  |  |  |  | #RayStationVersion: 13.1.0.144
| FILE_CREATION_DATE= |  |  |  |  | #PatientName: RayTraining 40x40x40 Phantom
| LAST_MODIFIED=11-Jan-2023 11:56:47 |  |  |  |  | #PatientId: Watertank cube
| BEGIN_SCAN 1 |  |  |  |  | #CoordinateSystem: IEC patient coordinate system
|  | TASK_NAME=RayStation Phantom Profiles & PDDs |  |  |  | #######################################################
|  | PROGRAM=RayStation |  |  |  | #LineName: 1
|  | COMMENT=RayStation Script Export |  |  |  | #DoseName: Beam dose: 15x10 (Validation, CT 1)
|  | MEAS_DATE=11-Jan-2023 11:56:47 |  |  |  | #DoseEngine: Clinical: Collapsed Cone v5.7
|  | LINAC=RayStation |  |  |  | #TreatmentMachine: RSL_TrueBeam Commission time: 11/20/2018 14:29:57
|  | MODALITY=X |  |  |  | #X [cm] | Y [cm] | Z [cm] | Dose [Gy] | cGy
|  | ISOCENTER=1000.00 |  |  |  | -30.232 | 25 | 20.135 | 0 | 0
|  | INPLANE_AXIS=Y |  |  |  | -30.222 | 25 | 20.135 | 0 | 0
|  | CROSSPLANE_AXIS=X |  |  |  | -30.212 | 25 | 20.135 | 0 | 0
|  | DEPTH_AXIS=Z |  |  |  | -30.202 | 25 | 20.135 | 0 | 0
|  | INPLANE_AXIS_DIR=GUN_TARGET |  |  |  | -30.192 | 25 | 20.135 | 0 | 0
|  | CROSSPLANE_AXIS_DIR=LEFT_RIGHT |  |  |  | -30.182 | 25 | 20.135 | 0 | 0
|  | DEPTH_AXIS_DIR=UP_DOWN |  |  |  | -30.172 | 25 | 20.135 | 0 | 0
|  | ENERGY=6.3 |  |  |  | -30.162 | 25 | 20.136 | 0 | 0
|  | NOMINAL_DMAX=20.00 |  |  |  | -30.152 | 25 | 20.136 | 0 | 0
|  | SSD=1000.0 |  |  |  | -30.142 | 25 | 20.136 | 0 | 0
|  | SCD=1000.00 |  |  |  | -30.132 | 25 | 20.136 | 0 | 0
|  | WEDGE_ANGLE=0.00 |  |  |  | -30.122 | 25 | 20.136 | 0 | 0
|  | FIELD_INPLANE=150.0 |  |  |  | -30.112 | 25 | 20.136 | 0 | 0
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -30.102 | 25 | 20.136 | 0 | 0
|  | FIELD_TYPE=IRREGULAR |  |  |  | -30.092 | 25 | 20.136 | 0 | 0
|  | GANTRY=0.00 |  |  |  | -30.082 | 25 | 20.136 | 0 | 0
|  | GANTRY_UPRIGHT_POSITION=0 |  |  |  | -30.072 | 25 | 20.136 | 0 | 0
|  | GANTRY_ROTATION=CW |  |  |  | -30.062 | 25 | 20.136 | 0 | 0
|  | COLL_ANGLE=0.00 |  |  |  | -30.052 | 25 | 20.136 | 0 | 0
|  | REF_FIELD_DEPTH=50.00 |  |  |  | -30.042 | 25 | 20.136 | 0 | 0
|  | REF_FIELD_DEFINED=ISOCENTER |  |  |  | -30.032 | 25 | 20.136 | 0 | 0
|  | FIELD_INPLANE=150.0 |  |  |  | -30.022 | 25 | 20.136 | 0 | 0
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -30.012 | 25 | 20.136 | 0 | 0
|  | SCAN_CURVETYPE=CROSSPLANE_PROFILE |  |  |  | -30.002 | 25 | 20.136 | 0 | 0
|  | SCAN_DEPTH=100.0 |  |  |  | -29.992 | 25 | 20.136 | 0 | 0
|  | SCAN_OFFAXIS_INPLANE=0.00 |  |  |  | -29.982 | 25 | 20.136 | 0 | 0
|  | SCAN_OFFAXIS_CROSSPLANE=0.00 |  |  |  | -29.972 | 25 | 20.136 | 0 | 0
|  | SCAN_ANGLE=0.00 |  |  |  | -29.962 | 25 | 20.136 | 0 | 0
|  | SCAN_DIAGONAL=NOT_DIAGONAL |  |  |  | -29.952 | 25 | 20.136 | 0 | 0
|  | SCAN_DIRECTION=NEGATIVE |  |  |  | -29.942 | 25 | 20.136 | 0 | 0
|  | MEAS_MEDIUM=WATER |  |  |  | -29.932 | 25 | 20.136 | 0 | 0
|  | MEAS_UNIT=cGy |  |  |  | -29.922 | 25 | 20.136 | 0 | 0
|  | INCLINATION_ANGLE=0.00 |  |  |  | -29.912 | 25 | 20.136 | 0 | 0
|  | PRESSURE=1013.20 |  |  |  | -29.902 | 25 | 20.136 | 0 | 0
|  | TEMPERATURE=20.00 |  |  |  | -29.892 | 25 | 20.136 | 0 | 0
|  | NORM_TEMPERATURE=20.00 |  |  |  | -29.882 | 25 | 20.136 | 0 | 0
|  | CORRECTION_FACTOR=1.0000; |  |  |  | -29.872 | 25 | 20.136 | 0 | 0
|  | BEGIN_DATA |  |  |  | -29.862 | 25 | 20.136 | 0 | 0
|  |  | -30.05 | 5.7848e-05 | 1 | -29.852 | 25 | 20.136 | 0 | 0
|  |  | -30 | 8.6771e-05 | 1 | -29.842 | 25 | 20.136 | 0 | 0
|  |  | -29.95 | 0.00011569 | 1 | -29.832 | 25 | 20.136 | 0 | 0
|  |  | -29.9 | 0.00011723 | 1 | -29.822 | 25 | 20.136 | 0 | 0
|  |  | -29.85 | 0.00011876 | 1 | -29.812 | 25 | 20.136 | 0 | 0
|  |  | -29.8 | 0.00012029 | 1 | -29.802 | 25 | 20.136 | 0 | 0
|  |  | -29.75 | 0.00012183 | 1 | -29.792 | 25 | 20.136 | 0 | 0
|  |  | -29.7 | 0.00012337 | 1 | -29.782 | 25 | 20.136 | 0 | 0
|  |  | -29.65 | 0.00012492 | 1 | -29.772 | 25 | 20.136 | 0 | 0
|  |  | -29.6 | 0.00012646 | 1 | -29.762 | 25 | 20.136 | 0 | 0
|  |  | -29.55 | 0.00012801 | 1 | -29.753 | 25 | 20.136 | 0 | 0
|  |  | -29.5 | 0.00013007 | 1 | -29.743 | 25 | 20.136 | 0 | 0
|  |  | -29.45 | 0.00013214 | 1 | -29.733 | 25 | 20.136 | 0 | 0
|  |  | -29.4 | 0.0001342 | 1 | -29.723 | 25 | 20.136 | 0 | 0
|  |  | -29.35 | 0.00013626 | 1 | -29.713 | 25 | 20.136 | 0 | 0
|  |  | -29.3 | 0.00013779 | 1 | -29.703 | 25 | 20.136 | 0 | 0
|  |  | -29.25 | 0.00013931 | 1 | -29.693 | 25 | 20.136 | 0 | 0
|  |  | -29.2 | 0.00014083 | 1 | -29.683 | 25 | 20.136 | 0 | 0
|  |  | -29.15 | 0.00014236 | 1 | -29.673 | 25 | 20.136 | 0 | 0
|  |  | -29.1 | 0.0001446 | 1 | -29.663 | 25 | 20.136 | 0 | 0
|  |  | -29.05 | 0.00014683 | 1 | -29.653 | 25 | 20.136 | 0 | 0
|  |  | -29 | 0.00014907 | 1 | -29.643 | 25 | 20.136 | 0 | 0
|  |  | -28.95 | 0.00015131 | 1 | -29.633 | 25 | 20.136 | 0 | 0
|  |  | -28.9 | 0.00015368 | 1 | -29.623 | 25 | 20.136 | 0 | 0
|  |  | -28.85 | 0.00015605 | 1 | -29.613 | 25 | 20.136 | 0 | 0
|  |  | -28.8 | 0.00015842 | 1 | -29.603 | 25 | 20.136 | 0 | 0
|  |  | -28.75 | 0.00016079 | 1 | -29.593 | 25 | 20.136 | 0 | 0
|  |  | -28.7 | 0.00017078 | 1 | -29.583 | 25 | 20.136 | 0 | 0
|  |  | -28.65 | 0.00018078 | 1 | -29.573 | 25 | 20.136 | 0 | 0
|  |  | -28.6 | 0.00019077 | 1 | -29.563 | 25 | 20.136 | 0 | 0
|  |  | -28.55 | 0.00020076 | 1 | -29.553 | 25 | 20.136 | 0 | 0
|  |  | -28.5 | 0.00021955 | 1 | -29.543 | 25 | 20.136 | 0 | 0
|  |  | -28.45 | 0.00023835 | 1 | -29.533 | 25 | 20.137 | 0 | 0
|  |  | -28.4 | 0.00025714 | 1 | -29.523 | 25 | 20.137 | 0 | 0
|  |  | -28.35 | 0.00027593 | 1 | -29.513 | 25 | 20.137 | 0 | 0
|  |  | -28.3 | 0.0002941 | 1 | -29.503 | 25 | 20.137 | 0 | 0
|  |  | -28.25 | 0.00031226 | 1 | -29.493 | 25 | 20.137 | 0 | 0
|  |  | -28.2 | 0.00033042 | 1 | -29.483 | 25 | 20.137 | 0 | 0
|  |  | -28.15 | 0.00034859 | 1 | -29.473 | 25 | 20.137 | 0 | 0
|  |  | -28.1 | 0.00036788 | 1 | -29.463 | 25 | 20.137 | 0 | 0
|  |  | -28.05 | 0.00038717 | 1 | -29.453 | 25 | 20.137 | 0 | 0
|  |  | -28 | 0.00040647 | 1 | -29.443 | 25 | 20.137 | 0 | 0
|  |  | -27.95 | 0.00042576 | 1 | -29.433 | 25 | 20.137 | 0 | 0
|  |  | -27.9 | 0.00044501 | 1 | -29.423 | 25 | 20.137 | 0 | 0
|  |  | -27.85 | 0.00046425 | 1 | -29.413 | 25 | 20.137 | 0 | 0
|  |  | -27.8 | 0.0004835 | 1 | -29.403 | 25 | 20.137 | 0 | 0
|  |  | -27.75 | 0.00050275 | 1 | -29.393 | 25 | 20.137 | 0 | 0
|  |  | -27.7 | 0.00052282 | 1 | -29.383 | 25 | 20.137 | 0 | 0
|  |  | -27.65 | 0.00054289 | 1 | -29.373 | 25 | 20.137 | 0 | 0
|  |  | -27.6 | 0.00056296 | 1 | -29.363 | 25 | 20.137 | 0 | 0
|  |  | -27.55 | 0.00058304 | 1 | -29.353 | 25 | 20.137 | 0 | 0
|  |  | -27.5 | 0.00061092 | 1 | -29.343 | 25 | 20.137 | 0 | 0
|  |  | -27.45 | 0.00063881 | 1 | -29.333 | 25 | 20.137 | 0 | 0
|  |  | -27.4 | 0.0006667 | 1 | -29.323 | 25 | 20.137 | 0 | 0
|  |  | -27.35 | 0.00069459 | 1 | -29.313 | 25 | 20.137 | 0 | 0
|  |  | -27.3 | 0.00072209 | 1 | -29.303 | 25 | 20.137 | 0 | 0
|  |  | -27.25 | 0.00074959 | 1 | -29.293 | 25 | 20.137 | 0 | 0
|  |  | -27.2 | 0.00077709 | 1 | -29.283 | 25 | 20.137 | 0 | 0
|  |  | -27.15 | 0.00080459 | 1 | -29.273 | 25 | 20.137 | 0 | 0
|  |  | -27.1 | 0.00083378 | 1 | -29.263 | 25 | 20.137 | 0 | 0
|  |  | -27.05 | 0.00086297 | 1 | -29.253 | 25 | 20.137 | 0 | 0
|  |  | -27 | 0.00089216 | 1 | -29.243 | 25 | 20.137 | 0 | 0
|  |  | -26.95 | 0.00092135 | 1 | -29.233 | 25 | 20.137 | 0 | 0
|  |  | -26.9 | 0.00095109 | 1 | -29.223 | 25 | 20.137 | 0 | 0
|  |  | -26.85 | 0.00098082 | 1 | -29.213 | 25 | 20.137 | 0 | 0
|  |  | -26.8 | 0.0010106 | 1 | -29.203 | 25 | 20.137 | 0 | 0
|  |  | -26.75 | 0.0010403 | 1 | -29.193 | 25 | 20.137 | 0 | 0
|  |  | -26.7 | 0.0010682 | 1 | -29.183 | 25 | 20.137 | 0 | 0
|  |  | -26.65 | 0.0010962 | 1 | -29.173 | 25 | 20.137 | 0 | 0
|  |  | -26.6 | 0.0011242 | 1 | -29.163 | 25 | 20.137 | 0 | 0
|  |  | -26.55 | 0.0011521 | 1 | -29.153 | 25 | 20.137 | 0 | 0
|  |  | -26.5 | 0.0011732 | 1 | -29.143 | 25 | 20.137 | 0 | 0
|  |  | -26.45 | 0.0011942 | 1 | -29.133 | 25 | 20.137 | 0 | 0
|  |  | -26.4 | 0.0012153 | 1 | -29.123 | 25 | 20.137 | 0 | 0
|  |  | -26.35 | 0.0012363 | 1 | -29.113 | 25 | 20.137 | 0 | 0
|  |  | -26.3 | 0.0012439 | 1 | -29.103 | 25 | 20.137 | 0 | 0
|  |  | -26.25 | 0.0012514 | 1 | -29.093 | 25 | 20.137 | 0 | 0
|  |  | -26.2 | 0.001259 | 1 | -29.083 | 25 | 20.137 | 0 | 0
|  |  | -26.15 | 0.0012666 | 1 | -29.073 | 25 | 20.137 | 0 | 0
|  |  | -26.1 | 0.001276 | 1 | -29.063 | 25 | 20.137 | 0 | 0
|  |  | -26.05 | 0.0012855 | 1 | -29.053 | 25 | 20.137 | 0 | 0
|  |  | -26 | 0.001295 | 1 | -29.043 | 25 | 20.137 | 0 | 0
|  |  | -25.95 | 0.0013045 | 1 | -29.033 | 25 | 20.137 | 0 | 0
|  |  | -25.9 | 0.0013158 | 1 | -29.023 | 25 | 20.137 | 0 | 0
|  |  | -25.85 | 0.0013271 | 1 | -29.013 | 25 | 20.137 | 0 | 0
|  |  | -25.8 | 0.0013384 | 1 | -29.003 | 25 | 20.137 | 0 | 0
|  |  | -25.75 | 0.0013497 | 1 | -28.993 | 25 | 20.137 | 0 | 0
|  |  | -25.7 | 0.0013577 | 1 | -28.983 | 25 | 20.137 | 0 | 0
|  |  | -25.65 | 0.0013656 | 1 | -28.973 | 25 | 20.137 | 0 | 0
|  |  | -25.6 | 0.0013736 | 1 | -28.963 | 25 | 20.137 | 0 | 0
|  |  | -25.55 | 0.0013816 | 1 | -28.953 | 25 | 20.137 | 0 | 0
|  |  | -25.5 | 0.0013972 | 1 | -28.943 | 25 | 20.137 | 0 | 0
|  |  | -25.45 | 0.0014128 | 1 | -28.933 | 25 | 20.137 | 0 | 0
|  |  | -25.4 | 0.0014283 | 1 | -28.923 | 25 | 20.137 | 0 | 0
|  |  | -25.35 | 0.0014439 | 1 | -28.913 | 25 | 20.137 | 0 | 0
|  |  | -25.3 | 0.0014534 | 1 | -28.903 | 25 | 20.137 | 0 | 0
|  |  | -25.25 | 0.0014629 | 1 | -28.893 | 25 | 20.138 | 0 | 0
|  |  | -25.2 | 0.0014723 | 1 | -28.883 | 25 | 20.138 | 0 | 0
|  |  | -25.15 | 0.0014818 | 1 | -28.873 | 25 | 20.138 | 0 | 0
|  |  | -25.1 | 0.001494 | 1 | -28.863 | 25 | 20.138 | 0 | 0
|  |  | -25.05 | 0.0015062 | 1 | -28.853 | 25 | 20.138 | 0 | 0
|  |  | -25 | 0.0015185 | 1 | -28.843 | 25 | 20.138 | 0 | 0
|  |  | -24.95 | 0.0015307 | 1 | -28.833 | 25 | 20.138 | 0 | 0
|  |  | -24.9 | 0.0015487 | 1 | -28.823 | 25 | 20.138 | 0 | 0
|  |  | -24.85 | 0.0015666 | 1 | -28.813 | 25 | 20.138 | 0 | 0
|  |  | -24.8 | 0.0015846 | 1 | -28.803 | 25 | 20.138 | 0 | 0
|  |  | -24.75 | 0.0016026 | 1 | -28.793 | 25 | 20.138 | 0 | 0
|  |  | -24.7 | 0.0016127 | 1 | -28.783 | 25 | 20.138 | 0 | 0
|  |  | -24.65 | 0.0016227 | 1 | -28.773 | 25 | 20.138 | 0 | 0
|  |  | -24.6 | 0.0016328 | 1 | -28.763 | 25 | 20.138 | 0 | 0
|  |  | -24.55 | 0.0016429 | 1 | -28.753 | 25 | 20.138 | 0 | 0
|  |  | -24.5 | 0.0016678 | 1 | -28.743 | 25 | 20.138 | 0 | 0
|  |  | -24.45 | 0.0016928 | 1 | -28.733 | 25 | 20.138 | 0 | 0
|  |  | -24.4 | 0.0017178 | 1 | -28.723 | 25 | 20.138 | 0 | 0
|  |  | -24.35 | 0.0017427 | 1 | -28.713 | 25 | 20.138 | 0 | 0
|  |  | -24.3 | 0.0017604 | 1 | -28.703 | 25 | 20.138 | 0 | 0
|  |  | -24.25 | 0.0017781 | 1 | -28.693 | 25 | 20.138 | 0 | 0
|  |  | -24.2 | 0.0017958 | 1 | -28.683 | 25 | 20.138 | 0 | 0
|  |  | -24.15 | 0.0018135 | 1 | -28.673 | 25 | 20.138 | 0 | 0
|  |  | -24.1 | 0.0018346 | 1 | -28.663 | 25 | 20.138 | 0 | 0
|  |  | -24.05 | 0.0018556 | 1 | -28.653 | 25 | 20.138 | 0 | 0
|  |  | -24 | 0.0018767 | 1 | -28.643 | 25 | 20.138 | 0 | 0
|  |  | -23.95 | 0.0018978 | 1 | -28.633 | 25 | 20.138 | 0 | 0
|  |  | -23.9 | 0.0019349 | 1 | -28.623 | 25 | 20.138 | 0 | 0
|  |  | -23.85 | 0.001972 | 1 | -28.613 | 25 | 20.138 | 0 | 0
|  |  | -23.8 | 0.0020092 | 1 | -28.603 | 25 | 20.138 | 0 | 0
|  |  | -23.75 | 0.0020463 | 1 | -28.593 | 25 | 20.138 | 0 | 0
|  |  | -23.7 | 0.0020663 | 1 | -28.583 | 25 | 20.138 | 0 | 0
|  |  | -23.65 | 0.0020863 | 1 | -28.573 | 25 | 20.138 | 0 | 0
|  |  | -23.6 | 0.0021064 | 1 | -28.563 | 25 | 20.138 | 0 | 0
|  |  | -23.55 | 0.0021264 | 1 | -28.553 | 25 | 20.138 | 0 | 0
|  |  | -23.5 | 0.002176 | 1 | -28.543 | 25 | 20.138 | 0 | 0
|  |  | -23.45 | 0.0022255 | 1 | -28.533 | 25 | 20.138 | 0 | 0
|  |  | -23.4 | 0.0022751 | 1 | -28.523 | 25 | 20.138 | 0 | 0
|  |  | -23.35 | 0.0023247 | 1 | -28.513 | 25 | 20.138 | 0 | 0
|  |  | -23.3 | 0.0023634 | 1 | -28.503 | 25 | 20.138 | 0 | 0
|  |  | -23.25 | 0.0024022 | 1 | -28.493 | 25 | 20.138 | 0 | 0
|  |  | -23.2 | 0.0024409 | 1 | -28.483 | 25 | 20.138 | 0 | 0
|  |  | -23.15 | 0.0024797 | 1 | -28.473 | 25 | 20.138 | 0 | 0
|  |  | -23.1 | 0.0025182 | 1 | -28.463 | 25 | 20.138 | 0 | 0
|  |  | -23.05 | 0.0025568 | 1 | -28.453 | 25 | 20.138 | 0 | 0
|  |  | -23 | 0.0025954 | 1 | -28.443 | 25 | 20.138 | 0 | 0
|  |  | -22.95 | 0.002634 | 1 | -28.433 | 25 | 20.138 | 0 | 0
|  |  | -22.9 | 0.0027134 | 1 | -28.423 | 25 | 20.138 | 0 | 0
|  |  | -22.85 | 0.0027928 | 1 | -28.413 | 25 | 20.138 | 0 | 0
|  |  | -22.8 | 0.0028723 | 1 | -28.403 | 25 | 20.138 | 0 | 0
|  |  | -22.75 | 0.0029517 | 1 | -28.393 | 25 | 20.138 | 0 | 0
|  |  | -22.7 | 0.0029926 | 1 | -28.383 | 25 | 20.138 | 0 | 0
|  |  | -22.65 | 0.0030334 | 1 | -28.373 | 25 | 20.138 | 0 | 0
|  |  | -22.6 | 0.0030742 | 1 | -28.363 | 25 | 20.138 | 0 | 0
|  |  | -22.55 | 0.0031151 | 1 | -28.353 | 25 | 20.138 | 0 | 0
|  |  | -22.5 | 0.0032133 | 1 | -28.343 | 25 | 20.138 | 0 | 0
|  |  | -22.45 | 0.0033114 | 1 | -28.333 | 25 | 20.138 | 0 | 0
|  |  | -22.4 | 0.0034096 | 1 | -28.323 | 25 | 20.138 | 0 | 0
|  |  | -22.35 | 0.0035077 | 1 | -28.313 | 25 | 20.138 | 0 | 0
|  |  | -22.3 | 0.0035956 | 1 | -28.303 | 25 | 20.138 | 0 | 0
|  |  | -22.25 | 0.0036835 | 1 | -28.293 | 25 | 20.138 | 0 | 0
|  |  | -22.2 | 0.0037714 | 1 | -28.283 | 25 | 20.138 | 0 | 0
|  |  | -22.15 | 0.0038592 | 1 | -28.273 | 25 | 20.138 | 0 | 0
|  |  | -22.1 | 0.0039332 | 1 | -28.263 | 25 | 20.138 | 0 | 0
|  |  | -22.05 | 0.0040072 | 1 | -28.253 | 25 | 20.139 | 0 | 0
|  |  | -22 | 0.0040812 | 1 | -28.243 | 25 | 20.139 | 0 | 0
|  |  | -21.95 | 0.0041552 | 1 | -28.233 | 25 | 20.139 | 0 | 0
|  |  | -21.9 | 0.0043284 | 1 | -28.223 | 25 | 20.139 | 0 | 0
|  |  | -21.85 | 0.0045016 | 1 | -28.213 | 25 | 20.139 | 0 | 0
|  |  | -21.8 | 0.0046748 | 1 | -28.203 | 25 | 20.139 | 0 | 0
|  |  | -21.75 | 0.0048479 | 1 | -28.193 | 25 | 20.139 | 0 | 0
|  |  | -21.7 | 0.0049354 | 1 | -28.183 | 25 | 20.139 | 0 | 0
|  |  | -21.65 | 0.0050228 | 1 | -28.173 | 25 | 20.139 | 0 | 0
|  |  | -21.6 | 0.0051102 | 1 | -28.163 | 25 | 20.139 | 0 | 0
|  |  | -21.55 | 0.0051977 | 1 | -28.153 | 25 | 20.139 | 0 | 0
|  |  | -21.5 | 0.0053914 | 1 | -28.143 | 25 | 20.139 | 0 | 0
|  |  | -21.45 | 0.0055851 | 1 | -28.133 | 25 | 20.139 | 0 | 0
|  |  | -21.4 | 0.0057788 | 1 | -28.123 | 25 | 20.139 | 0 | 0
|  |  | -21.35 | 0.0059725 | 1 | -28.113 | 25 | 20.139 | 0 | 0
|  |  | -21.3 | 0.006173 | 1 | -28.103 | 25 | 20.139 | 0 | 0
|  |  | -21.25 | 0.0063735 | 1 | -28.093 | 25 | 20.139 | 0 | 0
|  |  | -21.2 | 0.006574 | 1 | -28.083 | 25 | 20.139 | 0 | 0
|  |  | -21.15 | 0.0067745 | 1 | -28.073 | 25 | 20.139 | 0 | 0
|  |  | -21.1 | 0.0069196 | 1 | -28.063 | 25 | 20.139 | 0 | 0
|  |  | -21.05 | 0.0070647 | 1 | -28.053 | 25 | 20.139 | 0 | 0
|  |  | -21 | 0.0072098 | 1 | -28.043 | 25 | 20.139 | 0 | 0
|  |  | -20.95 | 0.0073549 | 1 | -28.033 | 25 | 20.139 | 0 | 0
|  |  | -20.9 | 0.0077348 | 1 | -28.023 | 25 | 20.139 | 0 | 0
|  |  | -20.85 | 0.0081147 | 1 | -28.013 | 25 | 20.139 | 0 | 0
|  |  | -20.8 | 0.0084945 | 1 | -28.003 | 25 | 20.139 | 0 | 0
|  |  | -20.75 | 0.0088744 | 1 | -27.993 | 25 | 20.139 | 0 | 0
|  |  | -20.7 | 0.0090712 | 1 | -27.983 | 25 | 20.139 | 0 | 0
|  |  | -20.65 | 0.0092679 | 1 | -27.973 | 25 | 20.139 | 0 | 0
|  |  | -20.6 | 0.0094647 | 1 | -27.963 | 25 | 20.139 | 0 | 0
|  |  | -20.55 | 0.0096615 | 1 | -27.953 | 25 | 20.139 | 0 | 0
|  |  | -20.5 | 0.010048 | 1 | -27.943 | 25 | 20.139 | 0 | 0
|  |  | -20.45 | 0.010434 | 1 | -27.933 | 25 | 20.139 | 0 | 0
|  |  | -20.4 | 0.01082 | 1 | -27.923 | 25 | 20.139 | 0 | 0
|  |  | -20.35 | 0.011206 | 1 | -27.913 | 25 | 20.139 | 0 | 0
|  |  | -20.3 | 0.011668 | 1 | -27.903 | 25 | 20.139 | 0 | 0
|  |  | -20.25 | 0.01213 | 1 | -27.893 | 25 | 20.139 | 0 | 0
|  |  | -20.2 | 0.012592 | 1 | -27.883 | 25 | 20.139 | 0 | 0
|  |  | -20.15 | 0.013054 | 1 | -27.873 | 25 | 20.139 | 0 | 0
|  |  | -20.1 | 0.013343 | 1 | -27.863 | 25 | 20.139 | 0 | 0
|  |  | -20.05 | 0.013631 | 1 | -27.853 | 25 | 20.139 | 0 | 0
|  |  | -20 | 0.013919 | 1 | -27.843 | 25 | 20.139 | 0 | 0
|  |  | -19.95 | 0.014208 | 1 | -27.833 | 25 | 20.139 | 0 | 0
|  |  | -19.9 | 0.015007 | 1 | -27.823 | 25 | 20.139 | 0 | 0
|  |  | -19.85 | 0.015805 | 1 | -27.813 | 25 | 20.139 | 0 | 0
|  |  | -19.8 | 0.016604 | 1 | -27.803 | 25 | 20.139 | 0 | 0
|  |  | -19.75 | 0.017403 | 1 | -27.793 | 25 | 20.139 | 0 | 0
|  |  | -19.7 | 0.017832 | 1 | -27.783 | 25 | 20.139 | 0 | 0
|  |  | -19.65 | 0.018261 | 1 | -27.773 | 25 | 20.139 | 0 | 0
|  |  | -19.6 | 0.01869 | 1 | -27.763 | 25 | 20.139 | 0 | 0
|  |  | -19.55 | 0.019119 | 1 | -27.753 | 25 | 20.139 | 0 | 0
|  |  | -19.5 | 0.019856 | 1 | -27.743 | 25 | 20.139 | 0 | 0
|  |  | -19.45 | 0.020593 | 1 | -27.733 | 25 | 20.139 | 0 | 0
|  |  | -19.4 | 0.021329 | 1 | -27.723 | 25 | 20.139 | 0 | 0
|  |  | -19.35 | 0.022066 | 1 | -27.713 | 25 | 20.139 | 0 | 0
|  |  | -19.3 | 0.023058 | 1 | -27.703 | 25 | 20.139 | 0 | 0
|  |  | -19.25 | 0.02405 | 1 | -27.693 | 25 | 20.139 | 0 | 0
|  |  | -19.2 | 0.025042 | 1 | -27.683 | 25 | 20.139 | 0 | 0
|  |  | -19.15 | 0.026034 | 1 | -27.673 | 25 | 20.139 | 0 | 0
|  |  | -19.1 | 0.026592 | 1 | -27.663 | 25 | 20.139 | 0 | 0
|  |  | -19.05 | 0.02715 | 1 | -27.653 | 25 | 20.139 | 0 | 0
|  |  | -19 | 0.027708 | 1 | -27.643 | 25 | 20.139 | 0 | 0
|  |  | -18.95 | 0.028266 | 1 | -27.633 | 25 | 20.139 | 0 | 0
|  |  | -18.9 | 0.029839 | 1 | -27.623 | 25 | 20.14 | 0 | 0
|  |  | -18.85 | 0.031413 | 1 | -27.613 | 25 | 20.14 | 0 | 0
|  |  | -18.8 | 0.032986 | 1 | -27.603 | 25 | 20.14 | 0 | 0
|  |  | -18.75 | 0.034559 | 1 | -27.593 | 25 | 20.14 | 0 | 0
|  |  | -18.7 | 0.035471 | 1 | -27.583 | 25 | 20.14 | 0 | 0
|  |  | -18.65 | 0.036384 | 1 | -27.573 | 25 | 20.14 | 0 | 0
|  |  | -18.6 | 0.037296 | 1 | -27.563 | 25 | 20.14 | 0 | 0
|  |  | -18.55 | 0.038208 | 1 | -27.553 | 25 | 20.14 | 0 | 0
|  |  | -18.5 | 0.039521 | 1 | -27.543 | 25 | 20.14 | 0 | 0
|  |  | -18.45 | 0.040834 | 1 | -27.533 | 25 | 20.14 | 0 | 0
|  |  | -18.4 | 0.042147 | 1 | -27.523 | 25 | 20.14 | 0 | 0
|  |  | -18.35 | 0.043469 | 1 | -27.513 | 25 | 20.14 | 0 | 0
|  |  | -18.3 | 0.13977 | 1 | -27.503 | 25 | 20.14 | 0 | 0
|  |  | -18.25 | 0.23608 | 1 | -27.493 | 25 | 20.14 | 0 | 0
|  |  | -18.2 | 0.33239 | 1 | -27.483 | 25 | 20.14 | 0 | 0
|  |  | -18.15 | 0.42868 | 1 | -27.473 | 25 | 20.14 | 0 | 0
|  |  | -18.1 | 0.44055 | 1 | -27.463 | 25 | 20.14 | 0 | 0
|  |  | -18.05 | 0.45243 | 1 | -27.453 | 25 | 20.14 | 0 | 0
|  |  | -18 | 0.4643 | 1 | -27.443 | 25 | 20.14 | 0 | 0
|  |  | -17.95 | 0.47617 | 1 | -27.433 | 25 | 20.14 | 0 | 0
|  |  | -17.9 | 0.48805 | 1 | -27.423 | 25 | 20.14 | 0 | 0
|  |  | -17.85 | 0.49994 | 1 | -27.413 | 25 | 20.14 | 0 | 0
|  |  | -17.8 | 0.51182 | 1 | -27.403 | 25 | 20.14 | 0 | 0
|  |  | -17.75 | 0.52371 | 1 | -27.393 | 25 | 20.14 | 0 | 0
|  |  | -17.7 | 0.5356 | 1 | -27.383 | 25 | 20.14 | 0 | 0
|  |  | -17.65 | 0.5475 | 1 | -27.373 | 25 | 20.14 | 0 | 0
|  |  | -17.6 | 0.5594 | 1 | -27.363 | 25 | 20.14 | 0 | 0
|  |  | -17.55 | 0.5713 | 1 | -27.353 | 25 | 20.14 | 0 | 0
|  |  | -17.5 | 0.5832 | 1 | -27.343 | 25 | 20.14 | 0 | 0
|  |  | -17.45 | 0.59511 | 1 | -27.333 | 25 | 20.14 | 0 | 0
|  |  | -17.4 | 0.60702 | 1 | -27.323 | 25 | 20.14 | 0 | 0
|  |  | -17.35 | 0.61893 | 1 | -27.313 | 25 | 20.14 | 0 | 0
|  |  | -17.3 | 0.62442 | 1 | -27.303 | 25 | 20.14 | 0 | 0
|  |  | -17.25 | 0.6299 | 1 | -27.293 | 25 | 20.14 | 0 | 0
|  |  | -17.2 | 0.63538 | 1 | -27.283 | 25 | 20.14 | 0 | 0
|  |  | -17.15 | 0.64086 | 1 | -27.273 | 25 | 20.14 | 0 | 0
|  |  | -17.1 | 0.64635 | 1 | -27.263 | 25 | 20.14 | 0 | 0
|  |  | -17.05 | 0.65184 | 1 | -27.253 | 25 | 20.14 | 0 | 0
|  |  | -17 | 0.65732 | 1 | -27.243 | 25 | 20.14 | 0 | 0
|  |  | -16.95 | 0.66281 | 1 | -27.233 | 25 | 20.14 | 0 | 0
|  |  | -16.9 | 0.6683 | 1 | -27.223 | 25 | 20.14 | 0 | 0
|  |  | -16.85 | 0.67379 | 1 | -27.213 | 25 | 20.14 | 0 | 0
|  |  | -16.8 | 0.67929 | 1 | -27.203 | 25 | 20.14 | 0 | 0
|  |  | -16.75 | 0.68478 | 1 | -27.193 | 25 | 20.14 | 0 | 0
|  |  | -16.7 | 0.69027 | 1 | -27.183 | 25 | 20.14 | 0 | 0
|  |  | -16.65 | 0.69577 | 1 | -27.173 | 25 | 20.14 | 0 | 0
|  |  | -16.6 | 0.70127 | 1 | -27.163 | 25 | 20.14 | 0 | 0
|  |  | -16.55 | 0.70677 | 1 | -27.153 | 25 | 20.14 | 0 | 0
|  |  | -16.5 | 0.71367 | 1 | -27.143 | 25 | 20.14 | 0 | 0
|  |  | -16.45 | 0.72057 | 1 | -27.133 | 25 | 20.14 | 0 | 0
|  |  | -16.4 | 0.72747 | 1 | -27.123 | 25 | 20.14 | 0 | 0
|  |  | -16.35 | 0.73437 | 1 | -27.113 | 25 | 20.14 | 0 | 0
|  |  | -16.3 | 0.74128 | 1 | -27.103 | 25 | 20.14 | 0 | 0
|  |  | -16.25 | 0.74818 | 1 | -27.093 | 25 | 20.14 | 0 | 0
|  |  | -16.2 | 0.75509 | 1 | -27.083 | 25 | 20.14 | 0 | 0
|  |  | -16.15 | 0.762 | 1 | -27.073 | 25 | 20.14 | 0 | 0
|  |  | -16.1 | 0.76891 | 1 | -27.063 | 25 | 20.14 | 0 | 0
|  |  | -16.05 | 0.77582 | 1 | -27.053 | 25 | 20.14 | 0 | 0
|  |  | -16 | 0.78274 | 1 | -27.043 | 25 | 20.14 | 0 | 0
|  |  | -15.95 | 0.78965 | 1 | -27.033 | 25 | 20.14 | 0 | 0
|  |  | -15.9 | 0.79657 | 1 | -27.023 | 25 | 20.14 | 0 | 0
|  |  | -15.85 | 0.80349 | 1 | -27.013 | 25 | 20.14 | 0 | 0
|  |  | -15.8 | 0.81041 | 1 | -27.003 | 25 | 20.14 | 0 | 0
|  |  | -15.75 | 0.81733 | 1 | -26.993 | 25 | 20.14 | 0 | 0
|  |  | -15.7 | 0.82685 | 1 | -26.983 | 25 | 20.141 | 0 | 0
|  |  | -15.65 | 0.83637 | 1 | -26.973 | 25 | 20.141 | 0 | 0
|  |  | -15.6 | 0.84589 | 1 | -26.963 | 25 | 20.141 | 0 | 0
|  |  | -15.55 | 0.85542 | 1 | -26.953 | 25 | 20.141 | 0 | 0
|  |  | -15.5 | 0.86495 | 1 | -26.943 | 25 | 20.141 | 0 | 0
|  |  | -15.45 | 0.87448 | 1 | -26.933 | 25 | 20.141 | 0 | 0
|  |  | -15.4 | 0.88401 | 1 | -26.923 | 25 | 20.141 | 0 | 0
|  |  | -15.35 | 0.89354 | 1 | -26.913 | 25 | 20.141 | 0 | 0
|  |  | -15.3 | 0.90308 | 1 | -26.903 | 25 | 20.141 | 0 | 0
|  |  | -15.25 | 0.91262 | 1 | -26.893 | 25 | 20.141 | 0 | 0
|  |  | -15.2 | 0.92216 | 1 | -26.883 | 25 | 20.141 | 0 | 0
|  |  | -15.15 | 0.9317 | 1 | -26.873 | 25 | 20.141 | 0 | 0
|  |  | -15.1 | 0.94124 | 1 | -26.863 | 25 | 20.141 | 0 | 0
|  |  | -15.05 | 0.95079 | 1 | -26.853 | 25 | 20.141 | 0 | 0
|  |  | -15 | 0.96033 | 1 | -26.843 | 25 | 20.141 | 0 | 0
|  |  | -14.95 | 0.96988 | 1 | -26.833 | 25 | 20.141 | 0 | 0
|  |  | -14.9 | 0.9822 | 1 | -26.823 | 25 | 20.141 | 0 | 0
|  |  | -14.85 | 0.99452 | 1 | -26.813 | 25 | 20.141 | 0 | 0
|  |  | -14.8 | 1.0068 | 1 | -26.803 | 25 | 20.141 | 0 | 0
|  |  | -14.75 | 1.0192 | 1 | -26.793 | 25 | 20.141 | 0 | 0
|  |  | -14.7 | 1.0315 | 1 | -26.783 | 25 | 20.141 | 0 | 0
|  |  | -14.65 | 1.0438 | 1 | -26.773 | 25 | 20.141 | 0 | 0
|  |  | -14.6 | 1.0562 | 1 | -26.763 | 25 | 20.141 | 0 | 0
|  |  | -14.55 | 1.0685 | 1 | -26.753 | 25 | 20.141 | 0 | 0
|  |  | -14.5 | 1.0809 | 1 | -26.743 | 25 | 20.141 | 0 | 0
|  |  | -14.45 | 1.0932 | 1 | -26.733 | 25 | 20.141 | 0 | 0
|  |  | -14.4 | 1.1055 | 1 | -26.723 | 25 | 20.141 | 0 | 0
|  |  | -14.35 | 1.1179 | 1 | -26.713 | 25 | 20.141 | 0 | 0
|  |  | -14.3 | 1.1302 | 1 | -26.703 | 25 | 20.141 | 0 | 0
|  |  | -14.25 | 1.1426 | 1 | -26.693 | 25 | 20.141 | 0 | 0
|  |  | -14.2 | 1.1549 | 1 | -26.683 | 25 | 20.141 | 0 | 0
|  |  | -14.15 | 1.1673 | 1 | -26.673 | 25 | 20.141 | 0 | 0
|  |  | -14.1 | 1.1822 | 1 | -26.663 | 25 | 20.141 | 0 | 0
|  |  | -14.05 | 1.1971 | 1 | -26.653 | 25 | 20.141 | 0 | 0
|  |  | -14 | 1.2119 | 1 | -26.643 | 25 | 20.141 | 0 | 0
|  |  | -13.95 | 1.2268 | 1 | -26.633 | 25 | 20.141 | 0 | 0
|  |  | -13.9 | 1.2417 | 1 | -26.623 | 25 | 20.141 | 0 | 0
|  |  | -13.85 | 1.2566 | 1 | -26.613 | 25 | 20.141 | 0 | 0
|  |  | -13.8 | 1.2715 | 1 | -26.603 | 25 | 20.141 | 0 | 0
|  |  | -13.75 | 1.2864 | 1 | -26.593 | 25 | 20.141 | 0 | 0
|  |  | -13.7 | 1.3013 | 1 | -26.583 | 25 | 20.141 | 0 | 0
|  |  | -13.65 | 1.3162 | 1 | -26.573 | 25 | 20.141 | 0 | 0
|  |  | -13.6 | 1.3311 | 1 | -26.563 | 25 | 20.141 | 0 | 0
|  |  | -13.55 | 1.3461 | 1 | -26.553 | 25 | 20.141 | 0 | 0
|  |  | -13.5 | 1.361 | 1 | -26.543 | 25 | 20.141 | 0 | 0
|  |  | -13.45 | 1.3759 | 1 | -26.533 | 25 | 20.141 | 0 | 0
|  |  | -13.4 | 1.3908 | 1 | -26.523 | 25 | 20.141 | 0 | 0
|  |  | -13.35 | 1.4057 | 1 | -26.513 | 25 | 20.141 | 0 | 0
|  |  | -13.3 | 1.4414 | 1 | -26.503 | 25 | 20.141 | 0 | 0
|  |  | -13.25 | 1.4771 | 1 | -26.493 | 25 | 20.141 | 0 | 0
|  |  | -13.2 | 1.5128 | 1 | -26.483 | 25 | 20.141 | 0 | 0
|  |  | -13.15 | 1.5484 | 1 | -26.473 | 25 | 20.141 | 0 | 0
|  |  | -13.1 | 1.5841 | 1 | -26.463 | 25 | 20.141 | 0 | 0
|  |  | -13.05 | 1.6198 | 1 | -26.453 | 25 | 20.141 | 0 | 0
|  |  | -13 | 1.6555 | 1 | -26.443 | 25 | 20.141 | 0 | 0
|  |  | -12.95 | 1.6912 | 1 | -26.433 | 25 | 20.141 | 0 | 0
|  |  | -12.9 | 1.727 | 1 | -26.423 | 25 | 20.141 | 0 | 0
|  |  | -12.85 | 1.7627 | 1 | -26.413 | 25 | 20.141 | 0 | 0
|  |  | -12.8 | 1.7984 | 1 | -26.403 | 25 | 20.141 | 0 | 0
|  |  | -12.75 | 1.8342 | 1 | -26.393 | 25 | 20.141 | 0 | 0
|  |  | -12.7 | 1.8699 | 1 | -26.383 | 25 | 20.141 | 0 | 0
|  |  | -12.65 | 1.9057 | 1 | -26.373 | 25 | 20.141 | 0 | 0
|  |  | -12.6 | 1.9415 | 1 | -26.363 | 25 | 20.141 | 0 | 0
|  |  | -12.55 | 1.9772 | 1 | -26.353 | 25 | 20.142 | 0 | 0
|  |  | -12.5 | 2.0047 | 1 | -26.343 | 25 | 20.142 | 0 | 0
|  |  | -12.45 | 2.0323 | 1 | -26.333 | 25 | 20.142 | 0 | 0
|  |  | -12.4 | 2.0598 | 1 | -26.323 | 25 | 20.142 | 0 | 0
|  |  | -12.35 | 2.0873 | 1 | -26.313 | 25 | 20.142 | 0 | 0
|  |  | -12.3 | 2.1149 | 1 | -26.303 | 25 | 20.142 | 0 | 0
|  |  | -12.25 | 2.1424 | 1 | -26.293 | 25 | 20.142 | 0 | 0
|  |  | -12.2 | 2.17 | 1 | -26.283 | 25 | 20.142 | 0 | 0
|  |  | -12.15 | 2.1975 | 1 | -26.273 | 25 | 20.142 | 0 | 0
|  |  | -12.1 | 2.2251 | 1 | -26.263 | 25 | 20.142 | 0 | 0
|  |  | -12.05 | 2.2527 | 1 | -26.253 | 25 | 20.142 | 0 | 0
|  |  | -12 | 2.2802 | 1 | -26.243 | 25 | 20.142 | 0 | 0
|  |  | -11.95 | 2.3078 | 1 | -26.233 | 25 | 20.142 | 0 | 0
|  |  | -11.9 | 2.3354 | 1 | -26.223 | 25 | 20.142 | 0 | 0
|  |  | -11.85 | 2.363 | 1 | -26.213 | 25 | 20.142 | 0 | 0
|  |  | -11.8 | 2.3906 | 1 | -26.203 | 25 | 20.142 | 0 | 0
|  |  | -11.75 | 2.4182 | 1 | -26.193 | 25 | 20.142 | 0 | 0
|  |  | -11.7 | 2.4557 | 1 | -26.183 | 25 | 20.142 | 0 | 0
|  |  | -11.65 | 2.4932 | 1 | -26.173 | 25 | 20.142 | 0 | 0
|  |  | -11.6 | 2.5307 | 1 | -26.163 | 25 | 20.142 | 0 | 0
|  |  | -11.55 | 2.5682 | 1 | -26.153 | 25 | 20.142 | 0 | 0
|  |  | -11.5 | 2.6058 | 1 | -26.143 | 25 | 20.142 | 0 | 0
|  |  | -11.45 | 2.6433 | 1 | -26.133 | 25 | 20.142 | 0 | 0
|  |  | -11.4 | 2.6808 | 1 | -26.123 | 25 | 20.142 | 0 | 0
|  |  | -11.35 | 2.7184 | 1 | -26.113 | 25 | 20.142 | 0 | 0
|  |  | -11.3 | 2.7559 | 1 | -26.103 | 25 | 20.142 | 0 | 0
|  |  | -11.25 | 2.7935 | 1 | -26.093 | 25 | 20.142 | 0 | 0
|  |  | -11.2 | 2.831 | 1 | -26.083 | 25 | 20.142 | 0 | 0
|  |  | -11.15 | 2.8686 | 1 | -26.073 | 25 | 20.142 | 0 | 0
|  |  | -11.1 | 2.9062 | 1 | -26.063 | 25 | 20.142 | 0 | 0
|  |  | -11.05 | 2.9438 | 1 | -26.053 | 25 | 20.142 | 0 | 0
|  |  | -11 | 2.9813 | 1 | -26.043 | 25 | 20.142 | 0 | 0
|  |  | -10.95 | 3.0189 | 1 | -26.033 | 25 | 20.142 | 0 | 0
|  |  | -10.9 | 3.0723 | 1 | -26.023 | 25 | 20.142 | 0 | 0
|  |  | -10.85 | 3.1256 | 1 | -26.013 | 25 | 20.142 | 0 | 0
|  |  | -10.8 | 3.1789 | 1 | -26.003 | 25 | 20.142 | 0 | 0
|  |  | -10.75 | 3.2322 | 1 | -25.993 | 25 | 20.142 | 0 | 0
|  |  | -10.7 | 3.2856 | 1 | -25.983 | 25 | 20.142 | 0 | 0
|  |  | -10.65 | 3.3389 | 1 | -25.973 | 25 | 20.142 | 0 | 0
|  |  | -10.6 | 3.3923 | 1 | -25.963 | 25 | 20.142 | 0 | 0
|  |  | -10.55 | 3.4456 | 1 | -25.953 | 25 | 20.142 | 0 | 0
|  |  | -10.5 | 3.499 | 1 | -25.943 | 25 | 20.142 | 0 | 0
|  |  | -10.45 | 3.5524 | 1 | -25.933 | 25 | 20.142 | 0 | 0
|  |  | -10.4 | 3.6058 | 1 | -25.923 | 25 | 20.142 | 0 | 0
|  |  | -10.35 | 3.6591 | 1 | -25.913 | 25 | 20.142 | 0 | 0
|  |  | -10.3 | 3.7125 | 1 | -25.903 | 25 | 20.142 | 0 | 0
|  |  | -10.25 | 3.766 | 1 | -25.893 | 25 | 20.142 | 0 | 0
|  |  | -10.2 | 3.8194 | 1 | -25.883 | 25 | 20.142 | 0 | 0
|  |  | -10.15 | 3.8728 | 1 | -25.873 | 25 | 20.142 | 0 | 0
|  |  | -10.1 | 3.9376 | 1 | -25.863 | 25 | 20.142 | 0 | 0
|  |  | -10.05 | 4.0024 | 1 | -25.853 | 25 | 20.142 | 0 | 0
|  |  | -10 | 4.0672 | 1 | -25.843 | 25 | 20.142 | 0 | 0
|  |  | -9.95 | 4.1321 | 1 | -25.833 | 25 | 20.142 | 0 | 0
|  |  | -9.9 | 4.1969 | 1 | -25.823 | 25 | 20.142 | 0 | 0
|  |  | -9.85 | 4.2618 | 1 | -25.813 | 25 | 20.142 | 0 | 0
|  |  | -9.8 | 4.3266 | 1 | -25.803 | 25 | 20.142 | 0 | 0
|  |  | -9.75 | 4.3915 | 1 | -25.793 | 25 | 20.142 | 0 | 0
|  |  | -9.7 | 4.456 | 1 | -25.783 | 25 | 20.142 | 0 | 0
|  |  | -9.65 | 4.5206 | 1 | -25.773 | 25 | 20.142 | 0 | 0
|  |  | -9.6 | 4.5851 | 1 | -25.763 | 25 | 20.142 | 0 | 0
|  |  | -9.55 | 4.6497 | 1 | -25.753 | 25 | 20.142 | 0 | 0
|  |  | -9.5 | 4.7143 | 1 | -25.743 | 25 | 20.142 | 0 | 0
|  |  | -9.45 | 4.7788 | 1 | -25.733 | 25 | 20.142 | 0 | 0
|  |  | -9.4 | 4.8434 | 1 | -25.723 | 25 | 20.142 | 0 | 0
|  |  | -9.35 | 4.908 | 1 | -25.713 | 25 | 20.143 | 0 | 0
|  |  | -9.3 | 5.0416 | 1 | -25.703 | 25 | 20.143 | 0 | 0
|  |  | -9.25 | 5.1752 | 1 | -25.693 | 25 | 20.143 | 0 | 0
|  |  | -9.2 | 5.3088 | 1 | -25.683 | 25 | 20.143 | 0 | 0
|  |  | -9.15 | 5.4424 | 1 | -25.673 | 25 | 20.143 | 0 | 0
|  |  | -9.1 | 5.5759 | 1 | -25.663 | 25 | 20.143 | 0 | 0
|  |  | -9.05 | 5.7094 | 1 | -25.653 | 25 | 20.143 | 0 | 0
|  |  | -9 | 5.8429 | 1 | -25.643 | 25 | 20.143 | 0 | 0
|  |  | -8.95 | 5.9765 | 1 | -25.633 | 25 | 20.143 | 0 | 0
|  |  | -8.9 | 6.2701 | 1 | -25.623 | 25 | 20.143 | 0 | 0
|  |  | -8.85 | 6.5637 | 1 | -25.613 | 25 | 20.143 | 0 | 0
|  |  | -8.8 | 6.8574 | 1 | -25.603 | 25 | 20.143 | 0 | 0
|  |  | -8.75 | 7.1511 | 1 | -25.593 | 25 | 20.143 | 0 | 0
|  |  | -8.7 | 7.9774 | 1 | -25.583 | 25 | 20.143 | 0 | 0
|  |  | -8.65 | 8.8038 | 1 | -25.573 | 25 | 20.143 | 0 | 0
|  |  | -8.6 | 9.6302 | 1 | -25.563 | 25 | 20.143 | 0 | 0
|  |  | -8.55 | 10.457 | 1 | -25.553 | 25 | 20.143 | 0 | 0
|  |  | -8.5 | 13.814 | 1 | -25.543 | 25 | 20.143 | 0 | 0
|  |  | -8.45 | 17.172 | 1 | -25.533 | 25 | 20.143 | 0 | 0
|  |  | -8.4 | 20.53 | 1 | -25.523 | 25 | 20.143 | 0 | 0
|  |  | -8.35 | 23.888 | 1 | -25.513 | 25 | 20.143 | 0 | 0
|  |  | -8.3 | 27.882 | 1 | -25.503 | 25 | 20.143 | 0 | 0
|  |  | -8.25 | 31.876 | 1 | -25.493 | 25 | 20.143 | 0 | 0
|  |  | -8.2 | 35.87 | 1 | -25.483 | 25 | 20.143 | 0 | 0
|  |  | -8.15 | 39.864 | 1 | -25.473 | 25 | 20.143 | 0 | 0
|  |  | -8.1 | 41.327 | 1 | -25.463 | 25 | 20.143 | 0 | 0
|  |  | -8.05 | 42.79 | 1 | -25.453 | 25 | 20.143 | 0 | 0
|  |  | -8 | 44.253 | 1 | -25.443 | 25 | 20.143 | 0 | 0
|  |  | -7.95 | 45.716 | 1 | -25.433 | 25 | 20.143 | 0 | 0
|  |  | -7.9 | 46.184 | 1 | -25.423 | 25 | 20.143 | 0 | 0
|  |  | -7.85 | 46.651 | 1 | -25.413 | 25 | 20.143 | 0 | 0
|  |  | -7.8 | 47.119 | 1 | -25.403 | 25 | 20.143 | 0 | 0
|  |  | -7.75 | 47.587 | 1 | -25.393 | 25 | 20.143 | 0 | 0
|  |  | -7.7 | 47.829 | 1 | -25.383 | 25 | 20.143 | 0 | 0
|  |  | -7.65 | 48.07 | 1 | -25.373 | 25 | 20.143 | 0 | 0
|  |  | -7.6 | 48.311 | 1 | -25.363 | 25 | 20.143 | 0 | 0
|  |  | -7.55 | 48.553 | 1 | -25.353 | 25 | 20.143 | 0 | 0
|  |  | -7.5 | 48.707 | 1 | -25.343 | 25 | 20.143 | 0 | 0
|  |  | -7.45 | 48.862 | 1 | -25.333 | 25 | 20.143 | 0 | 0
|  |  | -7.4 | 49.017 | 1 | -25.323 | 25 | 20.143 | 0 | 0
|  |  | -7.35 | 49.172 | 1 | -25.313 | 25 | 20.143 | 0 | 0
|  |  | -7.3 | 49.276 | 1 | -25.303 | 25 | 20.143 | 0 | 0
|  |  | -7.25 | 49.38 | 1 | -25.293 | 25 | 20.143 | 0 | 0
|  |  | -7.2 | 49.485 | 1 | -25.283 | 25 | 20.143 | 0 | 0
|  |  | -7.15 | 49.589 | 1 | -25.273 | 25 | 20.143 | 0 | 0
|  |  | -7.1 | 49.698 | 1 | -25.263 | 25 | 20.143 | 0 | 0
|  |  | -7.05 | 49.806 | 1 | -25.253 | 25 | 20.143 | 0 | 0
|  |  | -7 | 49.915 | 1 | -25.243 | 25 | 20.143 | 0 | 0
|  |  | -6.95 | 50.023 | 1 | -25.233 | 25 | 20.143 | 0 | 0
|  |  | -6.9 | 50.086 | 1 | -25.223 | 25 | 20.143 | 0 | 0
|  |  | -6.85 | 50.148 | 1 | -25.213 | 25 | 20.143 | 0 | 0
|  |  | -6.8 | 50.211 | 1 | -25.203 | 25 | 20.143 | 0 | 0
|  |  | -6.75 | 50.274 | 1 | -25.193 | 25 | 20.143 | 0 | 0
|  |  | -6.7 | 50.336 | 1 | -25.183 | 25 | 20.143 | 0 | 0
|  |  | -6.65 | 50.399 | 1 | -25.173 | 25 | 20.143 | 0 | 0
|  |  | -6.6 | 50.461 | 1 | -25.163 | 25 | 20.143 | 0 | 0
|  |  | -6.55 | 50.524 | 1 | -25.153 | 25 | 20.143 | 0 | 0
|  |  | -6.5 | 50.586 | 1 | -25.143 | 25 | 20.143 | 0 | 0
|  |  | -6.45 | 50.648 | 1 | -25.133 | 25 | 20.143 | 0 | 0
|  |  | -6.4 | 50.711 | 1 | -25.123 | 25 | 20.143 | 0 | 0
|  |  | -6.35 | 50.773 | 1 | -25.113 | 25 | 20.143 | 0 | 0
|  |  | -6.3 | 50.836 | 1 | -25.103 | 25 | 20.143 | 0 | 0
|  |  | -6.25 | 50.898 | 1 | -25.093 | 25 | 20.143 | 0 | 0
|  |  | -6.2 | 50.96 | 1 | -25.083 | 25 | 20.143 | 0 | 0
|  |  | -6.15 | 51.023 | 1 | -25.073 | 25 | 20.144 | 0 | 0
|  |  | -6.1 | 51.063 | 1 | -25.063 | 25 | 20.144 | 0 | 0
|  |  | -6.05 | 51.102 | 1 | -25.053 | 25 | 20.144 | 0 | 0
|  |  | -6 | 51.142 | 1 | -25.043 | 25 | 20.144 | 0 | 0
|  |  | -5.95 | 51.182 | 1 | -25.033 | 25 | 20.144 | 0 | 0
|  |  | -5.9 | 51.221 | 1 | -25.023 | 25 | 20.144 | 0 | 0
|  |  | -5.85 | 51.261 | 1 | -25.013 | 25 | 20.144 | 0 | 0
|  |  | -5.8 | 51.3 | 1 | -25.003 | 25 | 20.144 | 0 | 0
|  |  | -5.75 | 51.34 | 1 | -24.993 | 25 | 20.144 | 0 | 0
|  |  | -5.7 | 51.38 | 1 | -24.983 | 25 | 20.144 | 0 | 0
|  |  | -5.65 | 51.419 | 1 | -24.973 | 25 | 20.144 | 0 | 0
|  |  | -5.6 | 51.459 | 1 | -24.963 | 25 | 20.144 | 0 | 0
|  |  | -5.55 | 51.498 | 1 | -24.953 | 25 | 20.144 | 0 | 0
|  |  | -5.5 | 51.538 | 1 | -24.943 | 25 | 20.144 | 0 | 0
|  |  | -5.45 | 51.577 | 1 | -24.933 | 25 | 20.144 | 0 | 0
|  |  | -5.4 | 51.617 | 1 | -24.923 | 25 | 20.144 | 0 | 0
|  |  | -5.35 | 51.656 | 1 | -24.913 | 25 | 20.144 | 0 | 0
|  |  | -5.3 | 51.687 | 1 | -24.903 | 25 | 20.144 | 0 | 0
|  |  | -5.25 | 51.717 | 1 | -24.893 | 25 | 20.144 | 0 | 0
|  |  | -5.2 | 51.748 | 1 | -24.883 | 25 | 20.144 | 0 | 0
|  |  | -5.15 | 51.779 | 1 | -24.873 | 25 | 20.144 | 0 | 0
|  |  | -5.1 | 51.809 | 1 | -24.863 | 25 | 20.144 | 0 | 0
|  |  | -5.05 | 51.84 | 1 | -24.853 | 25 | 20.144 | 0 | 0
|  |  | -5 | 51.87 | 1 | -24.843 | 25 | 20.144 | 0 | 0
|  |  | -4.95 | 51.901 | 1 | -24.833 | 25 | 20.144 | 0 | 0
|  |  | -4.9 | 51.931 | 1 | -24.823 | 25 | 20.144 | 0 | 0
|  |  | -4.85 | 51.962 | 1 | -24.813 | 25 | 20.144 | 0 | 0
|  |  | -4.8 | 51.992 | 1 | -24.803 | 25 | 20.144 | 0 | 0
|  |  | -4.75 | 52.023 | 1 | -24.793 | 25 | 20.144 | 0 | 0
|  |  | -4.7 | 52.053 | 1 | -24.783 | 25 | 20.144 | 0 | 0
|  |  | -4.65 | 52.083 | 1 | -24.773 | 25 | 20.144 | 0 | 0
|  |  | -4.6 | 52.114 | 1 | -24.763 | 25 | 20.144 | 0 | 0
|  |  | -4.55 | 52.144 | 1 | -24.754 | 25 | 20.144 | 0 | 0
|  |  | -4.5 | 52.16 | 1 | -24.744 | 25 | 20.144 | 0 | 0
|  |  | -4.45 | 52.175 | 1 | -24.734 | 25 | 20.144 | 0 | 0
|  |  | -4.4 | 52.19 | 1 | -24.724 | 25 | 20.144 | 0 | 0
|  |  | -4.35 | 52.206 | 1 | -24.714 | 25 | 20.144 | 0 | 0
|  |  | -4.3 | 52.221 | 1 | -24.704 | 25 | 20.144 | 0 | 0
|  |  | -4.25 | 52.236 | 1 | -24.694 | 25 | 20.144 | 0 | 0
|  |  | -4.2 | 52.251 | 1 | -24.684 | 25 | 20.144 | 0 | 0
|  |  | -4.15 | 52.266 | 1 | -24.674 | 25 | 20.144 | 0 | 0
|  |  | -4.1 | 52.282 | 1 | -24.664 | 25 | 20.144 | 0 | 0
|  |  | -4.05 | 52.297 | 1 | -24.654 | 25 | 20.144 | 0 | 0
|  |  | -4 | 52.312 | 1 | -24.644 | 25 | 20.144 | 0 | 0
|  |  | -3.95 | 52.327 | 1 | -24.634 | 25 | 20.144 | 0 | 0
|  |  | -3.9 | 52.342 | 1 | -24.624 | 25 | 20.144 | 0 | 0
|  |  | -3.85 | 52.357 | 1 | -24.614 | 25 | 20.144 | 0 | 0
|  |  | -3.8 | 52.372 | 1 | -24.604 | 25 | 20.144 | 0 | 0
|  |  | -3.75 | 52.387 | 1 | -24.594 | 25 | 20.144 | 0 | 0
|  |  | -3.7 | 52.397 | 1 | -24.584 | 25 | 20.144 | 0 | 0
|  |  | -3.65 | 52.406 | 1 | -24.574 | 25 | 20.144 | 0 | 0
|  |  | -3.6 | 52.415 | 1 | -24.564 | 25 | 20.144 | 0 | 0
|  |  | -3.55 | 52.424 | 1 | -24.554 | 25 | 20.144 | 0 | 0
|  |  | -3.5 | 52.433 | 1 | -24.544 | 25 | 20.144 | 0 | 0
|  |  | -3.45 | 52.443 | 1 | -24.534 | 25 | 20.144 | 0 | 0
|  |  | -3.4 | 52.452 | 1 | -24.524 | 25 | 20.144 | 0 | 0
|  |  | -3.35 | 52.461 | 1 | -24.514 | 25 | 20.144 | 0 | 0
|  |  | -3.3 | 52.47 | 1 | -24.504 | 25 | 20.144 | 0 | 0
|  |  | -3.25 | 52.479 | 1 | -24.494 | 25 | 20.144 | 0 | 0
|  |  | -3.2 | 52.488 | 1 | -24.484 | 25 | 20.144 | 0 | 0
|  |  | -3.15 | 52.497 | 1 | -24.474 | 25 | 20.144 | 0 | 0
|  |  | -3.1 | 52.506 | 1 | -24.464 | 25 | 20.144 | 0 | 0
|  |  | -3.05 | 52.515 | 1 | -24.454 | 25 | 20.144 | 0 | 0
|  |  | -3 | 52.524 | 1 | -24.444 | 25 | 20.145 | 0 | 0
|  |  | -2.95 | 52.533 | 1 | -24.434 | 25 | 20.145 | 0 | 0
|  |  | -2.9 | 52.538 | 1 | -24.424 | 25 | 20.145 | 0 | 0
|  |  | -2.85 | 52.544 | 1 | -24.414 | 25 | 20.145 | 0 | 0
|  |  | -2.8 | 52.549 | 1 | -24.404 | 25 | 20.145 | 0 | 0
|  |  | -2.75 | 52.554 | 1 | -24.394 | 25 | 20.145 | 0 | 0
|  |  | -2.7 | 52.56 | 1 | -24.384 | 25 | 20.145 | 0 | 0
|  |  | -2.65 | 52.565 | 1 | -24.374 | 25 | 20.145 | 0 | 0
|  |  | -2.6 | 52.57 | 1 | -24.364 | 25 | 20.145 | 0 | 0
|  |  | -2.55 | 52.575 | 1 | -24.354 | 25 | 20.145 | 0 | 0
|  |  | -2.5 | 52.58 | 1 | -24.344 | 25 | 20.145 | 0 | 0
|  |  | -2.45 | 52.586 | 1 | -24.334 | 25 | 20.145 | 0 | 0
|  |  | -2.4 | 52.591 | 1 | -24.324 | 25 | 20.145 | 0 | 0
|  |  | -2.35 | 52.596 | 1 | -24.314 | 25 | 20.145 | 0 | 0
|  |  | -2.3 | 52.601 | 1 | -24.304 | 25 | 20.145 | 0 | 0
|  |  | -2.25 | 52.606 | 1 | -24.294 | 25 | 20.145 | 0 | 0
|  |  | -2.2 | 52.611 | 1 | -24.284 | 25 | 20.145 | 0 | 0
|  |  | -2.15 | 52.616 | 1 | -24.274 | 25 | 20.145 | 0 | 0
|  |  | -2.1 | 52.618 | 1 | -24.264 | 25 | 20.145 | 0 | 0
|  |  | -2.05 | 52.62 | 1 | -24.254 | 25 | 20.145 | 0 | 0
|  |  | -2 | 52.622 | 1 | -24.244 | 25 | 20.145 | 0 | 0
|  |  | -1.95 | 52.624 | 1 | -24.234 | 25 | 20.145 | 0 | 0
|  |  | -1.9 | 52.626 | 1 | -24.224 | 25 | 20.145 | 0 | 0
|  |  | -1.85 | 52.627 | 1 | -24.214 | 25 | 20.145 | 0 | 0
|  |  | -1.8 | 52.629 | 1 | -24.204 | 25 | 20.145 | 0 | 0
|  |  | -1.75 | 52.631 | 1 | -24.194 | 25 | 20.145 | 0 | 0
|  |  | -1.7 | 52.633 | 1 | -24.184 | 25 | 20.145 | 0 | 0
|  |  | -1.65 | 52.634 | 1 | -24.174 | 25 | 20.145 | 0 | 0
|  |  | -1.6 | 52.636 | 1 | -24.164 | 25 | 20.145 | 0 | 0
|  |  | -1.55 | 52.638 | 1 | -24.154 | 25 | 20.145 | 0 | 0
|  |  | -1.5 | 52.64 | 1 | -24.144 | 25 | 20.145 | 0 | 0
|  |  | -1.45 | 52.641 | 1 | -24.134 | 25 | 20.145 | 0 | 0
|  |  | -1.4 | 52.643 | 1 | -24.124 | 25 | 20.145 | 0 | 0
|  |  | -1.35 | 52.644 | 1 | -24.114 | 25 | 20.145 | 0 | 0
|  |  | -1.3 | 52.641 | 1 | -24.104 | 25 | 20.145 | 0 | 0
|  |  | -1.25 | 52.637 | 1 | -24.094 | 25 | 20.145 | 0 | 0
|  |  | -1.2 | 52.633 | 1 | -24.084 | 25 | 20.145 | 0 | 0
|  |  | -1.15 | 52.63 | 1 | -24.074 | 25 | 20.145 | 0 | 0
|  |  | -1.1 | 52.626 | 1 | -24.064 | 25 | 20.145 | 0 | 0
|  |  | -1.05 | 52.622 | 1 | -24.054 | 25 | 20.145 | 0 | 0
|  |  | -1 | 52.618 | 1 | -24.044 | 25 | 20.145 | 0 | 0
|  |  | -0.95 | 52.614 | 1 | -24.034 | 25 | 20.145 | 0 | 0
|  |  | -0.9 | 52.61 | 1 | -24.024 | 25 | 20.145 | 0 | 0
|  |  | -0.85 | 52.607 | 1 | -24.014 | 25 | 20.145 | 0 | 0
|  |  | -0.8 | 52.603 | 1 | -24.004 | 25 | 20.145 | 0 | 0
|  |  | -0.75 | 52.599 | 1 | -23.994 | 25 | 20.145 | 0 | 0
|  |  | -0.7 | 52.595 | 1 | -23.984 | 25 | 20.145 | 0 | 0
|  |  | -0.65 | 52.591 | 1 | -23.974 | 25 | 20.145 | 0 | 0
|  |  | -0.6 | 52.587 | 1 | -23.964 | 25 | 20.145 | 0 | 0
|  |  | -0.55 | 52.583 | 1 | -23.954 | 25 | 20.145 | 0 | 0
|  |  | -0.5 | 52.583 | 1 | -23.944 | 25 | 20.145 | 0 | 0
|  |  | -0.45 | 52.582 | 1 | -23.934 | 25 | 20.145 | 0 | 0
|  |  | -0.4 | 52.582 | 1 | -23.924 | 25 | 20.145 | 0 | 0
|  |  | -0.35 | 52.582 | 1 | -23.914 | 25 | 20.145 | 0 | 0
|  |  | -0.3 | 52.581 | 1 | -23.904 | 25 | 20.145 | 0 | 0
|  |  | -0.25 | 52.581 | 1 | -23.894 | 25 | 20.145 | 0 | 0
|  |  | -0.2 | 52.581 | 1 | -23.884 | 25 | 20.145 | 0 | 0
|  |  | -0.15 | 52.58 | 1 | -23.874 | 25 | 20.145 | 0 | 0
|  |  | -0.1 | 52.58 | 1 | -23.864 | 25 | 20.145 | 0 | 0
|  |  | -0.05 | 52.579 | 1 | -23.854 | 25 | 20.145 | 0 | 0
|  |  | 0 | 52.579 | 1 | -23.844 | 25 | 20.145 | 0 | 0
|  |  | 0.05 | 52.578 | 1 | -23.834 | 25 | 20.145 | 0 | 0
|  |  | 0.1 | 52.578 | 1 | -23.824 | 25 | 20.145 | 0 | 0
|  |  | 0.15 | 52.577 | 1 | -23.814 | 25 | 20.145 | 0 | 0
|  |  | 0.2 | 52.577 | 1 | -23.804 | 25 | 20.146 | 0 | 0
|  |  | 0.25 | 52.576 | 1 | -23.794 | 25 | 20.146 | 0 | 0
|  |  | 0.3 | 52.579 | 1 | -23.784 | 25 | 20.146 | 0 | 0
|  |  | 0.35 | 52.581 | 1 | -23.774 | 25 | 20.146 | 0 | 0
|  |  | 0.4 | 52.583 | 1 | -23.764 | 25 | 20.146 | 0 | 0
|  |  | 0.45 | 52.586 | 1 | -23.754 | 25 | 20.146 | 0 | 0
|  |  | 0.5 | 52.588 | 1 | -23.744 | 25 | 20.146 | 0 | 0
|  |  | 0.55 | 52.59 | 1 | -23.734 | 25 | 20.146 | 0 | 0
|  |  | 0.6 | 52.593 | 1 | -23.724 | 25 | 20.146 | 0 | 0
|  |  | 0.65 | 52.595 | 1 | -23.714 | 25 | 20.146 | 0 | 0
|  |  | 0.7 | 52.597 | 1 | -23.704 | 25 | 20.146 | 0 | 0
|  |  | 0.75 | 52.599 | 1 | -23.694 | 25 | 20.146 | 0 | 0
|  |  | 0.8 | 52.602 | 1 | -23.684 | 25 | 20.146 | 0 | 0
|  |  | 0.85 | 52.604 | 1 | -23.674 | 25 | 20.146 | 0 | 0
|  |  | 0.9 | 52.606 | 1 | -23.664 | 25 | 20.146 | 0 | 0
|  |  | 0.95 | 52.608 | 1 | -23.654 | 25 | 20.146 | 0 | 0
|  |  | 1 | 52.61 | 1 | -23.644 | 25 | 20.146 | 0 | 0
|  |  | 1.05 | 52.612 | 1 | -23.634 | 25 | 20.146 | 0 | 0
|  |  | 1.1 | 52.614 | 1 | -23.624 | 25 | 20.146 | 0 | 0
|  |  | 1.15 | 52.615 | 1 | -23.614 | 25 | 20.146 | 0 | 0
|  |  | 1.2 | 52.617 | 1 | -23.604 | 25 | 20.146 | 0 | 0
|  |  | 1.25 | 52.618 | 1 | -23.594 | 25 | 20.146 | 0 | 0
|  |  | 1.3 | 52.62 | 1 | -23.584 | 25 | 20.146 | 0 | 0
|  |  | 1.35 | 52.621 | 1 | -23.574 | 25 | 20.146 | 0 | 0
|  |  | 1.4 | 52.622 | 1 | -23.564 | 25 | 20.146 | 0 | 0
|  |  | 1.45 | 52.624 | 1 | -23.554 | 25 | 20.146 | 0 | 0
|  |  | 1.5 | 52.625 | 1 | -23.544 | 25 | 20.146 | 0 | 0
|  |  | 1.55 | 52.626 | 1 | -23.534 | 25 | 20.146 | 0 | 0
|  |  | 1.6 | 52.628 | 1 | -23.524 | 25 | 20.146 | 0 | 0
|  |  | 1.65 | 52.629 | 1 | -23.514 | 25 | 20.146 | 0 | 0
|  |  | 1.7 | 52.63 | 1 | -23.504 | 25 | 20.146 | 0 | 0
|  |  | 1.75 | 52.631 | 1 | -23.494 | 25 | 20.146 | 0 | 0
|  |  | 1.8 | 52.632 | 1 | -23.484 | 25 | 20.146 | 0 | 0
|  |  | 1.85 | 52.634 | 1 | -23.474 | 25 | 20.146 | 0 | 0
|  |  | 1.9 | 52.629 | 1 | -23.464 | 25 | 20.146 | 0 | 0
|  |  | 1.95 | 52.625 | 1 | -23.454 | 25 | 20.146 | 0 | 0
|  |  | 2 | 52.621 | 1 | -23.444 | 25 | 20.146 | 0 | 0
|  |  | 2.05 | 52.617 | 1 | -23.434 | 25 | 20.146 | 0 | 0
|  |  | 2.1 | 52.612 | 1 | -23.424 | 25 | 20.146 | 0 | 0
|  |  | 2.15 | 52.608 | 1 | -23.414 | 25 | 20.146 | 0 | 0
|  |  | 2.2 | 52.604 | 1 | -23.404 | 25 | 20.146 | 0 | 0
|  |  | 2.25 | 52.599 | 1 | -23.394 | 25 | 20.146 | 0 | 0
|  |  | 2.3 | 52.595 | 1 | -23.384 | 25 | 20.146 | 0 | 0
|  |  | 2.35 | 52.59 | 1 | -23.374 | 25 | 20.146 | 0 | 0
|  |  | 2.4 | 52.586 | 1 | -23.364 | 25 | 20.146 | 0 | 0
|  |  | 2.45 | 52.581 | 1 | -23.354 | 25 | 20.146 | 0 | 0
|  |  | 2.5 | 52.577 | 1 | -23.344 | 25 | 20.146 | 0 | 0
|  |  | 2.55 | 52.572 | 1 | -23.334 | 25 | 20.146 | 0 | 0
|  |  | 2.6 | 52.568 | 1 | -23.324 | 25 | 20.146 | 0 | 0
|  |  | 2.65 | 52.563 | 1 | -23.314 | 25 | 20.146 | 0 | 0
|  |  | 2.7 | 52.555 | 1 | -23.304 | 25 | 20.146 | 0 | 0
|  |  | 2.75 | 52.547 | 1 | -23.294 | 25 | 20.146 | 0 | 0
|  |  | 2.8 | 52.539 | 1 | -23.284 | 25 | 20.146 | 0 | 0
|  |  | 2.85 | 52.531 | 1 | -23.274 | 25 | 20.146 | 0 | 0
|  |  | 2.9 | 52.523 | 1 | -23.264 | 25 | 20.146 | 0 | 0
|  |  | 2.95 | 52.515 | 1 | -23.254 | 25 | 20.146 | 0 | 0
|  |  | 3 | 52.507 | 1 | -23.244 | 25 | 20.146 | 0 | 0
|  |  | 3.05 | 52.499 | 1 | -23.234 | 25 | 20.146 | 0 | 0
|  |  | 3.1 | 52.49 | 1 | -23.224 | 25 | 20.146 | 0 | 0
|  |  | 3.15 | 52.482 | 1 | -23.214 | 25 | 20.146 | 0 | 0
|  |  | 3.2 | 52.474 | 1 | -23.204 | 25 | 20.146 | 0 | 0
|  |  | 3.25 | 52.466 | 1 | -23.194 | 25 | 20.146 | 0 | 0
|  |  | 3.3 | 52.457 | 1 | -23.184 | 25 | 20.146 | 0 | 0
|  |  | 3.35 | 52.449 | 1 | -23.174 | 25 | 20.147 | 0 | 0
|  |  | 3.4 | 52.441 | 1 | -23.164 | 25 | 20.147 | 0 | 0
|  |  | 3.45 | 52.433 | 1 | -23.154 | 25 | 20.147 | 0 | 0
|  |  | 3.5 | 52.419 | 1 | -23.144 | 25 | 20.147 | 0 | 0
|  |  | 3.55 | 52.406 | 1 | -23.134 | 25 | 20.147 | 0 | 0
|  |  | 3.6 | 52.393 | 1 | -23.124 | 25 | 20.147 | 0 | 0
|  |  | 3.65 | 52.379 | 1 | -23.114 | 25 | 20.147 | 0 | 0
|  |  | 3.7 | 52.366 | 1 | -23.104 | 25 | 20.147 | 0 | 0
|  |  | 3.75 | 52.352 | 1 | -23.094 | 25 | 20.147 | 0 | 0
|  |  | 3.8 | 52.339 | 1 | -23.084 | 25 | 20.147 | 0 | 0
|  |  | 3.85 | 52.326 | 1 | -23.074 | 25 | 20.147 | 0 | 0
|  |  | 3.9 | 52.312 | 1 | -23.064 | 25 | 20.147 | 0 | 0
|  |  | 3.95 | 52.299 | 1 | -23.054 | 25 | 20.147 | 0 | 0
|  |  | 4 | 52.285 | 1 | -23.044 | 25 | 20.147 | 0 | 0
|  |  | 4.05 | 52.272 | 1 | -23.034 | 25 | 20.147 | 0 | 0
|  |  | 4.1 | 52.258 | 1 | -23.024 | 25 | 20.147 | 0 | 0
|  |  | 4.15 | 52.244 | 1 | -23.014 | 25 | 20.147 | 0 | 0
|  |  | 4.2 | 52.231 | 1 | -23.004 | 25 | 20.147 | 0 | 0
|  |  | 4.25 | 52.217 | 1 | -22.994 | 25 | 20.147 | 0 | 0
|  |  | 4.3 | 52.191 | 1 | -22.984 | 25 | 20.147 | 0 | 0
|  |  | 4.35 | 52.165 | 1 | -22.974 | 25 | 20.147 | 0 | 0
|  |  | 4.4 | 52.139 | 1 | -22.964 | 25 | 20.147 | 0 | 0
|  |  | 4.45 | 52.113 | 1 | -22.954 | 25 | 20.147 | 0 | 0
|  |  | 4.5 | 52.087 | 1 | -22.944 | 25 | 20.147 | 0 | 0
|  |  | 4.55 | 52.061 | 1 | -22.934 | 25 | 20.147 | 0 | 0
|  |  | 4.6 | 52.035 | 1 | -22.924 | 25 | 20.147 | 0 | 0
|  |  | 4.65 | 52.009 | 1 | -22.914 | 25 | 20.147 | 0 | 0
|  |  | 4.7 | 51.982 | 1 | -22.904 | 25 | 20.147 | 0 | 0
|  |  | 4.75 | 51.956 | 1 | -22.894 | 25 | 20.147 | 0 | 0
|  |  | 4.8 | 51.93 | 1 | -22.884 | 25 | 20.147 | 0 | 0
|  |  | 4.85 | 51.904 | 1 | -22.874 | 25 | 20.147 | 0 | 0
|  |  | 4.9 | 51.878 | 1 | -22.864 | 25 | 20.147 | 0 | 0
|  |  | 4.95 | 51.851 | 1 | -22.854 | 25 | 20.147 | 0 | 0
|  |  | 5 | 51.825 | 1 | -22.844 | 25 | 20.147 | 0 | 0
|  |  | 5.05 | 51.799 | 1 | -22.834 | 25 | 20.147 | 0 | 0
|  |  | 5.1 | 51.761 | 1 | -22.824 | 25 | 20.147 | 0 | 0
|  |  | 5.15 | 51.722 | 1 | -22.814 | 25 | 20.147 | 0 | 0
|  |  | 5.2 | 51.684 | 1 | -22.804 | 25 | 20.147 | 0 | 0
|  |  | 5.25 | 51.646 | 1 | -22.794 | 25 | 20.147 | 0 | 0
|  |  | 5.3 | 51.607 | 1 | -22.784 | 25 | 20.147 | 0 | 0
|  |  | 5.35 | 51.569 | 1 | -22.774 | 25 | 20.147 | 0 | 0
|  |  | 5.4 | 51.531 | 1 | -22.764 | 25 | 20.147 | 0 | 0
|  |  | 5.45 | 51.492 | 1 | -22.754 | 25 | 20.147 | 0 | 0
|  |  | 5.5 | 51.454 | 1 | -22.744 | 25 | 20.147 | 0 | 0
|  |  | 5.55 | 51.415 | 1 | -22.734 | 25 | 20.147 | 0 | 0
|  |  | 5.6 | 51.377 | 1 | -22.724 | 25 | 20.147 | 0 | 0
|  |  | 5.65 | 51.338 | 1 | -22.714 | 25 | 20.147 | 0 | 0
|  |  | 5.7 | 51.3 | 1 | -22.704 | 25 | 20.147 | 0 | 0
|  |  | 5.75 | 51.261 | 1 | -22.694 | 25 | 20.147 | 0 | 0
|  |  | 5.8 | 51.223 | 1 | -22.684 | 25 | 20.147 | 0 | 0
|  |  | 5.85 | 51.184 | 1 | -22.674 | 25 | 20.147 | 0 | 0
|  |  | 5.9 | 51.13 | 1 | -22.664 | 25 | 20.147 | 0 | 0
|  |  | 5.95 | 51.076 | 1 | -22.654 | 25 | 20.147 | 0 | 0
|  |  | 6 | 51.022 | 1 | -22.644 | 25 | 20.147 | 0 | 0
|  |  | 6.05 | 50.968 | 1 | -22.634 | 25 | 20.147 | 0 | 0
|  |  | 6.1 | 50.914 | 1 | -22.624 | 25 | 20.147 | 0 | 0
|  |  | 6.15 | 50.86 | 1 | -22.614 | 25 | 20.147 | 0 | 0
|  |  | 6.2 | 50.805 | 1 | -22.604 | 25 | 20.147 | 0 | 0
|  |  | 6.25 | 50.751 | 1 | -22.594 | 25 | 20.147 | 0 | 0
|  |  | 6.3 | 50.697 | 1 | -22.584 | 25 | 20.147 | 0 | 0
|  |  | 6.35 | 50.643 | 1 | -22.574 | 25 | 20.147 | 0 | 0
|  |  | 6.4 | 50.588 | 1 | -22.564 | 25 | 20.147 | 0 | 0
|  |  | 6.45 | 50.534 | 1 | -22.554 | 25 | 20.147 | 0 | 0
|  |  | 6.5 | 50.48 | 1 | -22.544 | 25 | 20.147 | 0 | 0
|  |  | 6.55 | 50.426 | 1 | -22.534 | 25 | 20.148 | 0 | 0
|  |  | 6.6 | 50.371 | 1 | -22.524 | 25 | 20.148 | 0 | 0
|  |  | 6.65 | 50.317 | 1 | -22.514 | 25 | 20.148 | 0 | 0
|  |  | 6.7 | 50.206 | 1 | -22.504 | 25 | 20.148 | 0 | 0
|  |  | 6.75 | 50.095 | 1 | -22.494 | 25 | 20.148 | 0 | 0
|  |  | 6.8 | 49.983 | 1 | -22.484 | 25 | 20.148 | 0 | 0
|  |  | 6.85 | 49.872 | 1 | -22.474 | 25 | 20.148 | 0 | 0
|  |  | 6.9 | 49.761 | 1 | -22.464 | 25 | 20.148 | 0 | 0
|  |  | 6.95 | 49.649 | 1 | -22.454 | 25 | 20.148 | 0 | 0
|  |  | 7 | 49.538 | 1 | -22.444 | 25 | 20.148 | 0 | 0
|  |  | 7.05 | 49.427 | 1 | -22.434 | 25 | 20.148 | 0 | 0
|  |  | 7.1 | 49.364 | 1 | -22.424 | 25 | 20.148 | 0 | 0
|  |  | 7.15 | 49.3 | 1 | -22.414 | 25 | 20.148 | 0 | 0
|  |  | 7.2 | 49.237 | 1 | -22.404 | 25 | 20.148 | 0 | 0
|  |  | 7.25 | 49.173 | 1 | -22.394 | 25 | 20.148 | 0 | 0
|  |  | 7.3 | 49.017 | 1 | -22.384 | 25 | 20.148 | 0 | 0
|  |  | 7.35 | 48.861 | 1 | -22.374 | 25 | 20.148 | 0 | 0
|  |  | 7.4 | 48.705 | 1 | -22.364 | 25 | 20.148 | 0 | 0
|  |  | 7.45 | 48.549 | 1 | -22.354 | 25 | 20.148 | 0 | 0
|  |  | 7.5 | 48.307 | 1 | -22.344 | 25 | 20.148 | 0 | 0
|  |  | 7.55 | 48.066 | 1 | -22.334 | 25 | 20.148 | 0 | 0
|  |  | 7.6 | 47.825 | 1 | -22.324 | 25 | 20.148 | 0 | 0
|  |  | 7.65 | 47.584 | 1 | -22.314 | 25 | 20.148 | 0 | 0
|  |  | 7.7 | 47.116 | 1 | -22.304 | 25 | 20.148 | 0 | 0
|  |  | 7.75 | 46.648 | 1 | -22.294 | 25 | 20.148 | 0 | 0
|  |  | 7.8 | 46.181 | 1 | -22.284 | 25 | 20.148 | 0 | 0
|  |  | 7.85 | 45.713 | 1 | -22.274 | 25 | 20.148 | 0 | 0
|  |  | 7.9 | 44.251 | 1 | -22.264 | 25 | 20.148 | 0 | 0
|  |  | 7.95 | 42.789 | 1 | -22.254 | 25 | 20.148 | 0 | 0
|  |  | 8 | 41.327 | 1 | -22.244 | 25 | 20.148 | 0 | 0
|  |  | 8.05 | 39.864 | 1 | -22.234 | 25 | 20.148 | 0 | 0
|  |  | 8.1 | 35.87 | 1 | -22.224 | 25 | 20.148 | 0 | 0
|  |  | 8.15 | 31.876 | 1 | -22.214 | 25 | 20.148 | 0 | 0
|  |  | 8.2 | 27.881 | 1 | -22.204 | 25 | 20.148 | 0 | 0
|  |  | 8.25 | 23.887 | 1 | -22.194 | 25 | 20.148 | 0 | 0
|  |  | 8.3 | 20.53 | 1 | -22.184 | 25 | 20.148 | 0 | 0
|  |  | 8.35 | 17.172 | 1 | -22.174 | 25 | 20.148 | 0 | 0
|  |  | 8.4 | 13.814 | 1 | -22.164 | 25 | 20.148 | 0 | 0
|  |  | 8.45 | 10.457 | 1 | -22.154 | 25 | 20.148 | 0 | 0
|  |  | 8.5 | 9.6306 | 1 | -22.144 | 25 | 20.148 | 0 | 0
|  |  | 8.55 | 8.8046 | 1 | -22.134 | 25 | 20.148 | 0 | 0
|  |  | 8.6 | 7.9786 | 1 | -22.124 | 25 | 20.148 | 0 | 0
|  |  | 8.65 | 7.1527 | 1 | -22.114 | 25 | 20.148 | 0 | 0
|  |  | 8.7 | 6.9258 | 1 | -22.104 | 25 | 20.148 | 0 | 0
|  |  | 8.75 | 6.6988 | 1 | -22.094 | 25 | 20.148 | 0 | 0
|  |  | 8.8 | 6.4719 | 1 | -22.084 | 25 | 20.148 | 0 | 0
|  |  | 8.85 | 6.2449 | 1 | -22.074 | 25 | 20.148 | 0 | 0
|  |  | 8.9 | 6.0181 | 1 | -22.064 | 25 | 20.148 | 0 | 0
|  |  | 8.95 | 5.7913 | 1 | -22.054 | 25 | 20.148 | 0 | 0
|  |  | 9 | 5.5645 | 1 | -22.044 | 25 | 20.148 | 0 | 0
|  |  | 9.05 | 5.3377 | 1 | -22.034 | 25 | 20.148 | 0 | 0
|  |  | 9.1 | 5.2595 | 1 | -22.024 | 25 | 20.148 | 0 | 0
|  |  | 9.15 | 5.1814 | 1 | -22.014 | 25 | 20.148 | 0 | 0
|  |  | 9.2 | 5.1032 | 1 | -22.004 | 25 | 20.148 | 0 | 0
|  |  | 9.25 | 5.0251 | 1 | -21.994 | 25 | 20.148 | 0 | 0
|  |  | 9.3 | 4.947 | 1 | -21.984 | 25 | 20.148 | 0 | 0
|  |  | 9.35 | 4.8689 | 1 | -21.974 | 25 | 20.148 | 0 | 0
|  |  | 9.4 | 4.7908 | 1 | -21.964 | 25 | 20.148 | 0 | 0
|  |  | 9.45 | 4.7127 | 1 | -21.954 | 25 | 20.148 | 0 | 0
|  |  | 9.5 | 4.6346 | 1 | -21.944 | 25 | 20.148 | 0 | 0
|  |  | 9.55 | 4.5566 | 1 | -21.934 | 25 | 20.148 | 0 | 0
|  |  | 9.6 | 4.4785 | 1 | -21.924 | 25 | 20.148 | 0 | 0
|  |  | 9.65 | 4.4004 | 1 | -21.914 | 25 | 20.148 | 0 | 0
|  |  | 9.7 | 4.3224 | 1 | -21.904 | 25 | 20.148 | 0 | 0
|  |  | 9.75 | 4.2444 | 1 | -21.894 | 25 | 20.149 | 0 | 0
|  |  | 9.8 | 4.1664 | 1 | -21.884 | 25 | 20.149 | 0 | 0
|  |  | 9.85 | 4.0884 | 1 | -21.874 | 25 | 20.149 | 0 | 0
|  |  | 9.9 | 4.0339 | 1 | -21.864 | 25 | 20.149 | 0 | 0
|  |  | 9.95 | 3.9795 | 1 | -21.854 | 25 | 20.149 | 0 | 0
|  |  | 10 | 3.925 | 1 | -21.844 | 25 | 20.149 | 0 | 0
|  |  | 10.05 | 3.8706 | 1 | -21.834 | 25 | 20.149 | 0 | 0
|  |  | 10.1 | 3.8162 | 1 | -21.824 | 25 | 20.149 | 0 | 0
|  |  | 10.15 | 3.7617 | 1 | -21.814 | 25 | 20.149 | 0 | 0
|  |  | 10.2 | 3.7073 | 1 | -21.804 | 25 | 20.149 | 0 | 0
|  |  | 10.25 | 3.6529 | 1 | -21.794 | 25 | 20.149 | 0 | 0
|  |  | 10.3 | 3.5985 | 1 | -21.784 | 25 | 20.149 | 0 | 0
|  |  | 10.35 | 3.5441 | 1 | -21.774 | 25 | 20.149 | 0 | 0
|  |  | 10.4 | 3.4897 | 1 | -21.764 | 25 | 20.149 | 0 | 0
|  |  | 10.45 | 3.4353 | 1 | -21.754 | 25 | 20.149 | 0 | 0
|  |  | 10.5 | 3.381 | 1 | -21.744 | 25 | 20.149 | 0 | 0
|  |  | 10.55 | 3.3266 | 1 | -21.734 | 25 | 20.149 | 0 | 0
|  |  | 10.6 | 3.2722 | 1 | -21.724 | 25 | 20.149 | 0 | 0
|  |  | 10.65 | 3.2179 | 1 | -21.714 | 25 | 20.149 | 0 | 0
|  |  | 10.7 | 3.1759 | 1 | -21.704 | 25 | 20.149 | 0 | 0
|  |  | 10.75 | 3.1339 | 1 | -21.694 | 25 | 20.149 | 0 | 0
|  |  | 10.8 | 3.092 | 1 | -21.684 | 25 | 20.149 | 0 | 0
|  |  | 10.85 | 3.05 | 1 | -21.674 | 25 | 20.149 | 0 | 0
|  |  | 10.9 | 3.0081 | 1 | -21.664 | 25 | 20.149 | 0 | 0
|  |  | 10.95 | 2.9662 | 1 | -21.654 | 25 | 20.149 | 0 | 0
|  |  | 11 | 2.9242 | 1 | -21.644 | 25 | 20.149 | 0 | 0
|  |  | 11.05 | 2.8823 | 1 | -21.634 | 25 | 20.149 | 0 | 0
|  |  | 11.1 | 2.8404 | 1 | -21.624 | 25 | 20.149 | 0 | 0
|  |  | 11.15 | 2.7985 | 1 | -21.614 | 25 | 20.149 | 0 | 0
|  |  | 11.2 | 2.7566 | 1 | -21.604 | 25 | 20.149 | 0 | 0
|  |  | 11.25 | 2.7146 | 1 | -21.594 | 25 | 20.149 | 0 | 0
|  |  | 11.3 | 2.6728 | 1 | -21.584 | 25 | 20.149 | 0 | 0
|  |  | 11.35 | 2.6309 | 1 | -21.574 | 25 | 20.149 | 0 | 0
|  |  | 11.4 | 2.589 | 1 | -21.564 | 25 | 20.149 | 0 | 0
|  |  | 11.45 | 2.5471 | 1 | -21.554 | 25 | 20.149 | 0 | 0
|  |  | 11.5 | 2.5179 | 1 | -21.544 | 25 | 20.149 | 0 | 0
|  |  | 11.55 | 2.4887 | 1 | -21.534 | 25 | 20.149 | 0 | 0
|  |  | 11.6 | 2.4595 | 1 | -21.524 | 25 | 20.149 | 0 | 0
|  |  | 11.65 | 2.4303 | 1 | -21.514 | 25 | 20.149 | 0 | 0
|  |  | 11.7 | 2.4011 | 1 | -21.504 | 25 | 20.149 | 0 | 0
|  |  | 11.75 | 2.3719 | 1 | -21.494 | 25 | 20.149 | 0 | 0
|  |  | 11.8 | 2.3427 | 1 | -21.484 | 25 | 20.149 | 0 | 0
|  |  | 11.85 | 2.3135 | 1 | -21.474 | 25 | 20.149 | 0 | 0
|  |  | 11.9 | 2.2843 | 1 | -21.464 | 25 | 20.149 | 0 | 0
|  |  | 11.95 | 2.2551 | 1 | -21.454 | 25 | 20.149 | 0 | 0
|  |  | 12 | 2.226 | 1 | -21.444 | 25 | 20.149 | 0 | 0
|  |  | 12.05 | 2.1968 | 1 | -21.434 | 25 | 20.149 | 0 | 0
|  |  | 12.1 | 2.1676 | 1 | -21.424 | 25 | 20.149 | 0 | 0
|  |  | 12.15 | 2.1385 | 1 | -21.414 | 25 | 20.149 | 0 | 0
|  |  | 12.2 | 2.1093 | 1 | -21.404 | 25 | 20.149 | 0 | 0
|  |  | 12.25 | 2.0802 | 1 | -21.394 | 25 | 20.149 | 0 | 0
|  |  | 12.3 | 2.0514 | 1 | -21.384 | 25 | 20.149 | 0 | 0
|  |  | 12.35 | 2.0227 | 1 | -21.374 | 25 | 20.149 | 0 | 0
|  |  | 12.4 | 1.994 | 1 | -21.364 | 25 | 20.149 | 0 | 0
|  |  | 12.45 | 1.9653 | 1 | -21.354 | 25 | 20.149 | 0 | 0
|  |  | 12.5 | 1.9366 | 1 | -21.344 | 25 | 20.149 | 0 | 0
|  |  | 12.55 | 1.9078 | 1 | -21.334 | 25 | 20.149 | 0 | 0
|  |  | 12.6 | 1.8791 | 1 | -21.324 | 25 | 20.149 | 0 | 0
|  |  | 12.65 | 1.8504 | 1 | -21.314 | 25 | 20.149 | 0 | 0
|  |  | 12.7 | 1.8217 | 1 | -21.304 | 25 | 20.149 | 0 | 0
|  |  | 12.75 | 1.7931 | 1 | -21.294 | 25 | 20.149 | 0 | 0
|  |  | 12.8 | 1.7644 | 1 | -21.284 | 25 | 20.149 | 0 | 0
|  |  | 12.85 | 1.7357 | 1 | -21.274 | 25 | 20.149 | 0 | 0
|  |  | 12.9 | 1.707 | 1 | -21.264 | 25 | 20.15 | 0 | 0
|  |  | 12.95 | 1.6784 | 1 | -21.254 | 25 | 20.15 | 0 | 0
|  |  | 13 | 1.6497 | 1 | -21.244 | 25 | 20.15 | 0 | 0
|  |  | 13.05 | 1.621 | 1 | -21.234 | 25 | 20.15 | 0 | 0
|  |  | 13.1 | 1.596 | 1 | -21.224 | 25 | 20.15 | 0 | 0
|  |  | 13.15 | 1.571 | 1 | -21.214 | 25 | 20.15 | 0 | 0
|  |  | 13.2 | 1.546 | 1 | -21.204 | 25 | 20.15 | 0 | 0
|  |  | 13.25 | 1.521 | 1 | -21.194 | 25 | 20.15 | 0 | 0
|  |  | 13.3 | 1.496 | 1 | -21.184 | 25 | 20.15 | 0 | 0
|  |  | 13.35 | 1.471 | 1 | -21.174 | 25 | 20.15 | 0 | 0
|  |  | 13.4 | 1.446 | 1 | -21.164 | 25 | 20.15 | 0 | 0
|  |  | 13.45 | 1.421 | 1 | -21.154 | 25 | 20.15 | 0 | 0
|  |  | 13.5 | 1.396 | 1 | -21.144 | 25 | 20.15 | 0 | 0
|  |  | 13.55 | 1.371 | 1 | -21.134 | 25 | 20.15 | 0 | 0
|  |  | 13.6 | 1.346 | 1 | -21.124 | 25 | 20.15 | 0 | 0
|  |  | 13.65 | 1.321 | 1 | -21.114 | 25 | 20.15 | 0 | 0
|  |  | 13.7 | 1.2961 | 1 | -21.104 | 25 | 20.15 | 0 | 0
|  |  | 13.75 | 1.2711 | 1 | -21.094 | 25 | 20.15 | 0 | 0
|  |  | 13.8 | 1.2461 | 1 | -21.084 | 25 | 20.15 | 0 | 0
|  |  | 13.85 | 1.2212 | 1 | -21.074 | 25 | 20.15 | 0 | 0
|  |  | 13.9 | 1.2083 | 1 | -21.064 | 25 | 20.15 | 0 | 0
|  |  | 13.95 | 1.1955 | 1 | -21.054 | 25 | 20.15 | 0 | 0
|  |  | 14 | 1.1827 | 1 | -21.044 | 25 | 20.15 | 0 | 0
|  |  | 14.05 | 1.1699 | 1 | -21.034 | 25 | 20.15 | 0 | 0
|  |  | 14.1 | 1.157 | 1 | -21.024 | 25 | 20.15 | 0 | 0
|  |  | 14.15 | 1.1442 | 1 | -21.014 | 25 | 20.15 | 0 | 0
|  |  | 14.2 | 1.1314 | 1 | -21.004 | 25 | 20.15 | 0 | 0
|  |  | 14.25 | 1.1186 | 1 | -20.994 | 25 | 20.15 | 0 | 0
|  |  | 14.3 | 1.1058 | 1 | -20.984 | 25 | 20.15 | 0 | 0
|  |  | 14.35 | 1.093 | 1 | -20.974 | 25 | 20.15 | 0 | 0
|  |  | 14.4 | 1.0802 | 1 | -20.964 | 25 | 20.15 | 0 | 0
|  |  | 14.45 | 1.0674 | 1 | -20.954 | 25 | 20.15 | 0 | 0
|  |  | 14.5 | 1.0546 | 1 | -20.944 | 25 | 20.15 | 0 | 0
|  |  | 14.55 | 1.0418 | 1 | -20.934 | 25 | 20.15 | 0 | 0
|  |  | 14.6 | 1.029 | 1 | -20.924 | 25 | 20.15 | 0 | 0
|  |  | 14.65 | 1.0162 | 1 | -20.914 | 25 | 20.15 | 0 | 0
|  |  | 14.7 | 1.0058 | 1 | -20.904 | 25 | 20.15 | 0 | 0
|  |  | 14.75 | 0.9954 | 1 | -20.894 | 25 | 20.15 | 0 | 0
|  |  | 14.8 | 0.98503 | 1 | -20.884 | 25 | 20.15 | 0 | 0
|  |  | 14.85 | 0.97465 | 1 | -20.874 | 25 | 20.15 | 0 | 0
|  |  | 14.9 | 0.96429 | 1 | -20.864 | 25 | 20.15 | 0 | 0
|  |  | 14.95 | 0.95392 | 1 | -20.854 | 25 | 20.15 | 0 | 0
|  |  | 15 | 0.94355 | 1 | -20.844 | 25 | 20.15 | 0 | 0
|  |  | 15.05 | 0.93319 | 1 | -20.834 | 25 | 20.15 | 0 | 0
|  |  | 15.1 | 0.92283 | 1 | -20.824 | 25 | 20.15 | 0 | 0
|  |  | 15.15 | 0.91247 | 1 | -20.814 | 25 | 20.15 | 0 | 0
|  |  | 15.2 | 0.90211 | 1 | -20.804 | 25 | 20.15 | 0 | 0
|  |  | 15.25 | 0.89175 | 1 | -20.794 | 25 | 20.15 | 0 | 0
|  |  | 15.3 | 0.8814 | 1 | -20.784 | 25 | 20.15 | 0 | 0
|  |  | 15.35 | 0.87105 | 1 | -20.774 | 25 | 20.15 | 0 | 0
|  |  | 15.4 | 0.8607 | 1 | -20.764 | 25 | 20.15 | 0 | 0
|  |  | 15.45 | 0.85035 | 1 | -20.754 | 25 | 20.15 | 0 | 0
|  |  | 15.5 | 0.84292 | 1 | -20.744 | 25 | 20.15 | 0 | 0
|  |  | 15.55 | 0.83549 | 1 | -20.734 | 25 | 20.15 | 0 | 0
|  |  | 15.6 | 0.82806 | 1 | -20.724 | 25 | 20.15 | 0 | 0
|  |  | 15.65 | 0.82063 | 1 | -20.714 | 25 | 20.15 | 0 | 0
|  |  | 15.7 | 0.8132 | 1 | -20.704 | 25 | 20.15 | 0 | 0
|  |  | 15.75 | 0.80577 | 1 | -20.694 | 25 | 20.15 | 0 | 0
|  |  | 15.8 | 0.79835 | 1 | -20.684 | 25 | 20.15 | 0 | 0
|  |  | 15.85 | 0.79092 | 1 | -20.674 | 25 | 20.15 | 0 | 0
|  |  | 15.9 | 0.7835 | 1 | -20.664 | 25 | 20.15 | 0 | 0
|  |  | 15.95 | 0.77609 | 1 | -20.654 | 25 | 20.15 | 0 | 0
|  |  | 16 | 0.76867 | 1 | -20.644 | 25 | 20.15 | 0 | 0
|  |  | 16.05 | 0.76125 | 1 | -20.634 | 25 | 20.15 | 0 | 0
|  |  | 16.1 | 0.75384 | 1 | -20.624 | 25 | 20.151 | 0 | 0
|  |  | 16.15 | 0.74642 | 1 | -20.614 | 25 | 20.151 | 0 | 0
|  |  | 16.2 | 0.73901 | 1 | -20.604 | 25 | 20.151 | 0 | 0
|  |  | 16.25 | 0.7316 | 1 | -20.594 | 25 | 20.151 | 0 | 0
|  |  | 16.3 | 0.72584 | 1 | -20.584 | 25 | 20.151 | 0 | 0
|  |  | 16.35 | 0.72009 | 1 | -20.574 | 25 | 20.151 | 0 | 0
|  |  | 16.4 | 0.71433 | 1 | -20.564 | 25 | 20.151 | 0 | 0
|  |  | 16.45 | 0.70858 | 1 | -20.554 | 25 | 20.151 | 0 | 0
|  |  | 16.5 | 0.70283 | 1 | -20.544 | 25 | 20.151 | 0 | 0
|  |  | 16.55 | 0.69708 | 1 | -20.534 | 25 | 20.151 | 0 | 0
|  |  | 16.6 | 0.69133 | 1 | -20.524 | 25 | 20.151 | 0 | 0
|  |  | 16.65 | 0.68558 | 1 | -20.514 | 25 | 20.151 | 0 | 0
|  |  | 16.7 | 0.67984 | 1 | -20.504 | 25 | 20.151 | 0 | 0
|  |  | 16.75 | 0.67409 | 1 | -20.494 | 25 | 20.151 | 0 | 0
|  |  | 16.8 | 0.66834 | 1 | -20.484 | 25 | 20.151 | 0 | 0
|  |  | 16.85 | 0.6626 | 1 | -20.474 | 25 | 20.151 | 0 | 0
|  |  | 16.9 | 0.65686 | 1 | -20.464 | 25 | 20.151 | 0 | 0
|  |  | 16.95 | 0.65112 | 1 | -20.454 | 25 | 20.151 | 0 | 0
|  |  | 17 | 0.64538 | 1 | -20.444 | 25 | 20.151 | 0 | 0
|  |  | 17.05 | 0.63964 | 1 | -20.434 | 25 | 20.151 | 0 | 0
|  |  | 17.1 | 0.63476 | 1 | -20.424 | 25 | 20.151 | 0 | 0
|  |  | 17.15 | 0.62989 | 1 | -20.414 | 25 | 20.151 | 0 | 0
|  |  | 17.2 | 0.62501 | 1 | -20.404 | 25 | 20.151 | 0 | 0
|  |  | 17.25 | 0.62013 | 1 | -20.394 | 25 | 20.151 | 0 | 0
|  |  | 17.3 | 0.61525 | 1 | -20.384 | 25 | 20.151 | 0 | 0
|  |  | 17.35 | 0.61038 | 1 | -20.374 | 25 | 20.151 | 0 | 0
|  |  | 17.4 | 0.6055 | 1 | -20.364 | 25 | 20.151 | 0 | 0
|  |  | 17.45 | 0.60063 | 1 | -20.354 | 25 | 20.151 | 0 | 0
|  |  | 17.5 | 0.59576 | 1 | -20.344 | 25 | 20.151 | 0 | 0
|  |  | 17.55 | 0.59089 | 1 | -20.334 | 25 | 20.151 | 0 | 0
|  |  | 17.6 | 0.58602 | 1 | -20.324 | 25 | 20.151 | 0 | 0
|  |  | 17.65 | 0.58115 | 1 | -20.314 | 25 | 20.151 | 0 | 0
|  |  | 17.7 | 0.57628 | 1 | -20.304 | 25 | 20.151 | 0 | 0
|  |  | 17.75 | 0.57142 | 1 | -20.294 | 25 | 20.151 | 0 | 0
|  |  | 17.8 | 0.56655 | 1 | -20.284 | 25 | 20.151 | 0 | 0
|  |  | 17.85 | 0.56168 | 1 | -20.274 | 25 | 20.151 | 0 | 0
|  |  | 17.9 | 0.49687 | 1 | -20.264 | 25 | 20.151 | 0 | 0
|  |  | 17.95 | 0.43207 | 1 | -20.254 | 25 | 20.151 | 0 | 0
|  |  | 18 | 0.36726 | 1 | -20.244 | 25 | 20.151 | 0 | 0
|  |  | 18.05 | 0.30246 | 1 | -20.234 | 25 | 20.151 | 0 | 0
|  |  | 18.1 | 0.2377 | 1 | -20.224 | 25 | 20.151 | 0 | 0
|  |  | 18.15 | 0.17295 | 1 | -20.214 | 25 | 20.151 | 0 | 0
|  |  | 18.2 | 0.1082 | 1 | -20.204 | 25 | 20.151 | 0 | 0
|  |  | 18.25 | 0.04345 | 1 | -20.194 | 25 | 20.151 | 0 | 0
|  |  | 18.3 | 0.042138 | 1 | -20.184 | 25 | 20.151 | 0 | 0
|  |  | 18.35 | 0.040826 | 1 | -20.174 | 25 | 20.151 | 0 | 0
|  |  | 18.4 | 0.039514 | 1 | -20.164 | 25 | 20.151 | 0 | 0
|  |  | 18.45 | 0.038201 | 1 | -20.154 | 25 | 20.151 | 0 | 0
|  |  | 18.5 | 0.037291 | 1 | -20.144 | 25 | 20.151 | 0 | 0
|  |  | 18.55 | 0.03638 | 1 | -20.134 | 25 | 20.151 | 0 | 0
|  |  | 18.6 | 0.035469 | 1 | -20.124 | 25 | 20.151 | 0 | 0
|  |  | 18.65 | 0.034558 | 1 | -20.114 | 25 | 20.151 | 0 | 0
|  |  | 18.7 | 0.032985 | 1 | -20.104 | 25 | 20.151 | 0 | 0
|  |  | 18.75 | 0.031413 | 1 | -20.094 | 25 | 20.151 | 0 | 0
|  |  | 18.8 | 0.02984 | 1 | -20.084 | 25 | 20.151 | 0 | 0
|  |  | 18.85 | 0.028268 | 1 | -20.074 | 25 | 20.151 | 0 | 0
|  |  | 18.9 | 0.02771 | 1 | -20.064 | 25 | 20.151 | 0 | 0
|  |  | 18.95 | 0.027153 | 1 | -20.054 | 25 | 20.151 | 0 | 0
|  |  | 19 | 0.026596 | 1 | -20.044 | 25 | 20.151 | 0 | 0
|  |  | 19.05 | 0.026038 | 1 | -20.034 | 25 | 20.151 | 0 | 0
|  |  | 19.1 | 0.025046 | 1 | -20.024 | 25 | 20.151 | 0 | 0
|  |  | 19.15 | 0.024053 | 1 | -20.014 | 25 | 20.151 | 0 | 0
|  |  | 19.2 | 0.023061 | 1 | -20.004 | 25 | 20.151 | 0 | 0
|  |  | 19.25 | 0.022068 | 1 | -19.994 | 25 | 20.152 | 0 | 0
|  |  | 19.3 | 0.021332 | 1 | -19.984 | 25 | 20.152 | 0 | 0
|  |  | 19.35 | 0.020595 | 1 | -19.974 | 25 | 20.152 | 0 | 0
|  |  | 19.4 | 0.019858 | 1 | -19.964 | 25 | 20.152 | 0 | 0
|  |  | 19.45 | 0.019121 | 1 | -19.954 | 25 | 20.152 | 0 | 0
|  |  | 19.5 | 0.018691 | 1 | -19.944 | 25 | 20.152 | 0 | 0
|  |  | 19.55 | 0.018262 | 1 | -19.934 | 25 | 20.152 | 0 | 0
|  |  | 19.6 | 0.017832 | 1 | -19.924 | 25 | 20.152 | 0 | 0
|  |  | 19.65 | 0.017402 | 1 | -19.914 | 25 | 20.152 | 0 | 0
|  |  | 19.7 | 0.016603 | 1 | -19.904 | 25 | 20.152 | 0 | 0
|  |  | 19.75 | 0.015805 | 1 | -19.894 | 25 | 20.152 | 0 | 0
|  |  | 19.8 | 0.015006 | 1 | -19.884 | 25 | 20.152 | 0 | 0
|  |  | 19.85 | 0.014207 | 1 | -19.874 | 25 | 20.152 | 0 | 0
|  |  | 19.9 | 0.013919 | 1 | -19.864 | 25 | 20.152 | 0 | 0
|  |  | 19.95 | 0.01363 | 1 | -19.854 | 25 | 20.152 | 0 | 0
|  |  | 20 | 0.013342 | 1 | -19.844 | 25 | 20.152 | 0 | 0
|  |  | 20.05 | 0.013054 | 1 | -19.834 | 25 | 20.152 | 0 | 0
|  |  | 20.1 | 0.012592 | 1 | -19.824 | 25 | 20.152 | 0 | 0
|  |  | 20.15 | 0.01213 | 1 | -19.814 | 25 | 20.152 | 0 | 0
|  |  | 20.2 | 0.011667 | 1 | -19.804 | 25 | 20.152 | 0 | 0
|  |  | 20.25 | 0.011205 | 1 | -19.794 | 25 | 20.152 | 0 | 0
|  |  | 20.3 | 0.01082 | 1 | -19.784 | 25 | 20.152 | 0 | 0
|  |  | 20.35 | 0.010434 | 1 | -19.774 | 25 | 20.152 | 0 | 0
|  |  | 20.4 | 0.010049 | 1 | -19.764 | 25 | 20.152 | 0 | 0
|  |  | 20.45 | 0.0096633 | 1 | -19.755 | 25 | 20.152 | 0 | 0
|  |  | 20.5 | 0.0094664 | 1 | -19.745 | 25 | 20.152 | 0 | 0
|  |  | 20.55 | 0.0092694 | 1 | -19.735 | 25 | 20.152 | 0 | 0
|  |  | 20.6 | 0.0090725 | 1 | -19.725 | 25 | 20.152 | 0 | 0
|  |  | 20.65 | 0.0088755 | 1 | -19.715 | 25 | 20.152 | 0 | 0
|  |  | 20.7 | 0.0084955 | 1 | -19.705 | 25 | 20.152 | 0 | 0
|  |  | 20.75 | 0.0081156 | 1 | -19.695 | 25 | 20.152 | 0 | 0
|  |  | 20.8 | 0.0077356 | 1 | -19.685 | 25 | 20.152 | 0 | 0
|  |  | 20.85 | 0.0073557 | 1 | -19.675 | 25 | 20.152 | 0 | 0
|  |  | 20.9 | 0.0072103 | 1 | -19.665 | 25 | 20.152 | 0 | 0
|  |  | 20.95 | 0.0070649 | 1 | -19.655 | 25 | 20.152 | 0 | 0
|  |  | 21 | 0.0069196 | 1 | -19.645 | 25 | 20.152 | 0 | 0
|  |  | 21.05 | 0.0067742 | 1 | -19.635 | 25 | 20.152 | 0 | 0
|  |  | 21.1 | 0.0065734 | 1 | -19.625 | 25 | 20.152 | 0 | 0
|  |  | 21.15 | 0.0063726 | 1 | -19.615 | 25 | 20.152 | 0 | 0
|  |  | 21.2 | 0.0061718 | 1 | -19.605 | 25 | 20.152 | 0 | 0
|  |  | 21.25 | 0.005971 | 1 | -19.595 | 25 | 20.152 | 0 | 0
|  |  | 21.3 | 0.0057773 | 1 | -19.585 | 25 | 20.152 | 0 | 0
|  |  | 21.35 | 0.0055836 | 1 | -19.575 | 25 | 20.152 | 0 | 0
|  |  | 21.4 | 0.0053899 | 1 | -19.565 | 25 | 20.152 | 0 | 0
|  |  | 21.45 | 0.0051962 | 1 | -19.555 | 25 | 20.152 | 0 | 0
|  |  | 21.5 | 0.0051089 | 1 | -19.545 | 25 | 20.152 | 0 | 0
|  |  | 21.55 | 0.0050215 | 1 | -19.535 | 25 | 20.152 | 0 | 0
|  |  | 21.6 | 0.0049342 | 1 | -19.525 | 25 | 20.152 | 0 | 0
|  |  | 21.65 | 0.0048469 | 1 | -19.515 | 25 | 20.152 | 0 | 0
|  |  | 21.7 | 0.0046739 | 1 | -19.505 | 25 | 20.152 | 0 | 0
|  |  | 21.75 | 0.0045009 | 1 | -19.495 | 25 | 20.152 | 0 | 0
|  |  | 21.8 | 0.004328 | 1 | -19.485 | 25 | 20.152 | 0 | 0
|  |  | 21.85 | 0.004155 | 1 | -19.475 | 25 | 20.152 | 0 | 0
|  |  | 21.9 | 0.0040812 | 1 | -19.465 | 25 | 20.152 | 0 | 0
|  |  | 21.95 | 0.0040073 | 1 | -19.455 | 25 | 20.152 | 0 | 0
|  |  | 22 | 0.0039335 | 1 | -19.445 | 25 | 20.152 | 0 | 0
|  |  | 22.05 | 0.0038597 | 1 | -19.435 | 25 | 20.152 | 0 | 0
|  |  | 22.1 | 0.0037719 | 1 | -19.425 | 25 | 20.152 | 0 | 0
|  |  | 22.15 | 0.0036842 | 1 | -19.415 | 25 | 20.152 | 0 | 0
|  |  | 22.2 | 0.0035965 | 1 | -19.405 | 25 | 20.152 | 0 | 0
|  |  | 22.25 | 0.0035088 | 1 | -19.395 | 25 | 20.152 | 0 | 0
|  |  | 22.3 | 0.0034105 | 1 | -19.385 | 25 | 20.152 | 0 | 0
|  |  | 22.35 | 0.0033122 | 1 | -19.375 | 25 | 20.152 | 0 | 0
|  |  | 22.4 | 0.0032139 | 1 | -19.365 | 25 | 20.152 | 0 | 0
|  |  | 22.45 | 0.0031157 | 1 | -19.355 | 25 | 20.153 | 0 | 0
|  |  | 22.5 | 0.0030746 | 1 | -19.345 | 25 | 20.153 | 0 | 0
|  |  | 22.55 | 0.0030335 | 1 | -19.335 | 25 | 20.153 | 0 | 0
|  |  | 22.6 | 0.0029924 | 1 | -19.325 | 25 | 20.153 | 0 | 0
|  |  | 22.65 | 0.0029514 | 1 | -19.315 | 25 | 20.153 | 0 | 0
|  |  | 22.7 | 0.0028719 | 1 | -19.305 | 25 | 20.153 | 0 | 0
|  |  | 22.75 | 0.0027925 | 1 | -19.295 | 25 | 20.153 | 0 | 0
|  |  | 22.8 | 0.0027131 | 1 | -19.285 | 25 | 20.153 | 0 | 0
|  |  | 22.85 | 0.0026336 | 1 | -19.275 | 25 | 20.153 | 0 | 0
|  |  | 22.9 | 0.0025951 | 1 | -19.265 | 25 | 20.153 | 0 | 0
|  |  | 22.95 | 0.0025565 | 1 | -19.255 | 25 | 20.153 | 0 | 0
|  |  | 23 | 0.0025179 | 1 | -19.245 | 25 | 20.153 | 0 | 0
|  |  | 23.05 | 0.0024793 | 1 | -19.235 | 25 | 20.153 | 0 | 0
|  |  | 23.1 | 0.0024406 | 1 | -19.225 | 25 | 20.153 | 0 | 0
|  |  | 23.15 | 0.0024019 | 1 | -19.215 | 25 | 20.153 | 0 | 0
|  |  | 23.2 | 0.0023632 | 1 | -19.205 | 25 | 20.153 | 0 | 0
|  |  | 23.25 | 0.0023245 | 1 | -19.195 | 25 | 20.153 | 0 | 0
|  |  | 23.3 | 0.0022749 | 1 | -19.185 | 25 | 20.153 | 0 | 0
|  |  | 23.35 | 0.0022254 | 1 | -19.175 | 25 | 20.153 | 0 | 0
|  |  | 23.4 | 0.0021759 | 1 | -19.165 | 25 | 20.153 | 0 | 0
|  |  | 23.45 | 0.0021263 | 1 | -19.155 | 25 | 20.153 | 0 | 0
|  |  | 23.5 | 0.0021064 | 1 | -19.145 | 25 | 20.153 | 0 | 0
|  |  | 23.55 | 0.0020865 | 1 | -19.135 | 25 | 20.153 | 0 | 0
|  |  | 23.6 | 0.0020666 | 1 | -19.125 | 25 | 20.153 | 0 | 0
|  |  | 23.65 | 0.0020467 | 1 | -19.115 | 25 | 20.153 | 0 | 0
|  |  | 23.7 | 0.0020096 | 1 | -19.105 | 25 | 20.153 | 0 | 0
|  |  | 23.75 | 0.0019726 | 1 | -19.095 | 25 | 20.153 | 0 | 0
|  |  | 23.8 | 0.0019355 | 1 | -19.085 | 25 | 20.153 | 0 | 0
|  |  | 23.85 | 0.0018984 | 1 | -19.075 | 25 | 20.153 | 0 | 0
|  |  | 23.9 | 0.0018772 | 1 | -19.065 | 25 | 20.153 | 0 | 0
|  |  | 23.95 | 0.001856 | 1 | -19.055 | 25 | 20.153 | 0 | 0
|  |  | 24 | 0.0018348 | 1 | -19.045 | 25 | 20.153 | 0 | 0
|  |  | 24.05 | 0.0018137 | 1 | -19.035 | 25 | 20.153 | 0 | 0
|  |  | 24.1 | 0.0017958 | 1 | -19.025 | 25 | 20.153 | 0 | 0
|  |  | 24.15 | 0.001778 | 1 | -19.015 | 25 | 20.153 | 0 | 0
|  |  | 24.2 | 0.0017602 | 1 | -19.005 | 25 | 20.153 | 0 | 0
|  |  | 24.25 | 0.0017424 | 1 | -18.995 | 25 | 20.153 | 0 | 0
|  |  | 24.3 | 0.0017175 | 1 | -18.985 | 25 | 20.153 | 0 | 0
|  |  | 24.35 | 0.0016926 | 1 | -18.975 | 25 | 20.153 | 0 | 0
|  |  | 24.4 | 0.0016677 | 1 | -18.965 | 25 | 20.153 | 0 | 0
|  |  | 24.45 | 0.0016428 | 1 | -18.955 | 25 | 20.153 | 0 | 0
|  |  | 24.5 | 0.0016327 | 1 | -18.945 | 25 | 20.153 | 0 | 0
|  |  | 24.55 | 0.0016226 | 1 | -18.935 | 25 | 20.153 | 0 | 0
|  |  | 24.6 | 0.0016125 | 1 | -18.925 | 25 | 20.153 | 0 | 0
|  |  | 24.65 | 0.0016024 | 1 | -18.915 | 25 | 20.153 | 0 | 0
|  |  | 24.7 | 0.0015844 | 1 | -18.905 | 25 | 20.153 | 0 | 0
|  |  | 24.75 | 0.0015663 | 1 | -18.895 | 25 | 20.153 | 0 | 0
|  |  | 24.8 | 0.0015483 | 1 | -18.885 | 25 | 20.153 | 0 | 0
|  |  | 24.85 | 0.0015302 | 1 | -18.875 | 25 | 20.153 | 0 | 0
|  |  | 24.9 | 0.001518 | 1 | -18.865 | 25 | 20.153 | 0 | 0
|  |  | 24.95 | 0.0015058 | 1 | -18.855 | 25 | 20.153 | 0 | 0
|  |  | 25 | 0.0014935 | 1 | -18.845 | 25 | 20.153 | 0 | 0
|  |  | 25.05 | 0.0014813 | 1 | -18.835 | 25 | 20.153 | 0 | 0
|  |  | 25.1 | 0.0014719 | 1 | -18.825 | 25 | 20.153 | 0 | 0
|  |  | 25.15 | 0.0014625 | 1 | -18.815 | 25 | 20.153 | 0 | 0
|  |  | 25.2 | 0.0014531 | 1 | -18.805 | 25 | 20.153 | 0 | 0
|  |  | 25.25 | 0.0014437 | 1 | -18.795 | 25 | 20.153 | 0 | 0
|  |  | 25.3 | 0.0014282 | 1 | -18.785 | 25 | 20.153 | 0 | 0
|  |  | 25.35 | 0.0014128 | 1 | -18.775 | 25 | 20.153 | 0 | 0
|  |  | 25.4 | 0.0013974 | 1 | -18.765 | 25 | 20.153 | 0 | 0
|  |  | 25.45 | 0.0013819 | 1 | -18.755 | 25 | 20.153 | 0 | 0
|  |  | 25.5 | 0.001374 | 1 | -18.745 | 25 | 20.153 | 0 | 0
|  |  | 25.55 | 0.001366 | 1 | -18.735 | 25 | 20.153 | 0 | 0
|  |  | 25.6 | 0.001358 | 1 | -18.725 | 25 | 20.153 | 0 | 0
|  |  | 25.65 | 0.0013501 | 1 | -18.715 | 25 | 20.154 | 0 | 0
|  |  | 25.7 | 0.0013386 | 1 | -18.705 | 25 | 20.154 | 0 | 0
|  |  | 25.75 | 0.0013272 | 1 | -18.695 | 25 | 20.154 | 0 | 0
|  |  | 25.8 | 0.0013158 | 1 | -18.685 | 25 | 20.154 | 0 | 0
|  |  | 25.85 | 0.0013044 | 1 | -18.675 | 25 | 20.154 | 0 | 0
|  |  | 25.9 | 0.0012949 | 1 | -18.665 | 25 | 20.154 | 0 | 0
|  |  | 25.95 | 0.0012855 | 1 | -18.655 | 25 | 20.154 | 0 | 0
|  |  | 26 | 0.001276 | 1 | -18.645 | 25 | 20.154 | 0 | 0
|  |  | 26.05 | 0.0012666 | 1 | -18.635 | 25 | 20.154 | 0 | 0
|  |  | 26.1 | 0.001259 | 1 | -18.625 | 25 | 20.154 | 0 | 0
|  |  | 26.15 | 0.0012515 | 1 | -18.615 | 25 | 20.154 | 0 | 0
|  |  | 26.2 | 0.001244 | 1 | -18.605 | 25 | 20.154 | 0 | 0
|  |  | 26.25 | 0.0012365 | 1 | -18.595 | 25 | 20.154 | 0 | 0
|  |  | 26.3 | 0.0012153 | 1 | -18.585 | 25 | 20.154 | 0 | 0
|  |  | 26.35 | 0.0011942 | 1 | -18.575 | 25 | 20.154 | 0 | 0
|  |  | 26.4 | 0.0011731 | 1 | -18.565 | 25 | 20.154 | 0 | 0
|  |  | 26.45 | 0.001152 | 1 | -18.555 | 25 | 20.154 | 0 | 0
|  |  | 26.5 | 0.001124 | 1 | -18.545 | 25 | 20.154 | 0 | 0
|  |  | 26.55 | 0.0010961 | 1 | -18.535 | 25 | 20.154 | 0 | 0
|  |  | 26.6 | 0.0010681 | 1 | -18.525 | 25 | 20.154 | 0 | 0
|  |  | 26.65 | 0.0010402 | 1 | -18.515 | 25 | 20.154 | 0 | 0
|  |  | 26.7 | 0.0010104 | 1 | -18.505 | 25 | 20.154 | 0 | 0
|  |  | 26.75 | 0.0009807 | 1 | -18.495 | 25 | 20.154 | 0 | 0
|  |  | 26.8 | 0.00095098 | 1 | -18.485 | 25 | 20.154 | 0 | 0
|  |  | 26.85 | 0.00092125 | 1 | -18.475 | 25 | 20.154 | 0 | 0
|  |  | 26.9 | 0.00089212 | 1 | -18.465 | 25 | 20.154 | 0 | 0
|  |  | 26.95 | 0.000863 | 1 | -18.455 | 25 | 20.154 | 0 | 0
|  |  | 27 | 0.00083387 | 1 | -18.445 | 25 | 20.154 | 0 | 0
|  |  | 27.05 | 0.00080474 | 1 | -18.435 | 25 | 20.154 | 0 | 0
|  |  | 27.1 | 0.00077724 | 1 | -18.425 | 25 | 20.154 | 0 | 0
|  |  | 27.15 | 0.00074975 | 1 | -18.415 | 25 | 20.154 | 0 | 0
|  |  | 27.2 | 0.00072225 | 1 | -18.405 | 25 | 20.154 | 0 | 0
|  |  | 27.25 | 0.00069475 | 1 | -18.395 | 25 | 20.154 | 0 | 0
|  |  | 27.3 | 0.00066684 | 1 | -18.385 | 25 | 20.154 | 0 | 0
|  |  | 27.35 | 0.00063892 | 1 | -18.375 | 25 | 20.154 | 0 | 0
|  |  | 27.4 | 0.000611 | 1 | -18.365 | 25 | 20.154 | 0 | 0
|  |  | 27.45 | 0.00058309 | 1 | -18.355 | 25 | 20.154 | 0 | 0
|  |  | 27.5 | 0.00056298 | 1 | -18.345 | 25 | 20.154 | 0.001 | 0.1
|  |  | 27.55 | 0.00054288 | 1 | -18.335 | 25 | 20.154 | 0.001 | 0.1
|  |  | 27.6 | 0.00052278 | 1 | -18.325 | 25 | 20.154 | 0.001 | 0.1
|  |  | 27.65 | 0.00050268 | 1 | -18.315 | 25 | 20.154 | 0.001 | 0.1
|  |  | 27.7 | 0.00048345 | 1 | -18.305 | 25 | 20.154 | 0.001 | 0.1
|  |  | 27.75 | 0.00046421 | 1 | -18.295 | 25 | 20.154 | 0.002 | 0.2
|  |  | 27.8 | 0.00044498 | 1 | -18.285 | 25 | 20.154 | 0.002 | 0.2
|  |  | 27.85 | 0.00042575 | 1 | -18.275 | 25 | 20.154 | 0.002 | 0.2
|  |  | 27.9 | 0.00040645 | 1 | -18.265 | 25 | 20.154 | 0.002 | 0.2
|  |  | 27.95 | 0.00038716 | 1 | -18.255 | 25 | 20.154 | 0.002 | 0.2
|  |  | 28 | 0.00036786 | 1 | -18.245 | 25 | 20.154 | 0.002 | 0.2
|  |  | 28.05 | 0.00034857 | 1 | -18.235 | 25 | 20.154 | 0.003 | 0.3
|  |  | 28.1 | 0.0003304 | 1 | -18.225 | 25 | 20.154 | 0.003 | 0.3
|  |  | 28.15 | 0.00031223 | 1 | -18.215 | 25 | 20.154 | 0.003 | 0.3
|  |  | 28.2 | 0.00029406 | 1 | -18.205 | 25 | 20.154 | 0.003 | 0.3
|  |  | 28.25 | 0.00027589 | 1 | -18.195 | 25 | 20.154 | 0.003 | 0.3
|  |  | 28.3 | 0.00025711 | 1 | -18.185 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.35 | 0.00023832 | 1 | -18.175 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.4 | 0.00021954 | 1 | -18.165 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.45 | 0.00020076 | 1 | -18.155 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.5 | 0.00019077 | 1 | -18.145 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.55 | 0.00018078 | 1 | -18.135 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.6 | 0.00017079 | 1 | -18.125 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.65 | 0.00016081 | 1 | -18.115 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.7 | 0.00015844 | 1 | -18.105 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.75 | 0.00015608 | 1 | -18.095 | 25 | 20.154 | 0.004 | 0.4
|  |  | 28.8 | 0.00015371 | 1 | -18.085 | 25 | 20.155 | 0.004 | 0.4
|  |  | 28.85 | 0.00015135 | 1 | -18.075 | 25 | 20.155 | 0.004 | 0.4
|  |  | 28.9 | 0.0001491 | 1 | -18.065 | 25 | 20.155 | 0.004 | 0.4
|  |  | 28.95 | 0.00014686 | 1 | -18.055 | 25 | 20.155 | 0.004 | 0.4
|  |  | 29 | 0.00014461 | 1 | -18.045 | 25 | 20.155 | 0.004 | 0.4
|  |  | 29.05 | 0.00014237 | 1 | -18.035 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.1 | 0.00014084 | 1 | -18.025 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.15 | 0.00013931 | 1 | -18.015 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.2 | 0.00013778 | 1 | -18.005 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.25 | 0.00013625 | 1 | -17.995 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.3 | 0.00013419 | 1 | -17.985 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.35 | 0.00013212 | 1 | -17.975 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.4 | 0.00013006 | 1 | -17.965 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.45 | 0.000128 | 1 | -17.955 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.5 | 0.00012645 | 1 | -17.945 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.55 | 0.0001249 | 1 | -17.935 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.6 | 0.00012335 | 1 | -17.925 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.65 | 0.00012181 | 1 | -17.915 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.7 | 0.00012027 | 1 | -17.905 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.75 | 0.00011873 | 1 | -17.895 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.8 | 0.0001172 | 1 | -17.885 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.85 | 0.00011566 | 1 | -17.875 | 25 | 20.155 | 0.005 | 0.5
|  |  | 29.9 | 0.00010119 | 1 | -17.865 | 25 | 20.155 | 0.005 | 0.5
|  | END_DATA |  |  |  | -17.855 | 25 | 20.155 | 0.005 | 0.5
| END_SCAN 1 |  |  |  |  | -17.845 | 25 | 20.155 | 0.005 | 0.5
END_SCAN_DATA | [5 empty] | -17.835 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.825 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.815 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.805 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.795 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.785 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.775 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.765 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.755 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.745 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.735 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.725 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.715 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.705 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.695 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.685 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.675 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.665 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.655 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.645 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.635 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.625 | 25 | 20.155 | 0.005 | 0.5
[6 empty] | -17.615 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.605 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.595 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.585 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.575 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.565 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.555 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.545 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.535 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.525 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.515 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.505 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.495 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.485 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.475 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.465 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.455 | 25 | 20.155 | 0.006 | 0.6
[6 empty] | -17.445 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.435 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.425 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.415 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.405 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.395 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.385 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.375 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.365 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.355 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.345 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.335 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.325 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.315 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.305 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.295 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.285 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.275 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.265 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.255 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.245 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.235 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.225 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.215 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.205 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.195 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.185 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.175 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.165 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.155 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.145 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.135 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.125 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.115 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.105 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.095 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.085 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.075 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.065 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.055 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.045 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.035 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.025 | 25 | 20.156 | 0.006 | 0.6
[6 empty] | -17.015 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -17.005 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.995 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.985 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.975 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.965 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.955 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.945 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.935 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.925 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.915 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.905 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.895 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.885 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.875 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.865 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.855 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.845 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.835 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.825 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.815 | 25 | 20.156 | 0.007 | 0.7
[6 empty] | -16.805 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.795 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.785 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.775 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.765 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.755 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.745 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.735 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.725 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.715 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.705 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.695 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.685 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.675 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.665 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.655 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.645 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.635 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.625 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.615 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.605 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.595 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.585 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.575 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.565 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.555 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.545 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.535 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.525 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.515 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.505 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.495 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.485 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.475 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.465 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.455 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.445 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.435 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.425 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.415 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.405 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.395 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.385 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.375 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.365 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.355 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.345 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.335 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.325 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.315 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.305 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.295 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.285 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.275 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.265 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.255 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.245 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.235 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.225 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.215 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.205 | 25 | 20.157 | 0.007 | 0.7
[6 empty] | -16.195 | 25 | 20.157 | 0.008 | 0.8
[6 empty] | -16.185 | 25 | 20.157 | 0.008 | 0.8
[6 empty] | -16.175 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.165 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.155 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.145 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.135 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.125 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.115 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.105 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.095 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.085 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.075 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.065 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.055 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.045 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.035 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.025 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.015 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -16.005 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.995 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.985 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.975 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.965 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.955 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.945 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.935 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.925 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.915 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.905 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.895 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.885 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.875 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.865 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.855 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.845 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.835 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.825 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.815 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.805 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.795 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.785 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.775 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.765 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.755 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.745 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.735 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.725 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.715 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.705 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.695 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.685 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.675 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.665 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.655 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.645 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.635 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.625 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.615 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.605 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.595 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.585 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.575 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.565 | 25 | 20.158 | 0.008 | 0.8
[6 empty] | -15.555 | 25 | 20.158 | 0.009 | 0.9
[6 empty] | -15.545 | 25 | 20.158 | 0.009 | 0.9
[6 empty] | -15.535 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.525 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.515 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.505 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.495 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.485 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.475 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.465 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.455 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.445 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.435 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.425 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.415 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.405 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.395 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.385 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.375 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.365 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.355 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.345 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.335 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.325 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.315 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.305 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.295 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.285 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.275 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.265 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.255 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.245 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.235 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.225 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.215 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.205 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.195 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.185 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.175 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.165 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.155 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.145 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.135 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.125 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.115 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.105 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.095 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.085 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.075 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.065 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.055 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.045 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.035 | 25 | 20.159 | 0.009 | 0.9
[6 empty] | -15.025 | 25 | 20.159 | 0.01 | 1
[6 empty] | -15.015 | 25 | 20.159 | 0.01 | 1
[6 empty] | -15.005 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.995 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.985 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.975 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.965 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.955 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.945 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.935 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.925 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.915 | 25 | 20.159 | 0.01 | 1
[6 empty] | -14.905 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.895 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.885 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.875 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.865 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.855 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.845 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.835 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.825 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.815 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.805 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.795 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.785 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.775 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.765 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.756 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.746 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.736 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.726 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.716 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.706 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.696 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.686 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.676 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.666 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.656 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.646 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.636 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.626 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.616 | 25 | 20.16 | 0.01 | 1
[6 empty] | -14.606 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.596 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.586 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.576 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.566 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.556 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.546 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.536 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.526 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.516 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.506 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.496 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.486 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.476 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.466 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.456 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.446 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.436 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.426 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.416 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.406 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.396 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.386 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.376 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.366 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.356 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.346 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.336 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.326 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.316 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.306 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.296 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.286 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.276 | 25 | 20.16 | 0.011 | 1.1
[6 empty] | -14.266 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.256 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.246 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.236 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.226 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.216 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.206 | 25 | 20.161 | 0.011 | 1.1
[6 empty] | -14.196 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.186 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.176 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.166 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.156 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.146 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.136 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.126 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.116 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.106 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.096 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.086 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.076 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.066 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.056 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.046 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.036 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.026 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.016 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -14.006 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.996 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.986 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.976 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.966 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.956 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.946 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.936 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.926 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.916 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.906 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.896 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.886 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.876 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.866 | 25 | 20.161 | 0.012 | 1.2
[6 empty] | -13.856 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.846 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.836 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.826 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.816 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.806 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.796 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.786 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.776 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.766 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.756 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.746 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.736 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.726 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.716 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.706 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.696 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.686 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.676 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.666 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.656 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.646 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.636 | 25 | 20.161 | 0.013 | 1.3
[6 empty] | -13.626 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.616 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.606 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.596 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.586 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.576 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.566 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.556 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.546 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.536 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.526 | 25 | 20.162 | 0.013 | 1.3
[6 empty] | -13.516 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.506 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.496 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.486 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.476 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.466 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.456 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.446 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.436 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.426 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.416 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.406 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.396 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.386 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.376 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.366 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.356 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.346 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.336 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.326 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.316 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.306 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.296 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.286 | 25 | 20.162 | 0.014 | 1.4
[6 empty] | -13.276 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.266 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.256 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.246 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.236 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.226 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.216 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.206 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.196 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.186 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.176 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.166 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.156 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.146 | 25 | 20.162 | 0.015 | 1.5
[6 empty] | -13.136 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.126 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.116 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.106 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.096 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.086 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.076 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.066 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.056 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.046 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.036 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.026 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.016 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -13.006 | 25 | 20.162 | 0.016 | 1.6
[6 empty] | -12.996 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.986 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.976 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.966 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.956 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.946 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.936 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.926 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.916 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.906 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.896 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.886 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.876 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.866 | 25 | 20.163 | 0.017 | 1.7
[6 empty] | -12.856 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.846 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.836 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.826 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.816 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.806 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.796 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.786 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.776 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.766 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.756 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.746 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.736 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.726 | 25 | 20.163 | 0.018 | 1.8
[6 empty] | -12.716 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.706 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.696 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.686 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.676 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.666 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.656 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.646 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.636 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.626 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.616 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.606 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.596 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.586 | 25 | 20.163 | 0.019 | 1.9
[6 empty] | -12.576 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.566 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.556 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.546 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.536 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.526 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.516 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.506 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.496 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.486 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.476 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.466 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.456 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.446 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.436 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.426 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.416 | 25 | 20.163 | 0.02 | 2
[6 empty] | -12.406 | 25 | 20.163 | 0.021 | 2.1
[6 empty] | -12.396 | 25 | 20.163 | 0.021 | 2.1
[6 empty] | -12.386 | 25 | 20.163 | 0.021 | 2.1
[6 empty] | -12.376 | 25 | 20.163 | 0.021 | 2.1
[6 empty] | -12.366 | 25 | 20.163 | 0.021 | 2.1
[6 empty] | -12.356 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.346 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.336 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.326 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.316 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.306 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.296 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.286 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.276 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.266 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.256 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.246 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.236 | 25 | 20.164 | 0.021 | 2.1
[6 empty] | -12.226 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.216 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.206 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.196 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.186 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.176 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.166 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.156 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.146 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.136 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.126 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.116 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.106 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.096 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.086 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.076 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.066 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.056 | 25 | 20.164 | 0.022 | 2.2
[6 empty] | -12.046 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -12.036 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -12.026 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -12.016 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -12.006 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.996 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.986 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.976 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.966 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.956 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.946 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.936 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.926 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.916 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.906 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.896 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.886 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.876 | 25 | 20.164 | 0.023 | 2.3
[6 empty] | -11.866 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.856 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.846 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.836 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.826 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.816 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.806 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.796 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.786 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.776 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.766 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.756 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.746 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.736 | 25 | 20.164 | 0.024 | 2.4
[6 empty] | -11.726 | 25 | 20.165 | 0.024 | 2.4
[6 empty] | -11.716 | 25 | 20.165 | 0.024 | 2.4
[6 empty] | -11.706 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.696 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.686 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.676 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.666 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.656 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.646 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.636 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.626 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.616 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.606 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.596 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.586 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.576 | 25 | 20.165 | 0.025 | 2.5
[6 empty] | -11.566 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.556 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.546 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.536 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.526 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.516 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.506 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.496 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.486 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.476 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.466 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.456 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.446 | 25 | 20.165 | 0.026 | 2.6
[6 empty] | -11.436 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.426 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.416 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.406 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.396 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.386 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.376 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.366 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.356 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.346 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.336 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.326 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.316 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.306 | 25 | 20.165 | 0.027 | 2.7
[6 empty] | -11.296 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.286 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.276 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.266 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.256 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.246 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.236 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.226 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.216 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.206 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.196 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.186 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.176 | 25 | 20.165 | 0.028 | 2.8
[6 empty] | -11.166 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.156 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.146 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.136 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.126 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.116 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.106 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.096 | 25 | 20.165 | 0.029 | 2.9
[6 empty] | -11.086 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.076 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.066 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.056 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.046 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.036 | 25 | 20.166 | 0.029 | 2.9
[6 empty] | -11.026 | 25 | 20.166 | 0.03 | 3
[6 empty] | -11.016 | 25 | 20.166 | 0.03 | 3
[6 empty] | -11.006 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.996 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.986 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.976 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.966 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.956 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.946 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.936 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.926 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.916 | 25 | 20.166 | 0.03 | 3
[6 empty] | -10.906 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.896 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.886 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.876 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.866 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.856 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.846 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.836 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.826 | 25 | 20.166 | 0.031 | 3.1
[6 empty] | -10.816 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.806 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.796 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.786 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.776 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.766 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.756 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.746 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.736 | 25 | 20.166 | 0.032 | 3.2
[6 empty] | -10.726 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.716 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.706 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.696 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.686 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.676 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.666 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.656 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.646 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.636 | 25 | 20.166 | 0.033 | 3.3
[6 empty] | -10.626 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.616 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.606 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.596 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.586 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.576 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.566 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.556 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.546 | 25 | 20.166 | 0.034 | 3.4
[6 empty] | -10.536 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.526 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.516 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.506 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.496 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.486 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.476 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.466 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.456 | 25 | 20.166 | 0.035 | 3.5
[6 empty] | -10.446 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.436 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.426 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.416 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.406 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.396 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.386 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.376 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.366 | 25 | 20.167 | 0.036 | 3.6
[6 empty] | -10.356 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.346 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.336 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.326 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.316 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.306 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.296 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.286 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.276 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.266 | 25 | 20.167 | 0.037 | 3.7
[6 empty] | -10.256 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.246 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.236 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.226 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.216 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.206 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.196 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.186 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.176 | 25 | 20.167 | 0.038 | 3.8
[6 empty] | -10.166 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.156 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.146 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.136 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.126 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.116 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.106 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.096 | 25 | 20.167 | 0.039 | 3.9
[6 empty] | -10.086 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.076 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.066 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.056 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.046 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.036 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.026 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.016 | 25 | 20.167 | 0.04 | 4
[6 empty] | -10.006 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.996 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.986 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.976 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.966 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.956 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.946 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.936 | 25 | 20.167 | 0.041 | 4.1
[6 empty] | -9.926 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.916 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.906 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.896 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.886 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.876 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.866 | 25 | 20.167 | 0.042 | 4.2
[6 empty] | -9.856 | 25 | 20.167 | 0.043 | 4.3
[6 empty] | -9.846 | 25 | 20.167 | 0.043 | 4.3
[6 empty] | -9.836 | 25 | 20.167 | 0.043 | 4.3
[6 empty] | -9.826 | 25 | 20.167 | 0.043 | 4.3
[6 empty] | -9.816 | 25 | 20.168 | 0.043 | 4.3
[6 empty] | -9.806 | 25 | 20.168 | 0.043 | 4.3
[6 empty] | -9.796 | 25 | 20.168 | 0.043 | 4.3
[6 empty] | -9.786 | 25 | 20.168 | 0.043 | 4.3
[6 empty] | -9.776 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.766 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.756 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.747 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.737 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.727 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.717 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.707 | 25 | 20.168 | 0.044 | 4.4
[6 empty] | -9.697 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.687 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.677 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.667 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.657 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.647 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.637 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.627 | 25 | 20.168 | 0.045 | 4.5
[6 empty] | -9.617 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.607 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.597 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.587 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.577 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.567 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.557 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.547 | 25 | 20.168 | 0.046 | 4.6
[6 empty] | -9.537 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.527 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.517 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.507 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.497 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.487 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.477 | 25 | 20.168 | 0.047 | 4.7
[6 empty] | -9.467 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.457 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.447 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.437 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.427 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.417 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.407 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.397 | 25 | 20.168 | 0.048 | 4.8
[6 empty] | -9.387 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.377 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.367 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.357 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.347 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.337 | 25 | 20.168 | 0.049 | 4.9
[6 empty] | -9.327 | 25 | 20.168 | 0.05 | 5
[6 empty] | -9.317 | 25 | 20.168 | 0.05 | 5
[6 empty] | -9.307 | 25 | 20.168 | 0.05 | 5
[6 empty] | -9.297 | 25 | 20.168 | 0.05 | 5
[6 empty] | -9.287 | 25 | 20.168 | 0.051 | 5.1
[6 empty] | -9.277 | 25 | 20.168 | 0.051 | 5.1
[6 empty] | -9.267 | 25 | 20.168 | 0.051 | 5.1
[6 empty] | -9.257 | 25 | 20.168 | 0.052 | 5.2
[6 empty] | -9.247 | 25 | 20.168 | 0.052 | 5.2
[6 empty] | -9.237 | 25 | 20.168 | 0.052 | 5.2
[6 empty] | -9.227 | 25 | 20.168 | 0.052 | 5.2
[6 empty] | -9.217 | 25 | 20.168 | 0.053 | 5.3
[6 empty] | -9.207 | 25 | 20.168 | 0.053 | 5.3
[6 empty] | -9.197 | 25 | 20.168 | 0.053 | 5.3
[6 empty] | -9.187 | 25 | 20.168 | 0.053 | 5.3
[6 empty] | -9.177 | 25 | 20.169 | 0.054 | 5.4
[6 empty] | -9.167 | 25 | 20.169 | 0.054 | 5.4
[6 empty] | -9.157 | 25 | 20.169 | 0.054 | 5.4
[6 empty] | -9.147 | 25 | 20.169 | 0.054 | 5.4
[6 empty] | -9.137 | 25 | 20.169 | 0.055 | 5.5
[6 empty] | -9.127 | 25 | 20.169 | 0.055 | 5.5
[6 empty] | -9.117 | 25 | 20.169 | 0.055 | 5.5
[6 empty] | -9.107 | 25 | 20.169 | 0.055 | 5.5
[6 empty] | -9.097 | 25 | 20.169 | 0.056 | 5.6
[6 empty] | -9.087 | 25 | 20.169 | 0.056 | 5.6
[6 empty] | -9.077 | 25 | 20.169 | 0.056 | 5.6
[6 empty] | -9.067 | 25 | 20.169 | 0.057 | 5.7
[6 empty] | -9.057 | 25 | 20.169 | 0.057 | 5.7
[6 empty] | -9.047 | 25 | 20.169 | 0.057 | 5.7
[6 empty] | -9.037 | 25 | 20.169 | 0.057 | 5.7
[6 empty] | -9.027 | 25 | 20.169 | 0.058 | 5.8
[6 empty] | -9.017 | 25 | 20.169 | 0.058 | 5.8
[6 empty] | -9.007 | 25 | 20.169 | 0.058 | 5.8
[6 empty] | -8.997 | 25 | 20.169 | 0.058 | 5.8
[6 empty] | -8.987 | 25 | 20.169 | 0.059 | 5.9
[6 empty] | -8.977 | 25 | 20.169 | 0.059 | 5.9
[6 empty] | -8.967 | 25 | 20.169 | 0.059 | 5.9
[6 empty] | -8.957 | 25 | 20.169 | 0.059 | 5.9
[6 empty] | -8.947 | 25 | 20.169 | 0.06 | 6
[6 empty] | -8.937 | 25 | 20.169 | 0.06 | 6
[6 empty] | -8.927 | 25 | 20.169 | 0.061 | 6.1
[6 empty] | -8.917 | 25 | 20.169 | 0.062 | 6.2
[6 empty] | -8.907 | 25 | 20.169 | 0.062 | 6.2
[6 empty] | -8.897 | 25 | 20.169 | 0.063 | 6.3
[6 empty] | -8.887 | 25 | 20.169 | 0.063 | 6.3
[6 empty] | -8.877 | 25 | 20.169 | 0.064 | 6.4
[6 empty] | -8.867 | 25 | 20.169 | 0.064 | 6.4
[6 empty] | -8.857 | 25 | 20.169 | 0.065 | 6.5
[6 empty] | -8.847 | 25 | 20.169 | 0.066 | 6.6
[6 empty] | -8.837 | 25 | 20.169 | 0.066 | 6.6
[6 empty] | -8.827 | 25 | 20.169 | 0.067 | 6.7
[6 empty] | -8.817 | 25 | 20.169 | 0.067 | 6.7
[6 empty] | -8.807 | 25 | 20.169 | 0.068 | 6.8
[6 empty] | -8.797 | 25 | 20.169 | 0.068 | 6.8
[6 empty] | -8.787 | 25 | 20.169 | 0.069 | 6.9
[6 empty] | -8.777 | 25 | 20.169 | 0.07 | 7
[6 empty] | -8.767 | 25 | 20.169 | 0.07 | 7
[6 empty] | -8.757 | 25 | 20.169 | 0.071 | 7.1
[6 empty] | -8.747 | 25 | 20.169 | 0.072 | 7.2
[6 empty] | -8.737 | 25 | 20.169 | 0.073 | 7.3
[6 empty] | -8.727 | 25 | 20.169 | 0.075 | 7.5
[6 empty] | -8.717 | 25 | 20.169 | 0.077 | 7.7
[6 empty] | -8.707 | 25 | 20.169 | 0.078 | 7.8
[6 empty] | -8.697 | 25 | 20.169 | 0.08 | 8
[6 empty] | -8.687 | 25 | 20.169 | 0.081 | 8.1
[6 empty] | -8.677 | 25 | 20.169 | 0.083 | 8.3
[6 empty] | -8.667 | 25 | 20.169 | 0.084 | 8.4
[6 empty] | -8.657 | 25 | 20.169 | 0.086 | 8.6
[6 empty] | -8.647 | 25 | 20.169 | 0.088 | 8.8
[6 empty] | -8.637 | 25 | 20.169 | 0.089 | 8.9
[6 empty] | -8.627 | 25 | 20.169 | 0.091 | 9.1
[6 empty] | -8.617 | 25 | 20.169 | 0.092 | 9.2
[6 empty] | -8.607 | 25 | 20.169 | 0.094 | 9.4
[6 empty] | -8.597 | 25 | 20.169 | 0.095 | 9.5
[6 empty] | -8.587 | 25 | 20.169 | 0.097 | 9.7
[6 empty] | -8.577 | 25 | 20.169 | 0.099 | 9.9
[6 empty] | -8.567 | 25 | 20.169 | 0.1 | 10
[6 empty] | -8.557 | 25 | 20.169 | 0.102 | 10.2
[6 empty] | -8.547 | 25 | 20.17 | 0.106 | 10.6
[6 empty] | -8.537 | 25 | 20.17 | 0.113 | 11.3
[6 empty] | -8.527 | 25 | 20.17 | 0.119 | 11.9
[6 empty] | -8.517 | 25 | 20.17 | 0.125 | 12.5
[6 empty] | -8.507 | 25 | 20.17 | 0.132 | 13.2
[6 empty] | -8.497 | 25 | 20.17 | 0.138 | 13.8
[6 empty] | -8.487 | 25 | 20.17 | 0.145 | 14.5
[6 empty] | -8.477 | 25 | 20.17 | 0.151 | 15.1
[6 empty] | -8.467 | 25 | 20.17 | 0.158 | 15.8
[6 empty] | -8.457 | 25 | 20.17 | 0.164 | 16.4
[6 empty] | -8.447 | 25 | 20.17 | 0.171 | 17.1
[6 empty] | -8.437 | 25 | 20.17 | 0.177 | 17.7
[6 empty] | -8.427 | 25 | 20.17 | 0.183 | 18.3
[6 empty] | -8.417 | 25 | 20.17 | 0.19 | 19
[6 empty] | -8.407 | 25 | 20.17 | 0.196 | 19.6
[6 empty] | -8.397 | 25 | 20.17 | 0.203 | 20.3
[6 empty] | -8.387 | 25 | 20.17 | 0.209 | 20.9
[6 empty] | -8.377 | 25 | 20.17 | 0.216 | 21.6
[6 empty] | -8.367 | 25 | 20.17 | 0.222 | 22.2
[6 empty] | -8.357 | 25 | 20.17 | 0.229 | 22.9
[6 empty] | -8.347 | 25 | 20.17 | 0.236 | 23.6
[6 empty] | -8.337 | 25 | 20.17 | 0.244 | 24.4
[6 empty] | -8.327 | 25 | 20.17 | 0.253 | 25.3
[6 empty] | -8.317 | 25 | 20.17 | 0.261 | 26.1
[6 empty] | -8.307 | 25 | 20.17 | 0.269 | 26.9
[6 empty] | -8.297 | 25 | 20.17 | 0.278 | 27.8
[6 empty] | -8.287 | 25 | 20.17 | 0.286 | 28.6
[6 empty] | -8.277 | 25 | 20.17 | 0.294 | 29.4
[6 empty] | -8.267 | 25 | 20.17 | 0.302 | 30.2
[6 empty] | -8.257 | 25 | 20.17 | 0.311 | 31.1
[6 empty] | -8.247 | 25 | 20.17 | 0.319 | 31.9
[6 empty] | -8.237 | 25 | 20.17 | 0.327 | 32.7
[6 empty] | -8.227 | 25 | 20.17 | 0.336 | 33.6
[6 empty] | -8.217 | 25 | 20.17 | 0.344 | 34.4
[6 empty] | -8.207 | 25 | 20.17 | 0.352 | 35.2
[6 empty] | -8.197 | 25 | 20.17 | 0.361 | 36.1
[6 empty] | -8.187 | 25 | 20.17 | 0.369 | 36.9
[6 empty] | -8.177 | 25 | 20.17 | 0.377 | 37.7
[6 empty] | -8.167 | 25 | 20.17 | 0.386 | 38.6
[6 empty] | -8.157 | 25 | 20.17 | 0.394 | 39.4
[6 empty] | -8.147 | 25 | 20.17 | 0.399 | 39.9
[6 empty] | -8.137 | 25 | 20.17 | 0.402 | 40.2
[6 empty] | -8.127 | 25 | 20.17 | 0.405 | 40.5
[6 empty] | -8.117 | 25 | 20.17 | 0.409 | 40.9
[6 empty] | -8.107 | 25 | 20.17 | 0.412 | 41.2
[6 empty] | -8.097 | 25 | 20.17 | 0.415 | 41.5
[6 empty] | -8.087 | 25 | 20.17 | 0.418 | 41.8
[6 empty] | -8.077 | 25 | 20.17 | 0.421 | 42.1
[6 empty] | -8.067 | 25 | 20.17 | 0.424 | 42.4
[6 empty] | -8.057 | 25 | 20.17 | 0.427 | 42.7
[6 empty] | -8.047 | 25 | 20.17 | 0.43 | 43
[6 empty] | -8.037 | 25 | 20.17 | 0.433 | 43.3
[6 empty] | -8.027 | 25 | 20.17 | 0.436 | 43.6
[6 empty] | -8.017 | 25 | 20.17 | 0.439 | 43.9
[6 empty] | -8.007 | 25 | 20.17 | 0.443 | 44.3
[6 empty] | -7.997 | 25 | 20.17 | 0.446 | 44.6
[6 empty] | -7.987 | 25 | 20.17 | 0.449 | 44.9
[6 empty] | -7.977 | 25 | 20.17 | 0.452 | 45.2
[6 empty] | -7.967 | 25 | 20.17 | 0.455 | 45.5
[6 empty] | -7.957 | 25 | 20.17 | 0.458 | 45.8
[6 empty] | -7.947 | 25 | 20.17 | 0.46 | 46
[6 empty] | -7.937 | 25 | 20.17 | 0.461 | 46.1
[6 empty] | -7.927 | 25 | 20.17 | 0.462 | 46.2
[6 empty] | -7.917 | 25 | 20.17 | 0.463 | 46.3
[6 empty] | -7.907 | 25 | 20.171 | 0.464 | 46.4
[6 empty] | -7.897 | 25 | 20.171 | 0.465 | 46.5
[6 empty] | -7.887 | 25 | 20.171 | 0.466 | 46.6
[6 empty] | -7.877 | 25 | 20.171 | 0.467 | 46.7
[6 empty] | -7.867 | 25 | 20.171 | 0.468 | 46.8
[6 empty] | -7.857 | 25 | 20.171 | 0.469 | 46.9
[6 empty] | -7.847 | 25 | 20.171 | 0.47 | 47
[6 empty] | -7.837 | 25 | 20.171 | 0.47 | 47
[6 empty] | -7.827 | 25 | 20.171 | 0.471 | 47.1
[6 empty] | -7.817 | 25 | 20.171 | 0.472 | 47.2
[6 empty] | -7.807 | 25 | 20.171 | 0.473 | 47.3
[6 empty] | -7.797 | 25 | 20.171 | 0.474 | 47.4
[6 empty] | -7.787 | 25 | 20.171 | 0.475 | 47.5
[6 empty] | -7.777 | 25 | 20.171 | 0.476 | 47.6
[6 empty] | -7.767 | 25 | 20.171 | 0.477 | 47.7
[6 empty] | -7.757 | 25 | 20.171 | 0.478 | 47.8
[6 empty] | -7.747 | 25 | 20.171 | 0.479 | 47.9
[6 empty] | -7.737 | 25 | 20.171 | 0.479 | 47.9
[6 empty] | -7.727 | 25 | 20.171 | 0.48 | 48
[6 empty] | -7.717 | 25 | 20.171 | 0.48 | 48
[6 empty] | -7.707 | 25 | 20.171 | 0.481 | 48.1
[6 empty] | -7.697 | 25 | 20.171 | 0.481 | 48.1
[6 empty] | -7.687 | 25 | 20.171 | 0.482 | 48.2
[6 empty] | -7.677 | 25 | 20.171 | 0.482 | 48.2
[6 empty] | -7.667 | 25 | 20.171 | 0.483 | 48.3
[6 empty] | -7.657 | 25 | 20.171 | 0.483 | 48.3
[6 empty] | -7.647 | 25 | 20.171 | 0.484 | 48.4
[6 empty] | -7.637 | 25 | 20.171 | 0.484 | 48.4
[6 empty] | -7.627 | 25 | 20.171 | 0.485 | 48.5
[6 empty] | -7.617 | 25 | 20.171 | 0.485 | 48.5
[6 empty] | -7.607 | 25 | 20.171 | 0.486 | 48.6
[6 empty] | -7.597 | 25 | 20.171 | 0.486 | 48.6
[6 empty] | -7.587 | 25 | 20.171 | 0.487 | 48.7
[6 empty] | -7.577 | 25 | 20.171 | 0.487 | 48.7
[6 empty] | -7.567 | 25 | 20.171 | 0.488 | 48.8
[6 empty] | -7.557 | 25 | 20.171 | 0.488 | 48.8
[6 empty] | -7.547 | 25 | 20.171 | 0.489 | 48.9
[6 empty] | -7.537 | 25 | 20.171 | 0.489 | 48.9
[6 empty] | -7.527 | 25 | 20.171 | 0.489 | 48.9
[6 empty] | -7.517 | 25 | 20.171 | 0.49 | 49
[6 empty] | -7.507 | 25 | 20.171 | 0.49 | 49
[6 empty] | -7.497 | 25 | 20.171 | 0.49 | 49
[6 empty] | -7.487 | 25 | 20.171 | 0.491 | 49.1
[6 empty] | -7.477 | 25 | 20.171 | 0.491 | 49.1
[6 empty] | -7.467 | 25 | 20.171 | 0.491 | 49.1
[6 empty] | -7.457 | 25 | 20.171 | 0.492 | 49.2
[6 empty] | -7.447 | 25 | 20.171 | 0.492 | 49.2
[6 empty] | -7.437 | 25 | 20.171 | 0.492 | 49.2
[6 empty] | -7.427 | 25 | 20.171 | 0.493 | 49.3
[6 empty] | -7.417 | 25 | 20.171 | 0.493 | 49.3
[6 empty] | -7.407 | 25 | 20.171 | 0.493 | 49.3
[6 empty] | -7.397 | 25 | 20.171 | 0.493 | 49.3
[6 empty] | -7.387 | 25 | 20.171 | 0.494 | 49.4
[6 empty] | -7.377 | 25 | 20.171 | 0.494 | 49.4
[6 empty] | -7.367 | 25 | 20.171 | 0.494 | 49.4
[6 empty] | -7.357 | 25 | 20.171 | 0.495 | 49.5
[6 empty] | -7.347 | 25 | 20.171 | 0.495 | 49.5
[6 empty] | -7.337 | 25 | 20.171 | 0.495 | 49.5
[6 empty] | -7.327 | 25 | 20.171 | 0.495 | 49.5
[6 empty] | -7.317 | 25 | 20.171 | 0.496 | 49.6
[6 empty] | -7.307 | 25 | 20.171 | 0.496 | 49.6
[6 empty] | -7.297 | 25 | 20.171 | 0.496 | 49.6
[6 empty] | -7.287 | 25 | 20.171 | 0.496 | 49.6
[6 empty] | -7.277 | 25 | 20.171 | 0.497 | 49.7
[6 empty] | -7.267 | 25 | 20.172 | 0.497 | 49.7
[6 empty] | -7.257 | 25 | 20.172 | 0.497 | 49.7
[6 empty] | -7.247 | 25 | 20.172 | 0.497 | 49.7
[6 empty] | -7.237 | 25 | 20.172 | 0.497 | 49.7
[6 empty] | -7.227 | 25 | 20.172 | 0.498 | 49.8
[6 empty] | -7.217 | 25 | 20.172 | 0.498 | 49.8
[6 empty] | -7.207 | 25 | 20.172 | 0.498 | 49.8
[6 empty] | -7.197 | 25 | 20.172 | 0.498 | 49.8
[6 empty] | -7.187 | 25 | 20.172 | 0.499 | 49.9
[6 empty] | -7.177 | 25 | 20.172 | 0.499 | 49.9
[6 empty] | -7.167 | 25 | 20.172 | 0.499 | 49.9
[6 empty] | -7.157 | 25 | 20.172 | 0.499 | 49.9
[6 empty] | -7.147 | 25 | 20.172 | 0.499 | 49.9
[6 empty] | -7.137 | 25 | 20.172 | 0.5 | 50
[6 empty] | -7.127 | 25 | 20.172 | 0.5 | 50
[6 empty] | -7.117 | 25 | 20.172 | 0.5 | 50
[6 empty] | -7.107 | 25 | 20.172 | 0.5 | 50
[6 empty] | -7.097 | 25 | 20.172 | 0.5 | 50
[6 empty] | -7.087 | 25 | 20.172 | 0.501 | 50.1
[6 empty] | -7.077 | 25 | 20.172 | 0.501 | 50.1
[6 empty] | -7.067 | 25 | 20.172 | 0.501 | 50.1
[6 empty] | -7.057 | 25 | 20.172 | 0.501 | 50.1
[6 empty] | -7.047 | 25 | 20.172 | 0.502 | 50.2
[6 empty] | -7.037 | 25 | 20.172 | 0.502 | 50.2
[6 empty] | -7.027 | 25 | 20.172 | 0.502 | 50.2
[6 empty] | -7.017 | 25 | 20.172 | 0.502 | 50.2
[6 empty] | -7.007 | 25 | 20.172 | 0.502 | 50.2
[6 empty] | -6.997 | 25 | 20.172 | 0.503 | 50.3
[6 empty] | -6.987 | 25 | 20.172 | 0.503 | 50.3
[6 empty] | -6.977 | 25 | 20.172 | 0.503 | 50.3
[6 empty] | -6.967 | 25 | 20.172 | 0.503 | 50.3
[6 empty] | -6.957 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.947 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.937 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.927 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.917 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.907 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.897 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.887 | 25 | 20.172 | 0.504 | 50.4
[6 empty] | -6.877 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.867 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.857 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.847 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.837 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.827 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.817 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.807 | 25 | 20.172 | 0.505 | 50.5
[6 empty] | -6.797 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.787 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.777 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.767 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.757 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.747 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.737 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.727 | 25 | 20.172 | 0.506 | 50.6
[6 empty] | -6.717 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.707 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.697 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.687 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.677 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.667 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.657 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.647 | 25 | 20.172 | 0.507 | 50.7
[6 empty] | -6.637 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.627 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.617 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.607 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.597 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.587 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.577 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.567 | 25 | 20.173 | 0.508 | 50.8
[6 empty] | -6.557 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.547 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.537 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.527 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.517 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.507 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.497 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.487 | 25 | 20.173 | 0.509 | 50.9
[6 empty] | -6.477 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.467 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.457 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.447 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.437 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.427 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.417 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.407 | 25 | 20.173 | 0.51 | 51
[6 empty] | -6.397 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.387 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.377 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.367 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.357 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.347 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.337 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.327 | 25 | 20.173 | 0.511 | 51.1
[6 empty] | -6.317 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.307 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.297 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.287 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.277 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.267 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.257 | 25 | 20.173 | 0.512 | 51.2
[6 empty] | -6.247 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.237 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.227 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.217 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.207 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.197 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.187 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.177 | 25 | 20.173 | 0.513 | 51.3
[6 empty] | -6.167 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.157 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.147 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.137 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.127 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.117 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.107 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.097 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.087 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.077 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.067 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.057 | 25 | 20.173 | 0.514 | 51.4
[6 empty] | -6.047 | 25 | 20.173 | 0.515 | 51.5
[6 empty] | -6.037 | 25 | 20.173 | 0.515 | 51.5
[6 empty] | -6.027 | 25 | 20.173 | 0.515 | 51.5
[6 empty] | -6.017 | 25 | 20.173 | 0.515 | 51.5
[6 empty] | -6.007 | 25 | 20.173 | 0.515 | 51.5
[6 empty] | -5.997 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.987 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.977 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.967 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.957 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.947 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.937 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.927 | 25 | 20.174 | 0.515 | 51.5
[6 empty] | -5.917 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.907 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.897 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.887 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.877 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.867 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.857 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.847 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.837 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.827 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.817 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.807 | 25 | 20.174 | 0.516 | 51.6
[6 empty] | -5.797 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.787 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.777 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.767 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.757 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.747 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.737 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.727 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.717 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.707 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.697 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.687 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.677 | 25 | 20.174 | 0.517 | 51.7
[6 empty] | -5.667 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.657 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.647 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.637 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.627 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.617 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.607 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.597 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.587 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.577 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.567 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.557 | 25 | 20.174 | 0.518 | 51.8
[6 empty] | -5.547 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.537 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.527 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.517 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.507 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.497 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.487 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.477 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.467 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.457 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.447 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.437 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.427 | 25 | 20.174 | 0.519 | 51.9
[6 empty] | -5.417 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.407 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.397 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.387 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.377 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.367 | 25 | 20.174 | 0.52 | 52
[6 empty] | -5.357 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.347 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.337 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.327 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.317 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.307 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.297 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.287 | 25 | 20.175 | 0.52 | 52
[6 empty] | -5.277 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.267 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.257 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.247 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.237 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.227 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.217 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.207 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.197 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.187 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.177 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.167 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.157 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.147 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.137 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.127 | 25 | 20.175 | 0.521 | 52.1
[6 empty] | -5.117 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.107 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.097 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.087 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.077 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.067 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.057 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.047 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.037 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.027 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.017 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -5.007 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -4.997 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -4.987 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -4.977 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -4.967 | 25 | 20.175 | 0.522 | 52.2
[6 empty] | -4.957 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.947 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.937 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.927 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.917 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.907 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.897 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.887 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.877 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.867 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.857 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.847 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.837 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.827 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.817 | 25 | 20.175 | 0.523 | 52.3
[6 empty] | -4.807 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.797 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.787 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.777 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.767 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.757 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.748 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.738 | 25 | 20.175 | 0.524 | 52.4
[6 empty] | -4.728 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.718 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.708 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.698 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.688 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.678 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.668 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.658 | 25 | 20.176 | 0.524 | 52.4
[6 empty] | -4.648 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.638 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.628 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.618 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.608 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.598 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.588 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.578 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.568 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.558 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.548 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.538 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.528 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.518 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.508 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.498 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.488 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.478 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.468 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.458 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.448 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.438 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.428 | 25 | 20.176 | 0.525 | 52.5
[6 empty] | -4.418 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.408 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.398 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.388 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.378 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.368 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.358 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.348 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.338 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.328 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.318 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.308 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.298 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.288 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.278 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.268 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.258 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.248 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.238 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.228 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.218 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.208 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.198 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.188 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.178 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.168 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.158 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.148 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.138 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.128 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.118 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.108 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.098 | 25 | 20.176 | 0.526 | 52.6
[6 empty] | -4.088 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.078 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.068 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.058 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.048 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.038 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.028 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.018 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -4.008 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.998 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.988 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.978 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.968 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.958 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.948 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.938 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.928 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.918 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.908 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.898 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.888 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.878 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.868 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.858 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.848 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.838 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.828 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.818 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.808 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.798 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.788 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.778 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.768 | 25 | 20.177 | 0.527 | 52.7
[6 empty] | -3.758 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.748 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.738 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.728 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.718 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.708 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.698 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.688 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.678 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.668 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.658 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.648 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.638 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.628 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.618 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.608 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.598 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.588 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.578 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.568 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.558 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.548 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.538 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.528 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.518 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.508 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.498 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.488 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.478 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.468 | 25 | 20.177 | 0.528 | 52.8
[6 empty] | -3.458 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.448 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.438 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.428 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.418 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.408 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.398 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.388 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.378 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.368 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.358 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.348 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.338 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.328 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.318 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.308 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.298 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.288 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.278 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.268 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.258 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.248 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.238 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.228 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.218 | 25 | 20.178 | 0.528 | 52.8
[6 empty] | -3.208 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.198 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.188 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.178 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.168 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.158 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.148 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.138 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.128 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.118 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.108 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.098 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.088 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.078 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.068 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.058 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.048 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.038 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.028 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.018 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -3.008 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.998 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.988 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.978 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.968 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.958 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.948 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.938 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.928 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.918 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.908 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.898 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.888 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.878 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.868 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.858 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.848 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.838 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.828 | 25 | 20.178 | 0.529 | 52.9
[6 empty] | -2.818 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.808 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.798 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.788 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.778 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.768 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.758 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.748 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.738 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.728 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.718 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.708 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.698 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.688 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.678 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.668 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.658 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.648 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.638 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.628 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.618 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.608 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.598 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.588 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.578 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.568 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.558 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.548 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.538 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.528 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.518 | 25 | 20.179 | 0.529 | 52.9
[6 empty] | -2.508 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.498 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.488 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.478 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.468 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.458 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.448 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.438 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.428 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.418 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.408 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.398 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.388 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.378 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.368 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.358 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.348 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.338 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.328 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.318 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.308 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.298 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.288 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.278 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.268 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.258 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.248 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.238 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.228 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.218 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.208 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.198 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.188 | 25 | 20.179 | 0.53 | 53
[6 empty] | -2.178 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.168 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.158 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.148 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.138 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.128 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.118 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.108 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.098 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.088 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.078 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.068 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.058 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.048 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.038 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.028 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.018 | 25 | 20.18 | 0.53 | 53
[6 empty] | -2.008 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.998 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.988 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.978 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.968 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.958 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.948 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.938 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.928 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.918 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.908 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.898 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.888 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.878 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.868 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.858 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.848 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.838 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.828 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.818 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.808 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.798 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.788 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.778 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.768 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.758 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.748 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.738 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.728 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.718 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.708 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.698 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.688 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.678 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.668 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.658 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.648 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.638 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.628 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.618 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.608 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.598 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.588 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.578 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.568 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.558 | 25 | 20.18 | 0.53 | 53
[6 empty] | -1.548 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.538 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.528 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.518 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.508 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.498 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.488 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.478 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.468 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.458 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.448 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.438 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.428 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.418 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.408 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.398 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.388 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.378 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.368 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.358 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.348 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.338 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.328 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.318 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.308 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.298 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.288 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.278 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.268 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.258 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.248 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.238 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.228 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.218 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.208 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.198 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.188 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.178 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.168 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.158 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.148 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.138 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.128 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.118 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.108 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.098 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.088 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.078 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.068 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.058 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.048 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.038 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.028 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.018 | 25 | 20.181 | 0.53 | 53
[6 empty] | -1.008 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.998 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.988 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.978 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.968 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.958 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.948 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.938 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.928 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.918 | 25 | 20.181 | 0.53 | 53
[6 empty] | -0.908 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.898 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.888 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.878 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.868 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.858 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.848 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.838 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.828 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.818 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.808 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.798 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.788 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.778 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.768 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.758 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.748 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.738 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.728 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.718 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.708 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.698 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.688 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.678 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.668 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.658 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.648 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.638 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.628 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.618 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.608 | 25 | 20.182 | 0.53 | 53
[6 empty] | -0.598 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.588 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.578 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.568 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.558 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.548 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.538 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.528 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.518 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.508 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.498 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.488 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.478 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.468 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.458 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.448 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.438 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.428 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.418 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.408 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.398 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.388 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.378 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.368 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.358 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.348 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.338 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.328 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.318 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.308 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.298 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.288 | 25 | 20.182 | 0.529 | 52.9
[6 empty] | -0.278 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.268 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.258 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.248 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.238 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.228 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.218 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.208 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.198 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.188 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.178 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.168 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.158 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.148 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.138 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.128 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.118 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.108 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.098 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.088 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.078 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.068 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.058 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.048 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.038 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.028 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.018 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | -0.008 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.002 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.012 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.022 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.032 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.042 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.052 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.062 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.072 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.082 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.092 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.102 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.112 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.122 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.132 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.142 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.152 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.162 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.172 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.182 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.192 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.202 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.212 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.222 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.232 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.242 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.251 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.261 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.271 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.281 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.291 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.301 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.311 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.321 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.331 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.341 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.351 | 25 | 20.183 | 0.529 | 52.9
[6 empty] | 0.361 | 25 | 20.184 | 0.529 | 52.9
[6 empty] | 0.371 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.381 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.391 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.401 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.411 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.421 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.431 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.441 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.451 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.461 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.471 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.481 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.491 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.501 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.511 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.521 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.531 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.541 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.551 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.561 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.571 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.581 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.591 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.601 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.611 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.621 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.631 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.641 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.651 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.661 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.671 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.681 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.691 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.701 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.711 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.721 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.731 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.741 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.751 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.761 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.771 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.781 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.791 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.801 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.811 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.821 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.831 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.841 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.851 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.861 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.871 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.881 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.891 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.901 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.911 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.921 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.931 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.941 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.951 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.961 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.971 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.981 | 25 | 20.184 | 0.53 | 53
[6 empty] | 0.991 | 25 | 20.184 | 0.53 | 53
[6 empty] | 1.001 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.011 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.021 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.031 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.041 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.051 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.061 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.071 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.081 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.091 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.101 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.111 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.121 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.131 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.141 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.151 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.161 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.171 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.181 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.191 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.201 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.211 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.221 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.231 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.241 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.251 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.261 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.271 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.281 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.291 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.301 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.311 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.321 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.331 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.341 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.351 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.361 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.371 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.381 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.391 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.401 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.411 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.421 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.431 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.441 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.451 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.461 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.471 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.481 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.491 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.501 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.511 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.521 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.531 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.541 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.551 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.561 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.571 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.581 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.591 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.601 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.611 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.621 | 25 | 20.185 | 0.53 | 53
[6 empty] | 1.631 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.641 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.651 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.661 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.671 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.681 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.691 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.701 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.711 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.721 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.731 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.741 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.751 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.761 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.771 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.781 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.791 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.801 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.811 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.821 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.831 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.841 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.851 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.861 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.871 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.881 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.891 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.901 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.911 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.921 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.931 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.941 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.951 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.961 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.971 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.981 | 25 | 20.186 | 0.53 | 53
[6 empty] | 1.991 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.001 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.011 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.021 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.031 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.041 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.051 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.061 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.071 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.081 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.091 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.101 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.111 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.121 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.131 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.141 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.151 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.161 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.171 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.181 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.191 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.201 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.211 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.221 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.231 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.241 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.251 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.261 | 25 | 20.186 | 0.53 | 53
[6 empty] | 2.271 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.281 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.291 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.301 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.311 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.321 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.331 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.341 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.351 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.361 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.371 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.381 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.391 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.401 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.411 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.421 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.431 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.441 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.451 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.461 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.471 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.481 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.491 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.501 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.511 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.521 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.531 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.541 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.551 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.561 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.571 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.581 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.591 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.601 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.611 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.621 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.631 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.641 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.651 | 25 | 20.187 | 0.53 | 53
[6 empty] | 2.661 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.671 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.681 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.691 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.701 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.711 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.721 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.731 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.741 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.751 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.761 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.771 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.781 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.791 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.801 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.811 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.821 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.831 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.841 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.851 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.861 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.871 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.881 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.891 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.901 | 25 | 20.187 | 0.529 | 52.9
[6 empty] | 2.911 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.921 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.931 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.941 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.951 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.961 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.971 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.981 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 2.991 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.001 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.011 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.021 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.031 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.041 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.051 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.061 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.071 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.081 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.091 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.101 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.111 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.121 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.131 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.141 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.151 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.161 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.171 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.181 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.191 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.201 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.211 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.221 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.231 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.241 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.251 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.261 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.271 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.281 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.291 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.301 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.311 | 25 | 20.188 | 0.529 | 52.9
[6 empty] | 3.321 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.331 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.341 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.351 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.361 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.371 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.381 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.391 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.401 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.411 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.421 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.431 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.441 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.451 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.461 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.471 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.481 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.491 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.501 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.511 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.521 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.531 | 25 | 20.188 | 0.528 | 52.8
[6 empty] | 3.541 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.551 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.561 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.571 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.581 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.591 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.601 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.611 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.621 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.631 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.641 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.651 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.661 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.671 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.681 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.691 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.701 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.711 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.721 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.731 | 25 | 20.189 | 0.528 | 52.8
[6 empty] | 3.741 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.751 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.761 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.771 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.781 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.791 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.801 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.811 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.821 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.831 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.841 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.851 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.861 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.871 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.881 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.891 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.901 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.911 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.921 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.931 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.941 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.951 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.961 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.971 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.981 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 3.991 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.001 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.011 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.021 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.031 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.041 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.051 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.061 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.071 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.081 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.091 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.101 | 25 | 20.189 | 0.527 | 52.7
[6 empty] | 4.111 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.121 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.131 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.141 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.151 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.161 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.171 | 25 | 20.189 | 0.526 | 52.6
[6 empty] | 4.181 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.191 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.201 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.211 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.221 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.231 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.241 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.251 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.261 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.271 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.281 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.291 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.301 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.311 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.321 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.331 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.341 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.351 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.361 | 25 | 20.19 | 0.526 | 52.6
[6 empty] | 4.371 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.381 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.391 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.401 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.411 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.421 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.431 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.441 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.451 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.461 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.471 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.481 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.491 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.501 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.511 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.521 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.531 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.541 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.551 | 25 | 20.19 | 0.525 | 52.5
[6 empty] | 4.561 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.571 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.581 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.591 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.601 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.611 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.621 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.631 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.641 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.651 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.661 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.671 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.681 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.691 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.701 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.711 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.721 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.731 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.741 | 25 | 20.19 | 0.524 | 52.4
[6 empty] | 4.751 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.761 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.771 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.781 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.791 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.801 | 25 | 20.19 | 0.523 | 52.3
[6 empty] | 4.811 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.821 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.831 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.841 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.851 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.861 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.871 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.881 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.891 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.901 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.911 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.921 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.931 | 25 | 20.191 | 0.523 | 52.3
[6 empty] | 4.941 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 4.951 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 4.961 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 4.971 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 4.981 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 4.991 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.001 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.011 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.021 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.031 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.041 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.051 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.061 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.071 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.081 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.091 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.101 | 25 | 20.191 | 0.522 | 52.2
[6 empty] | 5.111 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.121 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.131 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.141 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.151 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.161 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.171 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.181 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.191 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.201 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.211 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.221 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.231 | 25 | 20.191 | 0.521 | 52.1
[6 empty] | 5.241 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.25 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.26 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.27 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.28 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.29 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.3 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.31 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.32 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.33 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.34 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.35 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.36 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.37 | 25 | 20.191 | 0.52 | 52
[6 empty] | 5.38 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.39 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.4 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.41 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.42 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.43 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.44 | 25 | 20.191 | 0.519 | 51.9
[6 empty] | 5.45 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.46 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.47 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.48 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.49 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.5 | 25 | 20.192 | 0.519 | 51.9
[6 empty] | 5.51 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.52 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.53 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.54 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.55 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.56 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.57 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.58 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.59 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.6 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.61 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.62 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.63 | 25 | 20.192 | 0.518 | 51.8
[6 empty] | 5.64 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.65 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.66 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.67 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.68 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.69 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.7 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.71 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.72 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.73 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.74 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.75 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.76 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.77 | 25 | 20.192 | 0.517 | 51.7
[6 empty] | 5.78 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.79 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.8 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.81 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.82 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.83 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.84 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.85 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.86 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.87 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.88 | 25 | 20.192 | 0.516 | 51.6
[6 empty] | 5.89 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.9 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.91 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.92 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.93 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.94 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.95 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.96 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.97 | 25 | 20.192 | 0.515 | 51.5
[6 empty] | 5.98 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 5.99 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.01 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.02 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.03 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.04 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.05 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.06 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.07 | 25 | 20.192 | 0.514 | 51.4
[6 empty] | 6.08 | 25 | 20.192 | 0.513 | 51.3
[6 empty] | 6.09 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.1 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.11 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.12 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.13 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.14 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.15 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.16 | 25 | 20.193 | 0.513 | 51.3
[6 empty] | 6.17 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.18 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.19 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.2 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.21 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.22 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.23 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.24 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.25 | 25 | 20.193 | 0.512 | 51.2
[6 empty] | 6.26 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.27 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.28 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.29 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.3 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.31 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.32 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.33 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.34 | 25 | 20.193 | 0.511 | 51.1
[6 empty] | 6.35 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.36 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.37 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.38 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.39 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.4 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.41 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.42 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.43 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.44 | 25 | 20.193 | 0.51 | 51
[6 empty] | 6.45 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.46 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.47 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.48 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.49 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.5 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.51 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.52 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.53 | 25 | 20.193 | 0.509 | 50.9
[6 empty] | 6.54 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.55 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.56 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.57 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.58 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.59 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.6 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.61 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.62 | 25 | 20.193 | 0.508 | 50.8
[6 empty] | 6.63 | 25 | 20.193 | 0.507 | 50.7
[6 empty] | 6.64 | 25 | 20.193 | 0.507 | 50.7
[6 empty] | 6.65 | 25 | 20.193 | 0.507 | 50.7
[6 empty] | 6.66 | 25 | 20.193 | 0.507 | 50.7
[6 empty] | 6.67 | 25 | 20.193 | 0.507 | 50.7
[6 empty] | 6.68 | 25 | 20.193 | 0.506 | 50.6
[6 empty] | 6.69 | 25 | 20.193 | 0.506 | 50.6
[6 empty] | 6.7 | 25 | 20.193 | 0.506 | 50.6
[6 empty] | 6.71 | 25 | 20.193 | 0.506 | 50.6
[6 empty] | 6.72 | 25 | 20.194 | 0.506 | 50.6
[6 empty] | 6.73 | 25 | 20.194 | 0.505 | 50.5
[6 empty] | 6.74 | 25 | 20.194 | 0.505 | 50.5
[6 empty] | 6.75 | 25 | 20.194 | 0.505 | 50.5
[6 empty] | 6.76 | 25 | 20.194 | 0.505 | 50.5
[6 empty] | 6.77 | 25 | 20.194 | 0.504 | 50.4
[6 empty] | 6.78 | 25 | 20.194 | 0.504 | 50.4
[6 empty] | 6.79 | 25 | 20.194 | 0.504 | 50.4
[6 empty] | 6.8 | 25 | 20.194 | 0.504 | 50.4
[6 empty] | 6.81 | 25 | 20.194 | 0.504 | 50.4
[6 empty] | 6.82 | 25 | 20.194 | 0.503 | 50.3
[6 empty] | 6.83 | 25 | 20.194 | 0.503 | 50.3
[6 empty] | 6.84 | 25 | 20.194 | 0.503 | 50.3
[6 empty] | 6.85 | 25 | 20.194 | 0.503 | 50.3
[6 empty] | 6.86 | 25 | 20.194 | 0.502 | 50.2
[6 empty] | 6.87 | 25 | 20.194 | 0.502 | 50.2
[6 empty] | 6.88 | 25 | 20.194 | 0.502 | 50.2
[6 empty] | 6.89 | 25 | 20.194 | 0.502 | 50.2
[6 empty] | 6.9 | 25 | 20.194 | 0.502 | 50.2
[6 empty] | 6.91 | 25 | 20.194 | 0.501 | 50.1
[6 empty] | 6.92 | 25 | 20.194 | 0.501 | 50.1
[6 empty] | 6.93 | 25 | 20.194 | 0.501 | 50.1
[6 empty] | 6.94 | 25 | 20.194 | 0.501 | 50.1
[6 empty] | 6.95 | 25 | 20.194 | 0.5 | 50
[6 empty] | 6.96 | 25 | 20.194 | 0.5 | 50
[6 empty] | 6.97 | 25 | 20.194 | 0.5 | 50
[6 empty] | 6.98 | 25 | 20.194 | 0.5 | 50
[6 empty] | 6.99 | 25 | 20.194 | 0.499 | 49.9
[6 empty] | 7 | 25 | 20.194 | 0.499 | 49.9
[6 empty] | 7.01 | 25 | 20.194 | 0.499 | 49.9
[6 empty] | 7.02 | 25 | 20.194 | 0.499 | 49.9
[6 empty] | 7.03 | 25 | 20.194 | 0.499 | 49.9
[6 empty] | 7.04 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.05 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.06 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.07 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.08 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.09 | 25 | 20.194 | 0.498 | 49.8
[6 empty] | 7.1 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.11 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.12 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.13 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.14 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.15 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.16 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.17 | 25 | 20.194 | 0.497 | 49.7
[6 empty] | 7.18 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.19 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.2 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.21 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.22 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.23 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.24 | 25 | 20.194 | 0.496 | 49.6
[6 empty] | 7.25 | 25 | 20.194 | 0.495 | 49.5
[6 empty] | 7.26 | 25 | 20.194 | 0.495 | 49.5
[6 empty] | 7.27 | 25 | 20.194 | 0.495 | 49.5
[6 empty] | 7.28 | 25 | 20.194 | 0.494 | 49.4
[6 empty] | 7.29 | 25 | 20.194 | 0.494 | 49.4
[6 empty] | 7.3 | 25 | 20.194 | 0.494 | 49.4
[6 empty] | 7.31 | 25 | 20.194 | 0.493 | 49.3
[6 empty] | 7.32 | 25 | 20.194 | 0.493 | 49.3
[6 empty] | 7.33 | 25 | 20.194 | 0.493 | 49.3
[6 empty] | 7.34 | 25 | 20.194 | 0.493 | 49.3
[6 empty] | 7.35 | 25 | 20.194 | 0.492 | 49.2
[6 empty] | 7.36 | 25 | 20.195 | 0.492 | 49.2
[6 empty] | 7.37 | 25 | 20.195 | 0.492 | 49.2
[6 empty] | 7.38 | 25 | 20.195 | 0.491 | 49.1
[6 empty] | 7.39 | 25 | 20.195 | 0.491 | 49.1
[6 empty] | 7.4 | 25 | 20.195 | 0.491 | 49.1
[6 empty] | 7.41 | 25 | 20.195 | 0.49 | 49
[6 empty] | 7.42 | 25 | 20.195 | 0.49 | 49
[6 empty] | 7.43 | 25 | 20.195 | 0.49 | 49
[6 empty] | 7.44 | 25 | 20.195 | 0.489 | 48.9
[6 empty] | 7.45 | 25 | 20.195 | 0.489 | 48.9
[6 empty] | 7.46 | 25 | 20.195 | 0.488 | 48.8
[6 empty] | 7.47 | 25 | 20.195 | 0.488 | 48.8
[6 empty] | 7.48 | 25 | 20.195 | 0.487 | 48.7
[6 empty] | 7.49 | 25 | 20.195 | 0.487 | 48.7
[6 empty] | 7.5 | 25 | 20.195 | 0.486 | 48.6
[6 empty] | 7.51 | 25 | 20.195 | 0.486 | 48.6
[6 empty] | 7.52 | 25 | 20.195 | 0.485 | 48.5
[6 empty] | 7.53 | 25 | 20.195 | 0.485 | 48.5
[6 empty] | 7.54 | 25 | 20.195 | 0.484 | 48.4
[6 empty] | 7.55 | 25 | 20.195 | 0.484 | 48.4
[6 empty] | 7.56 | 25 | 20.195 | 0.483 | 48.3
[6 empty] | 7.57 | 25 | 20.195 | 0.483 | 48.3
[6 empty] | 7.58 | 25 | 20.195 | 0.482 | 48.2
[6 empty] | 7.59 | 25 | 20.195 | 0.482 | 48.2
[6 empty] | 7.6 | 25 | 20.195 | 0.481 | 48.1
[6 empty] | 7.61 | 25 | 20.195 | 0.481 | 48.1
[6 empty] | 7.62 | 25 | 20.195 | 0.48 | 48
[6 empty] | 7.63 | 25 | 20.195 | 0.48 | 48
[6 empty] | 7.64 | 25 | 20.195 | 0.479 | 47.9
[6 empty] | 7.65 | 25 | 20.195 | 0.479 | 47.9
[6 empty] | 7.66 | 25 | 20.195 | 0.478 | 47.8
[6 empty] | 7.67 | 25 | 20.195 | 0.477 | 47.7
[6 empty] | 7.68 | 25 | 20.195 | 0.476 | 47.6
[6 empty] | 7.69 | 25 | 20.195 | 0.475 | 47.5
[6 empty] | 7.7 | 25 | 20.195 | 0.474 | 47.4
[6 empty] | 7.71 | 25 | 20.195 | 0.473 | 47.3
[6 empty] | 7.72 | 25 | 20.195 | 0.472 | 47.2
[6 empty] | 7.73 | 25 | 20.195 | 0.471 | 47.1
[6 empty] | 7.74 | 25 | 20.195 | 0.47 | 47
[6 empty] | 7.75 | 25 | 20.195 | 0.469 | 46.9
[6 empty] | 7.76 | 25 | 20.195 | 0.468 | 46.8
[6 empty] | 7.77 | 25 | 20.195 | 0.467 | 46.7
[6 empty] | 7.78 | 25 | 20.195 | 0.466 | 46.6
[6 empty] | 7.79 | 25 | 20.195 | 0.465 | 46.5
[6 empty] | 7.8 | 25 | 20.195 | 0.464 | 46.4
[6 empty] | 7.81 | 25 | 20.195 | 0.463 | 46.3
[6 empty] | 7.82 | 25 | 20.195 | 0.462 | 46.2
[6 empty] | 7.83 | 25 | 20.195 | 0.461 | 46.1
[6 empty] | 7.84 | 25 | 20.195 | 0.46 | 46
[6 empty] | 7.85 | 25 | 20.195 | 0.458 | 45.8
[6 empty] | 7.86 | 25 | 20.195 | 0.455 | 45.5
[6 empty] | 7.87 | 25 | 20.195 | 0.452 | 45.2
[6 empty] | 7.88 | 25 | 20.195 | 0.449 | 44.9
[6 empty] | 7.89 | 25 | 20.195 | 0.446 | 44.6
[6 empty] | 7.9 | 25 | 20.195 | 0.442 | 44.2
[6 empty] | 7.91 | 25 | 20.195 | 0.439 | 43.9
[6 empty] | 7.92 | 25 | 20.195 | 0.436 | 43.6
[6 empty] | 7.93 | 25 | 20.195 | 0.433 | 43.3
[6 empty] | 7.94 | 25 | 20.195 | 0.43 | 43
[6 empty] | 7.95 | 25 | 20.195 | 0.427 | 42.7
[6 empty] | 7.96 | 25 | 20.195 | 0.424 | 42.4
[6 empty] | 7.97 | 25 | 20.195 | 0.421 | 42.1
[6 empty] | 7.98 | 25 | 20.195 | 0.417 | 41.7
[6 empty] | 7.99 | 25 | 20.196 | 0.414 | 41.4
[6 empty] | 8 | 25 | 20.196 | 0.411 | 41.1
[6 empty] | 8.01 | 25 | 20.196 | 0.408 | 40.8
[6 empty] | 8.02 | 25 | 20.196 | 0.405 | 40.5
[6 empty] | 8.03 | 25 | 20.196 | 0.402 | 40.2
[6 empty] | 8.04 | 25 | 20.196 | 0.399 | 39.9
[6 empty] | 8.05 | 25 | 20.196 | 0.393 | 39.3
[6 empty] | 8.06 | 25 | 20.196 | 0.384 | 38.4
[6 empty] | 8.07 | 25 | 20.196 | 0.376 | 37.6
[6 empty] | 8.08 | 25 | 20.196 | 0.368 | 36.8
[6 empty] | 8.09 | 25 | 20.196 | 0.359 | 35.9
[6 empty] | 8.1 | 25 | 20.196 | 0.351 | 35.1
[6 empty] | 8.11 | 25 | 20.196 | 0.343 | 34.3
[6 empty] | 8.12 | 25 | 20.196 | 0.334 | 33.4
[6 empty] | 8.13 | 25 | 20.196 | 0.326 | 32.6
[6 empty] | 8.14 | 25 | 20.196 | 0.317 | 31.7
[6 empty] | 8.15 | 25 | 20.196 | 0.309 | 30.9
[6 empty] | 8.16 | 25 | 20.196 | 0.301 | 30.1
[6 empty] | 8.17 | 25 | 20.196 | 0.292 | 29.2
[6 empty] | 8.18 | 25 | 20.196 | 0.284 | 28.4
[6 empty] | 8.19 | 25 | 20.196 | 0.276 | 27.6
[6 empty] | 8.2 | 25 | 20.196 | 0.267 | 26.7
[6 empty] | 8.21 | 25 | 20.196 | 0.259 | 25.9
[6 empty] | 8.22 | 25 | 20.196 | 0.25 | 25
[6 empty] | 8.23 | 25 | 20.196 | 0.242 | 24.2
[6 empty] | 8.24 | 25 | 20.196 | 0.234 | 23.4
[6 empty] | 8.25 | 25 | 20.196 | 0.226 | 22.6
[6 empty] | 8.26 | 25 | 20.196 | 0.22 | 22
[6 empty] | 8.27 | 25 | 20.196 | 0.214 | 21.4
[6 empty] | 8.28 | 25 | 20.196 | 0.207 | 20.7
[6 empty] | 8.29 | 25 | 20.196 | 0.201 | 20.1
[6 empty] | 8.3 | 25 | 20.196 | 0.194 | 19.4
[6 empty] | 8.31 | 25 | 20.196 | 0.188 | 18.8
[6 empty] | 8.32 | 25 | 20.196 | 0.182 | 18.2
[6 empty] | 8.33 | 25 | 20.196 | 0.175 | 17.5
[6 empty] | 8.34 | 25 | 20.196 | 0.169 | 16.9
[6 empty] | 8.35 | 25 | 20.196 | 0.163 | 16.3
[6 empty] | 8.36 | 25 | 20.196 | 0.156 | 15.6
[6 empty] | 8.37 | 25 | 20.196 | 0.15 | 15
[6 empty] | 8.38 | 25 | 20.196 | 0.143 | 14.3
[6 empty] | 8.39 | 25 | 20.196 | 0.137 | 13.7
[6 empty] | 8.4 | 25 | 20.196 | 0.131 | 13.1
[6 empty] | 8.41 | 25 | 20.196 | 0.124 | 12.4
[6 empty] | 8.42 | 25 | 20.196 | 0.118 | 11.8
[6 empty] | 8.43 | 25 | 20.196 | 0.112 | 11.2
[6 empty] | 8.44 | 25 | 20.196 | 0.105 | 10.5
[6 empty] | 8.45 | 25 | 20.196 | 0.101 | 10.1
[6 empty] | 8.46 | 25 | 20.196 | 0.1 | 10
[6 empty] | 8.47 | 25 | 20.196 | 0.098 | 9.8
[6 empty] | 8.48 | 25 | 20.196 | 0.097 | 9.7
[6 empty] | 8.49 | 25 | 20.196 | 0.095 | 9.5
[6 empty] | 8.5 | 25 | 20.196 | 0.093 | 9.3
[6 empty] | 8.51 | 25 | 20.196 | 0.092 | 9.2
[6 empty] | 8.52 | 25 | 20.196 | 0.09 | 9
[6 empty] | 8.53 | 25 | 20.196 | 0.089 | 8.9
[6 empty] | 8.54 | 25 | 20.196 | 0.087 | 8.7
[6 empty] | 8.55 | 25 | 20.196 | 0.086 | 8.6
[6 empty] | 8.56 | 25 | 20.196 | 0.084 | 8.4
[6 empty] | 8.57 | 25 | 20.196 | 0.083 | 8.3
[6 empty] | 8.58 | 25 | 20.196 | 0.081 | 8.1
[6 empty] | 8.59 | 25 | 20.196 | 0.079 | 7.9
[6 empty] | 8.6 | 25 | 20.196 | 0.078 | 7.8
[6 empty] | 8.61 | 25 | 20.196 | 0.076 | 7.6
[6 empty] | 8.62 | 25 | 20.196 | 0.075 | 7.5
[6 empty] | 8.63 | 25 | 20.197 | 0.073 | 7.3
[6 empty] | 8.64 | 25 | 20.197 | 0.072 | 7.2
[6 empty] | 8.65 | 25 | 20.197 | 0.071 | 7.1
[6 empty] | 8.66 | 25 | 20.197 | 0.07 | 7
[6 empty] | 8.67 | 25 | 20.197 | 0.07 | 7
[6 empty] | 8.68 | 25 | 20.197 | 0.069 | 6.9
[6 empty] | 8.69 | 25 | 20.197 | 0.069 | 6.9
[6 empty] | 8.7 | 25 | 20.197 | 0.068 | 6.8
[6 empty] | 8.71 | 25 | 20.197 | 0.068 | 6.8
[6 empty] | 8.72 | 25 | 20.197 | 0.068 | 6.8
[6 empty] | 8.73 | 25 | 20.197 | 0.067 | 6.7
[6 empty] | 8.74 | 25 | 20.197 | 0.067 | 6.7
[6 empty] | 8.75 | 25 | 20.197 | 0.066 | 6.6
[6 empty] | 8.76 | 25 | 20.197 | 0.066 | 6.6
[6 empty] | 8.77 | 25 | 20.197 | 0.065 | 6.5
[6 empty] | 8.78 | 25 | 20.197 | 0.065 | 6.5
[6 empty] | 8.79 | 25 | 20.197 | 0.064 | 6.4
[6 empty] | 8.8 | 25 | 20.197 | 0.064 | 6.4
[6 empty] | 8.81 | 25 | 20.197 | 0.064 | 6.4
[6 empty] | 8.82 | 25 | 20.197 | 0.063 | 6.3
[6 empty] | 8.83 | 25 | 20.197 | 0.063 | 6.3
[6 empty] | 8.84 | 25 | 20.197 | 0.062 | 6.2
[6 empty] | 8.85 | 25 | 20.197 | 0.062 | 6.2
[6 empty] | 8.86 | 25 | 20.197 | 0.061 | 6.1
[6 empty] | 8.87 | 25 | 20.197 | 0.061 | 6.1
[6 empty] | 8.88 | 25 | 20.197 | 0.06 | 6
[6 empty] | 8.89 | 25 | 20.197 | 0.06 | 6
[6 empty] | 8.9 | 25 | 20.197 | 0.06 | 6
[6 empty] | 8.91 | 25 | 20.197 | 0.059 | 5.9
[6 empty] | 8.92 | 25 | 20.197 | 0.059 | 5.9
[6 empty] | 8.93 | 25 | 20.197 | 0.058 | 5.8
[6 empty] | 8.94 | 25 | 20.197 | 0.058 | 5.8
[6 empty] | 8.95 | 25 | 20.197 | 0.057 | 5.7
[6 empty] | 8.96 | 25 | 20.197 | 0.057 | 5.7
[6 empty] | 8.97 | 25 | 20.197 | 0.057 | 5.7
[6 empty] | 8.98 | 25 | 20.197 | 0.056 | 5.6
[6 empty] | 8.99 | 25 | 20.197 | 0.056 | 5.6
[6 empty] | 9 | 25 | 20.197 | 0.055 | 5.5
[6 empty] | 9.01 | 25 | 20.197 | 0.055 | 5.5
[6 empty] | 9.02 | 25 | 20.197 | 0.054 | 5.4
[6 empty] | 9.03 | 25 | 20.197 | 0.054 | 5.4
[6 empty] | 9.04 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.05 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.06 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.07 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.08 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.09 | 25 | 20.197 | 0.053 | 5.3
[6 empty] | 9.1 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.11 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.12 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.13 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.14 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.15 | 25 | 20.197 | 0.052 | 5.2
[6 empty] | 9.16 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.17 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.18 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.19 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.2 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.21 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.22 | 25 | 20.197 | 0.051 | 5.1
[6 empty] | 9.23 | 25 | 20.197 | 0.05 | 5
[6 empty] | 9.24 | 25 | 20.197 | 0.05 | 5
[6 empty] | 9.25 | 25 | 20.197 | 0.05 | 5
[6 empty] | 9.26 | 25 | 20.197 | 0.05 | 5
[6 empty] | 9.27 | 25 | 20.198 | 0.05 | 5
[6 empty] | 9.28 | 25 | 20.198 | 0.05 | 5
[6 empty] | 9.29 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.3 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.31 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.32 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.33 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.34 | 25 | 20.198 | 0.049 | 4.9
[6 empty] | 9.35 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.36 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.37 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.38 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.39 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.4 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.41 | 25 | 20.198 | 0.048 | 4.8
[6 empty] | 9.42 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.43 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.44 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.45 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.46 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.47 | 25 | 20.198 | 0.047 | 4.7
[6 empty] | 9.48 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.49 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.5 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.51 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.52 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.53 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.54 | 25 | 20.198 | 0.046 | 4.6
[6 empty] | 9.55 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.56 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.57 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.58 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.59 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.6 | 25 | 20.198 | 0.045 | 4.5
[6 empty] | 9.61 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.62 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.63 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.64 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.65 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.66 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.67 | 25 | 20.198 | 0.044 | 4.4
[6 empty] | 9.68 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.69 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.7 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.71 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.72 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.73 | 25 | 20.198 | 0.043 | 4.3
[6 empty] | 9.74 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.75 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.76 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.77 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.78 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.79 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.8 | 25 | 20.198 | 0.042 | 4.2
[6 empty] | 9.81 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.82 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.83 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.84 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.85 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.86 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.87 | 25 | 20.198 | 0.041 | 4.1
[6 empty] | 9.88 | 25 | 20.198 | 0.04 | 4
[6 empty] | 9.89 | 25 | 20.198 | 0.04 | 4
[6 empty] | 9.9 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.91 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.92 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.93 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.94 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.95 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.96 | 25 | 20.199 | 0.04 | 4
[6 empty] | 9.97 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 9.98 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 9.99 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.01 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.02 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.03 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.04 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.05 | 25 | 20.199 | 0.039 | 3.9
[6 empty] | 10.06 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.07 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.08 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.09 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.1 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.11 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.12 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.13 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.14 | 25 | 20.199 | 0.038 | 3.8
[6 empty] | 10.15 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.16 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.17 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.18 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.19 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.2 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.21 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.22 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.23 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.24 | 25 | 20.199 | 0.037 | 3.7
[6 empty] | 10.249 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.259 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.269 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.279 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.289 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.299 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.309 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.319 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.329 | 25 | 20.199 | 0.036 | 3.6
[6 empty] | 10.339 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.349 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.359 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.369 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.379 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.389 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.399 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.409 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.419 | 25 | 20.199 | 0.035 | 3.5
[6 empty] | 10.429 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.439 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.449 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.459 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.469 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.479 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.489 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.499 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.509 | 25 | 20.199 | 0.034 | 3.4
[6 empty] | 10.519 | 25 | 20.199 | 0.033 | 3.3
[6 empty] | 10.529 | 25 | 20.199 | 0.033 | 3.3
[6 empty] | 10.539 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.549 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.559 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.569 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.579 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.589 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.599 | 25 | 20.2 | 0.033 | 3.3
[6 empty] | 10.609 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.619 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.629 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.639 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.649 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.659 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.669 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.679 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.689 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.699 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.709 | 25 | 20.2 | 0.032 | 3.2
[6 empty] | 10.719 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.729 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.739 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.749 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.759 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.769 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.779 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.789 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.799 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.809 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.819 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.829 | 25 | 20.2 | 0.031 | 3.1
[6 empty] | 10.839 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.849 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.859 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.869 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.879 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.889 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.899 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.909 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.919 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.929 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.939 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.949 | 25 | 20.2 | 0.03 | 3
[6 empty] | 10.959 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 10.969 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 10.979 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 10.989 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 10.999 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.009 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.019 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.029 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.039 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.049 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.059 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.069 | 25 | 20.2 | 0.029 | 2.9
[6 empty] | 11.079 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.089 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.099 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.109 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.119 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.129 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.139 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.149 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.159 | 25 | 20.2 | 0.028 | 2.8
[6 empty] | 11.169 | 25 | 20.201 | 0.028 | 2.8
[6 empty] | 11.179 | 25 | 20.201 | 0.028 | 2.8
[6 empty] | 11.189 | 25 | 20.201 | 0.028 | 2.8
[6 empty] | 11.199 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.209 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.219 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.229 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.239 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.249 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.259 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.269 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.279 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.289 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.299 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.309 | 25 | 20.201 | 0.027 | 2.7
[6 empty] | 11.319 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.329 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.339 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.349 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.359 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.369 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.379 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.389 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.399 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.409 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.419 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.429 | 25 | 20.201 | 0.026 | 2.6
[6 empty] | 11.439 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.449 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.459 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.469 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.479 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.489 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.499 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.509 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.519 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.529 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.539 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.549 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.559 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.569 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.579 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.589 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.599 | 25 | 20.201 | 0.025 | 2.5
[6 empty] | 11.609 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.619 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.629 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.639 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.649 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.659 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.669 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.679 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.689 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.699 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.709 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.719 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.729 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.739 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.749 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.759 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.769 | 25 | 20.201 | 0.024 | 2.4
[6 empty] | 11.779 | 25 | 20.201 | 0.023 | 2.3
[6 empty] | 11.789 | 25 | 20.201 | 0.023 | 2.3
[6 empty] | 11.799 | 25 | 20.201 | 0.023 | 2.3
[6 empty] | 11.809 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.819 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.829 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.839 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.849 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.859 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.869 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.879 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.889 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.899 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.909 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.919 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.929 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.939 | 25 | 20.202 | 0.023 | 2.3
[6 empty] | 11.949 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 11.959 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 11.969 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 11.979 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 11.989 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 11.999 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.009 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.019 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.029 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.039 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.049 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.059 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.069 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.079 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.089 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.099 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.109 | 25 | 20.202 | 0.022 | 2.2
[6 empty] | 12.119 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.129 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.139 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.149 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.159 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.169 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.179 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.189 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.199 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.209 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.219 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.229 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.239 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.249 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.259 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.269 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.279 | 25 | 20.202 | 0.021 | 2.1
[6 empty] | 12.289 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.299 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.309 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.319 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.329 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.339 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.349 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.359 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.369 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.379 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.389 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.399 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.409 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.419 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.429 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.439 | 25 | 20.202 | 0.02 | 2
[6 empty] | 12.449 | 25 | 20.203 | 0.02 | 2
[6 empty] | 12.459 | 25 | 20.203 | 0.02 | 2
[6 empty] | 12.469 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.479 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.489 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.499 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.509 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.519 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.529 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.539 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.549 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.559 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.569 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.579 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.589 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.599 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.609 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.619 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.629 | 25 | 20.203 | 0.019 | 1.9
[6 empty] | 12.639 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.649 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.659 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.669 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.679 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.689 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.699 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.709 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.719 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.729 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.739 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.749 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.759 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.769 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.779 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.789 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.799 | 25 | 20.203 | 0.018 | 1.8
[6 empty] | 12.809 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.819 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.829 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.839 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.849 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.859 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.869 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.879 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.889 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.899 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.909 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.919 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.929 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.939 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.949 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.959 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.969 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.979 | 25 | 20.203 | 0.017 | 1.7
[6 empty] | 12.989 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 12.999 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.009 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.019 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.029 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.039 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.049 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.059 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.069 | 25 | 20.203 | 0.016 | 1.6
[6 empty] | 13.079 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.089 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.099 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.109 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.119 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.129 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.139 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.149 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.159 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.169 | 25 | 20.204 | 0.016 | 1.6
[6 empty] | 13.179 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.189 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.199 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.209 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.219 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.229 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.239 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.249 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.259 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.269 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.279 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.289 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.299 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.309 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.319 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.329 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.339 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.349 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.359 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.369 | 25 | 20.204 | 0.015 | 1.5
[6 empty] | 13.379 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.389 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.399 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.409 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.419 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.429 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.439 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.449 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.459 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.469 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.479 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.489 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.499 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.509 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.519 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.529 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.539 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.549 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.559 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.569 | 25 | 20.204 | 0.014 | 1.4
[6 empty] | 13.579 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.589 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.599 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.609 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.619 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.629 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.639 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.649 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.659 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.669 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.679 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.689 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.699 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.709 | 25 | 20.204 | 0.013 | 1.3
[6 empty] | 13.719 | 25 | 20.205 | 0.013 | 1.3
[6 empty] | 13.729 | 25 | 20.205 | 0.013 | 1.3
[6 empty] | 13.739 | 25 | 20.205 | 0.013 | 1.3
[6 empty] | 13.749 | 25 | 20.205 | 0.013 | 1.3
[6 empty] | 13.759 | 25 | 20.205 | 0.013 | 1.3
[6 empty] | 13.769 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.779 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.789 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.799 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.809 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.819 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.829 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.839 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.849 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.859 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.869 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.879 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.889 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.899 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.909 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.919 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.929 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.939 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.949 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.959 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.969 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.979 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.989 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 13.999 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.009 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.019 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.029 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.039 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.049 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.059 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.069 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.079 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.089 | 25 | 20.205 | 0.012 | 1.2
[6 empty] | 14.099 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.109 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.119 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.129 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.139 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.149 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.159 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.169 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.179 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.189 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.199 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.209 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.219 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.229 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.239 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.249 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.259 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.269 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.279 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.289 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.299 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.309 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.319 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.329 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.339 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.349 | 25 | 20.205 | 0.011 | 1.1
[6 empty] | 14.359 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.369 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.379 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.389 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.399 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.409 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.419 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.429 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.439 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.449 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.459 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.469 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.479 | 25 | 20.206 | 0.011 | 1.1
[6 empty] | 14.489 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.499 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.509 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.519 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.529 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.539 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.549 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.559 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.569 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.579 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.589 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.599 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.609 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.619 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.629 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.639 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.649 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.659 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.669 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.679 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.689 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.699 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.709 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.719 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.729 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.739 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.749 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.759 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.769 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.779 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.789 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.799 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.809 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.819 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.829 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.839 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.849 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.859 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.869 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.879 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.889 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.899 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.909 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.919 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.929 | 25 | 20.206 | 0.01 | 1
[6 empty] | 14.939 | 25 | 20.206 | 0.009 | 0.9
[6 empty] | 14.949 | 25 | 20.206 | 0.009 | 0.9
[6 empty] | 14.959 | 25 | 20.206 | 0.009 | 0.9
[6 empty] | 14.969 | 25 | 20.206 | 0.009 | 0.9
[6 empty] | 14.979 | 25 | 20.206 | 0.009 | 0.9
[6 empty] | 14.989 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 14.999 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.009 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.019 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.029 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.039 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.049 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.059 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.069 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.079 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.089 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.099 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.109 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.119 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.129 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.139 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.149 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.159 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.169 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.179 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.189 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.199 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.209 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.219 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.229 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.239 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.248 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.258 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.268 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.278 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.288 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.298 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.308 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.318 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.328 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.338 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.348 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.358 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.368 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.378 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.388 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.398 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.408 | 25 | 20.207 | 0.009 | 0.9
[6 empty] | 15.418 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.428 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.438 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.448 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.458 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.468 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.478 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.488 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.498 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.508 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.518 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.528 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.538 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.548 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.558 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.568 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.578 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.588 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.598 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.608 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.618 | 25 | 20.207 | 0.008 | 0.8
[6 empty] | 15.628 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.638 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.648 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.658 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.668 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.678 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.688 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.698 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.708 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.718 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.728 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.738 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.748 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.758 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.768 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.778 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.788 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.798 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.808 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.818 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.828 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.838 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.848 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.858 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.868 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.878 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.888 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.898 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.908 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.918 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.928 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.938 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.948 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.958 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.968 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.978 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.988 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 15.998 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.008 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.018 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.028 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.038 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.048 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.058 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.068 | 25 | 20.208 | 0.008 | 0.8
[6 empty] | 16.078 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.088 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.098 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.108 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.118 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.128 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.138 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.148 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.158 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.168 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.178 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.188 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.198 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.208 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.218 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.228 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.238 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.248 | 25 | 20.208 | 0.007 | 0.7
[6 empty] | 16.258 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.268 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.278 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.288 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.298 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.308 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.318 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.328 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.338 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.348 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.358 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.368 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.378 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.388 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.398 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.408 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.418 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.428 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.438 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.448 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.458 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.468 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.478 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.488 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.498 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.508 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.518 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.528 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.538 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.548 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.558 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.568 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.578 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.588 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.598 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.608 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.618 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.628 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.638 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.648 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.658 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.668 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.678 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.688 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.698 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.708 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.718 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.728 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.738 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.748 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.758 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.768 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.778 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.788 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.798 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.808 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.818 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.828 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.838 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.848 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.858 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.868 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.878 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.888 | 25 | 20.209 | 0.007 | 0.7
[6 empty] | 16.898 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.908 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.918 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.928 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.938 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.948 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.958 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.968 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.978 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.988 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 16.998 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.008 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.018 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.028 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.038 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.048 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.058 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.068 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.078 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.088 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.098 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.108 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.118 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.128 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.138 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.148 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.158 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.168 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.178 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.188 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.198 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.208 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.218 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.228 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.238 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.248 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.258 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.268 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.278 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.288 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.298 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.308 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.318 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.328 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.338 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.348 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.358 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.368 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.378 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.388 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.398 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.408 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.418 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.428 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.438 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.448 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.458 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.468 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.478 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.488 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.498 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.508 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.518 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.528 | 25 | 20.21 | 0.006 | 0.6
[6 empty] | 17.538 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.548 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.558 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.568 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.578 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.588 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.598 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.608 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.618 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.628 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.638 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.648 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.658 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.668 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.678 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.688 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.698 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.708 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.718 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.728 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.738 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.748 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.758 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.768 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.778 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.788 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.798 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.808 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.818 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.828 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.838 | 25 | 20.211 | 0.006 | 0.6
[6 empty] | 17.848 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.858 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.868 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.878 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.888 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.898 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.908 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.918 | 25 | 20.211 | 0.005 | 0.5
[6 empty] | 17.928 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.938 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.948 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.958 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.968 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.978 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.988 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 17.998 | 25 | 20.211 | 0.004 | 0.4
[6 empty] | 18.008 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.018 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.028 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.038 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.048 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.058 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.068 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.078 | 25 | 20.211 | 0.003 | 0.3
[6 empty] | 18.088 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.098 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.108 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.118 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.128 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.138 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.148 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.158 | 25 | 20.211 | 0.002 | 0.2
[6 empty] | 18.168 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.178 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.188 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.198 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.208 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.218 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.228 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.238 | 25 | 20.212 | 0.001 | 0.1
[6 empty] | 18.248 | 25 | 20.212 | 0 | 0
[6 empty] | 18.258 | 25 | 20.212 | 0 | 0
[6 empty] | 18.268 | 25 | 20.212 | 0 | 0
[6 empty] | 18.278 | 25 | 20.212 | 0 | 0
[6 empty] | 18.288 | 25 | 20.212 | 0 | 0
[6 empty] | 18.298 | 25 | 20.212 | 0 | 0
[6 empty] | 18.308 | 25 | 20.212 | 0 | 0
[6 empty] | 18.318 | 25 | 20.212 | 0 | 0
[6 empty] | 18.328 | 25 | 20.212 | 0 | 0
[6 empty] | 18.338 | 25 | 20.212 | 0 | 0
[6 empty] | 18.348 | 25 | 20.212 | 0 | 0
[6 empty] | 18.358 | 25 | 20.212 | 0 | 0
[6 empty] | 18.368 | 25 | 20.212 | 0 | 0
[6 empty] | 18.378 | 25 | 20.212 | 0 | 0
[6 empty] | 18.388 | 25 | 20.212 | 0 | 0
[6 empty] | 18.398 | 25 | 20.212 | 0 | 0
[6 empty] | 18.408 | 25 | 20.212 | 0 | 0
[6 empty] | 18.418 | 25 | 20.212 | 0 | 0
[6 empty] | 18.428 | 25 | 20.212 | 0 | 0
[6 empty] | 18.438 | 25 | 20.212 | 0 | 0
[6 empty] | 18.448 | 25 | 20.212 | 0 | 0
[6 empty] | 18.458 | 25 | 20.212 | 0 | 0
[6 empty] | 18.468 | 25 | 20.212 | 0 | 0
[6 empty] | 18.478 | 25 | 20.212 | 0 | 0
[6 empty] | 18.488 | 25 | 20.212 | 0 | 0
[6 empty] | 18.498 | 25 | 20.212 | 0 | 0
[6 empty] | 18.508 | 25 | 20.212 | 0 | 0
[6 empty] | 18.518 | 25 | 20.212 | 0 | 0
[6 empty] | 18.528 | 25 | 20.212 | 0 | 0
[6 empty] | 18.538 | 25 | 20.212 | 0 | 0
[6 empty] | 18.548 | 25 | 20.212 | 0 | 0
[6 empty] | 18.558 | 25 | 20.212 | 0 | 0
[6 empty] | 18.568 | 25 | 20.212 | 0 | 0
[6 empty] | 18.578 | 25 | 20.212 | 0 | 0
[6 empty] | 18.588 | 25 | 20.212 | 0 | 0
[6 empty] | 18.598 | 25 | 20.212 | 0 | 0
[6 empty] | 18.608 | 25 | 20.212 | 0 | 0
[6 empty] | 18.618 | 25 | 20.212 | 0 | 0
[6 empty] | 18.628 | 25 | 20.212 | 0 | 0
[6 empty] | 18.638 | 25 | 20.212 | 0 | 0
[6 empty] | 18.648 | 25 | 20.212 | 0 | 0
[6 empty] | 18.658 | 25 | 20.212 | 0 | 0
[6 empty] | 18.668 | 25 | 20.212 | 0 | 0
[6 empty] | 18.678 | 25 | 20.212 | 0 | 0
[6 empty] | 18.688 | 25 | 20.212 | 0 | 0
[6 empty] | 18.698 | 25 | 20.212 | 0 | 0
[6 empty] | 18.708 | 25 | 20.212 | 0 | 0
[6 empty] | 18.718 | 25 | 20.212 | 0 | 0
[6 empty] | 18.728 | 25 | 20.212 | 0 | 0
[6 empty] | 18.738 | 25 | 20.212 | 0 | 0
[6 empty] | 18.748 | 25 | 20.212 | 0 | 0
[6 empty] | 18.758 | 25 | 20.212 | 0 | 0
[6 empty] | 18.768 | 25 | 20.212 | 0 | 0
[6 empty] | 18.778 | 25 | 20.212 | 0 | 0
[6 empty] | 18.788 | 25 | 20.212 | 0 | 0
[6 empty] | 18.798 | 25 | 20.212 | 0 | 0
[6 empty] | 18.808 | 25 | 20.213 | 0 | 0
[6 empty] | 18.818 | 25 | 20.213 | 0 | 0
[6 empty] | 18.828 | 25 | 20.213 | 0 | 0
[6 empty] | 18.838 | 25 | 20.213 | 0 | 0
[6 empty] | 18.848 | 25 | 20.213 | 0 | 0
[6 empty] | 18.858 | 25 | 20.213 | 0 | 0
[6 empty] | 18.868 | 25 | 20.213 | 0 | 0
[6 empty] | 18.878 | 25 | 20.213 | 0 | 0
[6 empty] | 18.888 | 25 | 20.213 | 0 | 0
[6 empty] | 18.898 | 25 | 20.213 | 0 | 0
[6 empty] | 18.908 | 25 | 20.213 | 0 | 0
[6 empty] | 18.918 | 25 | 20.213 | 0 | 0
[6 empty] | 18.928 | 25 | 20.213 | 0 | 0
[6 empty] | 18.938 | 25 | 20.213 | 0 | 0
[6 empty] | 18.948 | 25 | 20.213 | 0 | 0
[6 empty] | 18.958 | 25 | 20.213 | 0 | 0
[6 empty] | 18.968 | 25 | 20.213 | 0 | 0
[6 empty] | 18.978 | 25 | 20.213 | 0 | 0
[6 empty] | 18.988 | 25 | 20.213 | 0 | 0
[6 empty] | 18.998 | 25 | 20.213 | 0 | 0
[6 empty] | 19.008 | 25 | 20.213 | 0 | 0
[6 empty] | 19.018 | 25 | 20.213 | 0 | 0
[6 empty] | 19.028 | 25 | 20.213 | 0 | 0
[6 empty] | 19.038 | 25 | 20.213 | 0 | 0
[6 empty] | 19.048 | 25 | 20.213 | 0 | 0
[6 empty] | 19.058 | 25 | 20.213 | 0 | 0
[6 empty] | 19.068 | 25 | 20.213 | 0 | 0
[6 empty] | 19.078 | 25 | 20.213 | 0 | 0
[6 empty] | 19.088 | 25 | 20.213 | 0 | 0
[6 empty] | 19.098 | 25 | 20.213 | 0 | 0
[6 empty] | 19.108 | 25 | 20.213 | 0 | 0
[6 empty] | 19.118 | 25 | 20.213 | 0 | 0
[6 empty] | 19.128 | 25 | 20.213 | 0 | 0
[6 empty] | 19.138 | 25 | 20.213 | 0 | 0
[6 empty] | 19.148 | 25 | 20.213 | 0 | 0
[6 empty] | 19.158 | 25 | 20.213 | 0 | 0
[6 empty] | 19.168 | 25 | 20.213 | 0 | 0
[6 empty] | 19.178 | 25 | 20.213 | 0 | 0
[6 empty] | 19.188 | 25 | 20.213 | 0 | 0
[6 empty] | 19.198 | 25 | 20.213 | 0 | 0
[6 empty] | 19.208 | 25 | 20.213 | 0 | 0
[6 empty] | 19.218 | 25 | 20.213 | 0 | 0
[6 empty] | 19.228 | 25 | 20.213 | 0 | 0
[6 empty] | 19.238 | 25 | 20.213 | 0 | 0
[6 empty] | 19.248 | 25 | 20.213 | 0 | 0
[6 empty] | 19.258 | 25 | 20.213 | 0 | 0
[6 empty] | 19.268 | 25 | 20.213 | 0 | 0
[6 empty] | 19.278 | 25 | 20.213 | 0 | 0
[6 empty] | 19.288 | 25 | 20.213 | 0 | 0
[6 empty] | 19.298 | 25 | 20.213 | 0 | 0
[6 empty] | 19.308 | 25 | 20.213 | 0 | 0
[6 empty] | 19.318 | 25 | 20.213 | 0 | 0
[6 empty] | 19.328 | 25 | 20.213 | 0 | 0
[6 empty] | 19.338 | 25 | 20.213 | 0 | 0
[6 empty] | 19.348 | 25 | 20.213 | 0 | 0
[6 empty] | 19.358 | 25 | 20.213 | 0 | 0
[6 empty] | 19.368 | 25 | 20.213 | 0 | 0
[6 empty] | 19.378 | 25 | 20.213 | 0 | 0
[6 empty] | 19.388 | 25 | 20.213 | 0 | 0
[6 empty] | 19.398 | 25 | 20.213 | 0 | 0
[6 empty] | 19.408 | 25 | 20.213 | 0 | 0
[6 empty] | 19.418 | 25 | 20.213 | 0 | 0
[6 empty] | 19.428 | 25 | 20.213 | 0 | 0
[6 empty] | 19.438 | 25 | 20.214 | 0 | 0
[6 empty] | 19.448 | 25 | 20.214 | 0 | 0
[6 empty] | 19.458 | 25 | 20.214 | 0 | 0
[6 empty] | 19.468 | 25 | 20.214 | 0 | 0
[6 empty] | 19.478 | 25 | 20.214 | 0 | 0
[6 empty] | 19.488 | 25 | 20.214 | 0 | 0
[6 empty] | 19.498 | 25 | 20.214 | 0 | 0
[6 empty] | 19.508 | 25 | 20.214 | 0 | 0
[6 empty] | 19.518 | 25 | 20.214 | 0 | 0
[6 empty] | 19.528 | 25 | 20.214 | 0 | 0
[6 empty] | 19.538 | 25 | 20.214 | 0 | 0
[6 empty] | 19.548 | 25 | 20.214 | 0 | 0
[6 empty] | 19.558 | 25 | 20.214 | 0 | 0
[6 empty] | 19.568 | 25 | 20.214 | 0 | 0
[6 empty] | 19.578 | 25 | 20.214 | 0 | 0
[6 empty] | 19.588 | 25 | 20.214 | 0 | 0
[6 empty] | 19.598 | 25 | 20.214 | 0 | 0
[6 empty] | 19.608 | 25 | 20.214 | 0 | 0
[6 empty] | 19.618 | 25 | 20.214 | 0 | 0
[6 empty] | 19.628 | 25 | 20.214 | 0 | 0
[6 empty] | 19.638 | 25 | 20.214 | 0 | 0
[6 empty] | 19.648 | 25 | 20.214 | 0 | 0
[6 empty] | 19.658 | 25 | 20.214 | 0 | 0
[6 empty] | 19.668 | 25 | 20.214 | 0 | 0
[6 empty] | 19.678 | 25 | 20.214 | 0 | 0
[6 empty] | 19.688 | 25 | 20.214 | 0 | 0
[6 empty] | 19.698 | 25 | 20.214 | 0 | 0
[6 empty] | 19.708 | 25 | 20.214 | 0 | 0
[6 empty] | 19.718 | 25 | 20.214 | 0 | 0
[6 empty] | 19.728 | 25 | 20.214 | 0 | 0
[6 empty] | 19.738 | 25 | 20.214 | 0 | 0
[6 empty] | 19.748 | 25 | 20.214 | 0 | 0
[6 empty] | 19.758 | 25 | 20.214 | 0 | 0
[6 empty] | 19.768 | 25 | 20.214 | 0 | 0
[6 empty] | 19.778 | 25 | 20.214 | 0 | 0
[6 empty] | 19.788 | 25 | 20.214 | 0 | 0
[6 empty] | 19.798 | 25 | 20.214 | 0 | 0
[6 empty] | 19.808 | 25 | 20.214 | 0 | 0
[6 empty] | 19.818 | 25 | 20.214 | 0 | 0
[6 empty] | 19.828 | 25 | 20.214 | 0 | 0
[6 empty] | 19.838 | 25 | 20.214 | 0 | 0
[6 empty] | 19.848 | 25 | 20.214 | 0 | 0
[6 empty] | 19.858 | 25 | 20.214 | 0 | 0
[6 empty] | 19.868 | 25 | 20.214 | 0 | 0
[6 empty] | 19.878 | 25 | 20.214 | 0 | 0
[6 empty] | 19.888 | 25 | 20.214 | 0 | 0
[6 empty] | 19.898 | 25 | 20.214 | 0 | 0
[6 empty] | 19.908 | 25 | 20.214 | 0 | 0
[6 empty] | 19.918 | 25 | 20.214 | 0 | 0
[6 empty] | 19.928 | 25 | 20.214 | 0 | 0
[6 empty] | 19.938 | 25 | 20.214 | 0 | 0
[6 empty] | 19.948 | 25 | 20.214 | 0 | 0
[6 empty] | 19.958 | 25 | 20.214 | 0 | 0
[6 empty] | 19.968 | 25 | 20.214 | 0 | 0
[6 empty] | 19.978 | 25 | 20.214 | 0 | 0
[6 empty] | 19.988 | 25 | 20.214 | 0 | 0
[6 empty] | 19.998 | 25 | 20.214 | 0 | 0
[6 empty] | 20.008 | 25 | 20.214 | 0 | 0
[6 empty] | 20.018 | 25 | 20.214 | 0 | 0
[6 empty] | 20.028 | 25 | 20.214 | 0 | 0
[6 empty] | 20.038 | 25 | 20.214 | 0 | 0
[6 empty] | 20.048 | 25 | 20.214 | 0 | 0
[6 empty] | 20.058 | 25 | 20.214 | 0 | 0
[6 empty] | 20.068 | 25 | 20.214 | 0 | 0
[6 empty] | 20.078 | 25 | 20.215 | 0 | 0
[6 empty] | 20.088 | 25 | 20.215 | 0 | 0
[6 empty] | 20.098 | 25 | 20.215 | 0 | 0
[6 empty] | 20.108 | 25 | 20.215 | 0 | 0
[6 empty] | 20.118 | 25 | 20.215 | 0 | 0
[6 empty] | 20.128 | 25 | 20.215 | 0 | 0
[6 empty] | 20.138 | 25 | 20.215 | 0 | 0
[6 empty] | 20.148 | 25 | 20.215 | 0 | 0
[6 empty] | 20.158 | 25 | 20.215 | 0 | 0
[6 empty] | 20.168 | 25 | 20.215 | 0 | 0
[6 empty] | 20.178 | 25 | 20.215 | 0 | 0
[6 empty] | 20.188 | 25 | 20.215 | 0 | 0
[6 empty] | 20.198 | 25 | 20.215 | 0 | 0
[6 empty] | 20.208 | 25 | 20.215 | 0 | 0
[6 empty] | 20.218 | 25 | 20.215 | 0 | 0
[6 empty] | 20.228 | 25 | 20.215 | 0 | 0
[6 empty] | 20.238 | 25 | 20.215 | 0 | 0
[6 empty] | 20.247 | 25 | 20.215 | 0 | 0
[6 empty] | 20.257 | 25 | 20.215 | 0 | 0
[6 empty] | 20.267 | 25 | 20.215 | 0 | 0
[6 empty] | 20.277 | 25 | 20.215 | 0 | 0
[6 empty] | 20.287 | 25 | 20.215 | 0 | 0
[6 empty] | 20.297 | 25 | 20.215 | 0 | 0
[6 empty] | 20.307 | 25 | 20.215 | 0 | 0
[6 empty] | 20.317 | 25 | 20.215 | 0 | 0
[6 empty] | 20.327 | 25 | 20.215 | 0 | 0
[6 empty] | 20.337 | 25 | 20.215 | 0 | 0
[6 empty] | 20.347 | 25 | 20.215 | 0 | 0
[6 empty] | 20.357 | 25 | 20.215 | 0 | 0
[6 empty] | 20.367 | 25 | 20.215 | 0 | 0
[6 empty] | 20.377 | 25 | 20.215 | 0 | 0
[6 empty] | 20.387 | 25 | 20.215 | 0 | 0
[6 empty] | 20.397 | 25 | 20.215 | 0 | 0
[6 empty] | 20.407 | 25 | 20.215 | 0 | 0
[6 empty] | 20.417 | 25 | 20.215 | 0 | 0
[6 empty] | 20.427 | 25 | 20.215 | 0 | 0
[6 empty] | 20.437 | 25 | 20.215 | 0 | 0
[6 empty] | 20.447 | 25 | 20.215 | 0 | 0
[6 empty] | 20.457 | 25 | 20.215 | 0 | 0
[6 empty] | 20.467 | 25 | 20.215 | 0 | 0
[6 empty] | 20.477 | 25 | 20.215 | 0 | 0
[6 empty] | 20.487 | 25 | 20.215 | 0 | 0
[6 empty] | 20.497 | 25 | 20.215 | 0 | 0
[6 empty] | 20.507 | 25 | 20.215 | 0 | 0
[6 empty] | 20.517 | 25 | 20.215 | 0 | 0
[6 empty] | 20.527 | 25 | 20.215 | 0 | 0
[6 empty] | 20.537 | 25 | 20.215 | 0 | 0
[6 empty] | 20.547 | 25 | 20.215 | 0 | 0
[6 empty] | 20.557 | 25 | 20.215 | 0 | 0
[6 empty] | 20.567 | 25 | 20.215 | 0 | 0
[6 empty] | 20.577 | 25 | 20.215 | 0 | 0
[6 empty] | 20.587 | 25 | 20.215 | 0 | 0
[6 empty] | 20.597 | 25 | 20.215 | 0 | 0
[6 empty] | 20.607 | 25 | 20.215 | 0 | 0
[6 empty] | 20.617 | 25 | 20.215 | 0 | 0
[6 empty] | 20.627 | 25 | 20.215 | 0 | 0
[6 empty] | 20.637 | 25 | 20.215 | 0 | 0
[6 empty] | 20.647 | 25 | 20.215 | 0 | 0
[6 empty] | 20.657 | 25 | 20.215 | 0 | 0
[6 empty] | 20.667 | 25 | 20.215 | 0 | 0
[6 empty] | 20.677 | 25 | 20.215 | 0 | 0
[6 empty] | 20.687 | 25 | 20.215 | 0 | 0
[6 empty] | 20.697 | 25 | 20.215 | 0 | 0
[6 empty] | 20.707 | 25 | 20.215 | 0 | 0
[6 empty] | 20.717 | 25 | 20.216 | 0 | 0
[6 empty] | 20.727 | 25 | 20.216 | 0 | 0
[6 empty] | 20.737 | 25 | 20.216 | 0 | 0
[6 empty] | 20.747 | 25 | 20.216 | 0 | 0
[6 empty] | 20.757 | 25 | 20.216 | 0 | 0
[6 empty] | 20.767 | 25 | 20.216 | 0 | 0
[6 empty] | 20.777 | 25 | 20.216 | 0 | 0
[6 empty] | 20.787 | 25 | 20.216 | 0 | 0
[6 empty] | 20.797 | 25 | 20.216 | 0 | 0
[6 empty] | 20.807 | 25 | 20.216 | 0 | 0
[6 empty] | 20.817 | 25 | 20.216 | 0 | 0
[6 empty] | 20.827 | 25 | 20.216 | 0 | 0
[6 empty] | 20.837 | 25 | 20.216 | 0 | 0
[6 empty] | 20.847 | 25 | 20.216 | 0 | 0
[6 empty] | 20.857 | 25 | 20.216 | 0 | 0
[6 empty] | 20.867 | 25 | 20.216 | 0 | 0
[6 empty] | 20.877 | 25 | 20.216 | 0 | 0
[6 empty] | 20.887 | 25 | 20.216 | 0 | 0
[6 empty] | 20.897 | 25 | 20.216 | 0 | 0
[6 empty] | 20.907 | 25 | 20.216 | 0 | 0
[6 empty] | 20.917 | 25 | 20.216 | 0 | 0
[6 empty] | 20.927 | 25 | 20.216 | 0 | 0
[6 empty] | 20.937 | 25 | 20.216 | 0 | 0
[6 empty] | 20.947 | 25 | 20.216 | 0 | 0
[6 empty] | 20.957 | 25 | 20.216 | 0 | 0
[6 empty] | 20.967 | 25 | 20.216 | 0 | 0
[6 empty] | 20.977 | 25 | 20.216 | 0 | 0
[6 empty] | 20.987 | 25 | 20.216 | 0 | 0
[6 empty] | 20.997 | 25 | 20.216 | 0 | 0
[6 empty] | 21.007 | 25 | 20.216 | 0 | 0
[6 empty] | 21.017 | 25 | 20.216 | 0 | 0
[6 empty] | 21.027 | 25 | 20.216 | 0 | 0
[6 empty] | 21.037 | 25 | 20.216 | 0 | 0
[6 empty] | 21.047 | 25 | 20.216 | 0 | 0
[6 empty] | 21.057 | 25 | 20.216 | 0 | 0
[6 empty] | 21.067 | 25 | 20.216 | 0 | 0
[6 empty] | 21.077 | 25 | 20.216 | 0 | 0
[6 empty] | 21.087 | 25 | 20.216 | 0 | 0
[6 empty] | 21.097 | 25 | 20.216 | 0 | 0
[6 empty] | 21.107 | 25 | 20.216 | 0 | 0
[6 empty] | 21.117 | 25 | 20.216 | 0 | 0
[6 empty] | 21.127 | 25 | 20.216 | 0 | 0
[6 empty] | 21.137 | 25 | 20.216 | 0 | 0
[6 empty] | 21.147 | 25 | 20.216 | 0 | 0
[6 empty] | 21.157 | 25 | 20.216 | 0 | 0
[6 empty] | 21.167 | 25 | 20.216 | 0 | 0
[6 empty] | 21.177 | 25 | 20.216 | 0 | 0
[6 empty] | 21.187 | 25 | 20.216 | 0 | 0
[6 empty] | 21.197 | 25 | 20.216 | 0 | 0
[6 empty] | 21.207 | 25 | 20.216 | 0 | 0
[6 empty] | 21.217 | 25 | 20.216 | 0 | 0
[6 empty] | 21.227 | 25 | 20.216 | 0 | 0
[6 empty] | 21.237 | 25 | 20.216 | 0 | 0
[6 empty] | 21.247 | 25 | 20.216 | 0 | 0
[6 empty] | 21.257 | 25 | 20.216 | 0 | 0
[6 empty] | 21.267 | 25 | 20.216 | 0 | 0
[6 empty] | 21.277 | 25 | 20.216 | 0 | 0
[6 empty] | 21.287 | 25 | 20.216 | 0 | 0
[6 empty] | 21.297 | 25 | 20.216 | 0 | 0
[6 empty] | 21.307 | 25 | 20.216 | 0 | 0
[6 empty] | 21.317 | 25 | 20.216 | 0 | 0
[6 empty] | 21.327 | 25 | 20.216 | 0 | 0
[6 empty] | 21.337 | 25 | 20.216 | 0 | 0
[6 empty] | 21.347 | 25 | 20.217 | 0 | 0
[6 empty] | 21.357 | 25 | 20.217 | 0 | 0
[6 empty] | 21.367 | 25 | 20.217 | 0 | 0
[6 empty] | 21.377 | 25 | 20.217 | 0 | 0
[6 empty] | 21.387 | 25 | 20.217 | 0 | 0
[6 empty] | 21.397 | 25 | 20.217 | 0 | 0
[6 empty] | 21.407 | 25 | 20.217 | 0 | 0
[6 empty] | 21.417 | 25 | 20.217 | 0 | 0
[6 empty] | 21.427 | 25 | 20.217 | 0 | 0
[6 empty] | 21.437 | 25 | 20.217 | 0 | 0
[6 empty] | 21.447 | 25 | 20.217 | 0 | 0
[6 empty] | 21.457 | 25 | 20.217 | 0 | 0
[6 empty] | 21.467 | 25 | 20.217 | 0 | 0
[6 empty] | 21.477 | 25 | 20.217 | 0 | 0
[6 empty] | 21.487 | 25 | 20.217 | 0 | 0
[6 empty] | 21.497 | 25 | 20.217 | 0 | 0
[6 empty] | 21.507 | 25 | 20.217 | 0 | 0
[6 empty] | 21.517 | 25 | 20.217 | 0 | 0
[6 empty] | 21.527 | 25 | 20.217 | 0 | 0
[6 empty] | 21.537 | 25 | 20.217 | 0 | 0
[6 empty] | 21.547 | 25 | 20.217 | 0 | 0
[6 empty] | 21.557 | 25 | 20.217 | 0 | 0
[6 empty] | 21.567 | 25 | 20.217 | 0 | 0
[6 empty] | 21.577 | 25 | 20.217 | 0 | 0
[6 empty] | 21.587 | 25 | 20.217 | 0 | 0
[6 empty] | 21.597 | 25 | 20.217 | 0 | 0
[6 empty] | 21.607 | 25 | 20.217 | 0 | 0
[6 empty] | 21.617 | 25 | 20.217 | 0 | 0
[6 empty] | 21.627 | 25 | 20.217 | 0 | 0
[6 empty] | 21.637 | 25 | 20.217 | 0 | 0
[6 empty] | 21.647 | 25 | 20.217 | 0 | 0
[6 empty] | 21.657 | 25 | 20.217 | 0 | 0
[6 empty] | 21.667 | 25 | 20.217 | 0 | 0
[6 empty] | 21.677 | 25 | 20.217 | 0 | 0
[6 empty] | 21.687 | 25 | 20.217 | 0 | 0
[6 empty] | 21.697 | 25 | 20.217 | 0 | 0
[6 empty] | 21.707 | 25 | 20.217 | 0 | 0
[6 empty] | 21.717 | 25 | 20.217 | 0 | 0
[6 empty] | 21.727 | 25 | 20.217 | 0 | 0
[6 empty] | 21.737 | 25 | 20.217 | 0 | 0
[6 empty] | 21.747 | 25 | 20.217 | 0 | 0
[6 empty] | 21.757 | 25 | 20.217 | 0 | 0
[6 empty] | 21.767 | 25 | 20.217 | 0 | 0
[6 empty] | 21.777 | 25 | 20.217 | 0 | 0
[6 empty] | 21.787 | 25 | 20.217 | 0 | 0
[6 empty] | 21.797 | 25 | 20.217 | 0 | 0
[6 empty] | 21.807 | 25 | 20.217 | 0 | 0
[6 empty] | 21.817 | 25 | 20.217 | 0 | 0
[6 empty] | 21.827 | 25 | 20.217 | 0 | 0
[6 empty] | 21.837 | 25 | 20.217 | 0 | 0
[6 empty] | 21.847 | 25 | 20.217 | 0 | 0
[6 empty] | 21.857 | 25 | 20.217 | 0 | 0
[6 empty] | 21.867 | 25 | 20.217 | 0 | 0
[6 empty] | 21.877 | 25 | 20.217 | 0 | 0
[6 empty] | 21.887 | 25 | 20.217 | 0 | 0
[6 empty] | 21.897 | 25 | 20.217 | 0 | 0
[6 empty] | 21.907 | 25 | 20.217 | 0 | 0
[6 empty] | 21.917 | 25 | 20.217 | 0 | 0
[6 empty] | 21.927 | 25 | 20.217 | 0 | 0
[6 empty] | 21.937 | 25 | 20.217 | 0 | 0
[6 empty] | 21.947 | 25 | 20.217 | 0 | 0
[6 empty] | 21.957 | 25 | 20.217 | 0 | 0
[6 empty] | 21.967 | 25 | 20.217 | 0 | 0
[6 empty] | 21.977 | 25 | 20.217 | 0 | 0
[6 empty] | 21.987 | 25 | 20.218 | 0 | 0
[6 empty] | 21.997 | 25 | 20.218 | 0 | 0
[6 empty] | 22.007 | 25 | 20.218 | 0 | 0
[6 empty] | 22.017 | 25 | 20.218 | 0 | 0
[6 empty] | 22.027 | 25 | 20.218 | 0 | 0
[6 empty] | 22.037 | 25 | 20.218 | 0 | 0
[6 empty] | 22.047 | 25 | 20.218 | 0 | 0
[6 empty] | 22.057 | 25 | 20.218 | 0 | 0
[6 empty] | 22.067 | 25 | 20.218 | 0 | 0
[6 empty] | 22.077 | 25 | 20.218 | 0 | 0
[6 empty] | 22.087 | 25 | 20.218 | 0 | 0
[6 empty] | 22.097 | 25 | 20.218 | 0 | 0
[6 empty] | 22.107 | 25 | 20.218 | 0 | 0
[6 empty] | 22.117 | 25 | 20.218 | 0 | 0
[6 empty] | 22.127 | 25 | 20.218 | 0 | 0
[6 empty] | 22.137 | 25 | 20.218 | 0 | 0
[6 empty] | 22.147 | 25 | 20.218 | 0 | 0
[6 empty] | 22.157 | 25 | 20.218 | 0 | 0
[6 empty] | 22.167 | 25 | 20.218 | 0 | 0
[6 empty] | 22.177 | 25 | 20.218 | 0 | 0
[6 empty] | 22.187 | 25 | 20.218 | 0 | 0
[6 empty] | 22.197 | 25 | 20.218 | 0 | 0
[6 empty] | 22.207 | 25 | 20.218 | 0 | 0
[6 empty] | 22.217 | 25 | 20.218 | 0 | 0
[6 empty] | 22.227 | 25 | 20.218 | 0 | 0
[6 empty] | 22.237 | 25 | 20.218 | 0 | 0
[6 empty] | 22.247 | 25 | 20.218 | 0 | 0
[6 empty] | 22.257 | 25 | 20.218 | 0 | 0
[6 empty] | 22.267 | 25 | 20.218 | 0 | 0
[6 empty] | 22.277 | 25 | 20.218 | 0 | 0
[6 empty] | 22.287 | 25 | 20.218 | 0 | 0
[6 empty] | 22.297 | 25 | 20.218 | 0 | 0
[6 empty] | 22.307 | 25 | 20.218 | 0 | 0
[6 empty] | 22.317 | 25 | 20.218 | 0 | 0
[6 empty] | 22.327 | 25 | 20.218 | 0 | 0
[6 empty] | 22.337 | 25 | 20.218 | 0 | 0
[6 empty] | 22.347 | 25 | 20.218 | 0 | 0
[6 empty] | 22.357 | 25 | 20.218 | 0 | 0
[6 empty] | 22.367 | 25 | 20.218 | 0 | 0
[6 empty] | 22.377 | 25 | 20.218 | 0 | 0
[6 empty] | 22.387 | 25 | 20.218 | 0 | 0
[6 empty] | 22.397 | 25 | 20.218 | 0 | 0
[6 empty] | 22.407 | 25 | 20.218 | 0 | 0
[6 empty] | 22.417 | 25 | 20.218 | 0 | 0
[6 empty] | 22.427 | 25 | 20.218 | 0 | 0
[6 empty] | 22.437 | 25 | 20.218 | 0 | 0
[6 empty] | 22.447 | 25 | 20.218 | 0 | 0
[6 empty] | 22.457 | 25 | 20.218 | 0 | 0
[6 empty] | 22.467 | 25 | 20.218 | 0 | 0
[6 empty] | 22.477 | 25 | 20.218 | 0 | 0
[6 empty] | 22.487 | 25 | 20.218 | 0 | 0
[6 empty] | 22.497 | 25 | 20.218 | 0 | 0
[6 empty] | 22.507 | 25 | 20.218 | 0 | 0
[6 empty] | 22.517 | 25 | 20.218 | 0 | 0
[6 empty] | 22.527 | 25 | 20.218 | 0 | 0
[6 empty] | 22.537 | 25 | 20.218 | 0 | 0
[6 empty] | 22.547 | 25 | 20.218 | 0 | 0
[6 empty] | 22.557 | 25 | 20.218 | 0 | 0
[6 empty] | 22.567 | 25 | 20.218 | 0 | 0
[6 empty] | 22.577 | 25 | 20.218 | 0 | 0
[6 empty] | 22.587 | 25 | 20.218 | 0 | 0
[6 empty] | 22.597 | 25 | 20.218 | 0 | 0
[6 empty] | 22.607 | 25 | 20.218 | 0 | 0
[6 empty] | 22.617 | 25 | 20.218 | 0 | 0
[6 empty] | 22.627 | 25 | 20.219 | 0 | 0
[6 empty] | 22.637 | 25 | 20.219 | 0 | 0
[6 empty] | 22.647 | 25 | 20.219 | 0 | 0
[6 empty] | 22.657 | 25 | 20.219 | 0 | 0
[6 empty] | 22.667 | 25 | 20.219 | 0 | 0
[6 empty] | 22.677 | 25 | 20.219 | 0 | 0
[6 empty] | 22.687 | 25 | 20.219 | 0 | 0
[6 empty] | 22.697 | 25 | 20.219 | 0 | 0
[6 empty] | 22.707 | 25 | 20.219 | 0 | 0
[6 empty] | 22.717 | 25 | 20.219 | 0 | 0
[6 empty] | 22.727 | 25 | 20.219 | 0 | 0
[6 empty] | 22.737 | 25 | 20.219 | 0 | 0
[6 empty] | 22.747 | 25 | 20.219 | 0 | 0
[6 empty] | 22.757 | 25 | 20.219 | 0 | 0
[6 empty] | 22.767 | 25 | 20.219 | 0 | 0
[6 empty] | 22.777 | 25 | 20.219 | 0 | 0
[6 empty] | 22.787 | 25 | 20.219 | 0 | 0
[6 empty] | 22.797 | 25 | 20.219 | 0 | 0
[6 empty] | 22.807 | 25 | 20.219 | 0 | 0
[6 empty] | 22.817 | 25 | 20.219 | 0 | 0
[6 empty] | 22.827 | 25 | 20.219 | 0 | 0
[6 empty] | 22.837 | 25 | 20.219 | 0 | 0
[6 empty] | 22.847 | 25 | 20.219 | 0 | 0
[6 empty] | 22.857 | 25 | 20.219 | 0 | 0
[6 empty] | 22.867 | 25 | 20.219 | 0 | 0
[6 empty] | 22.877 | 25 | 20.219 | 0 | 0
[6 empty] | 22.887 | 25 | 20.219 | 0 | 0
[6 empty] | 22.897 | 25 | 20.219 | 0 | 0
[6 empty] | 22.907 | 25 | 20.219 | 0 | 0
[6 empty] | 22.917 | 25 | 20.219 | 0 | 0
[6 empty] | 22.927 | 25 | 20.219 | 0 | 0
[6 empty] | 22.937 | 25 | 20.219 | 0 | 0
[6 empty] | 22.947 | 25 | 20.219 | 0 | 0
[6 empty] | 22.957 | 25 | 20.219 | 0 | 0
[6 empty] | 22.967 | 25 | 20.219 | 0 | 0
[6 empty] | 22.977 | 25 | 20.219 | 0 | 0
[6 empty] | 22.987 | 25 | 20.219 | 0 | 0
[6 empty] | 22.997 | 25 | 20.219 | 0 | 0
[6 empty] | 23.007 | 25 | 20.219 | 0 | 0
[6 empty] | 23.017 | 25 | 20.219 | 0 | 0
[6 empty] | 23.027 | 25 | 20.219 | 0 | 0
[6 empty] | 23.037 | 25 | 20.219 | 0 | 0
[6 empty] | 23.047 | 25 | 20.219 | 0 | 0
[6 empty] | 23.057 | 25 | 20.219 | 0 | 0
[6 empty] | 23.067 | 25 | 20.219 | 0 | 0
[6 empty] | 23.077 | 25 | 20.219 | 0 | 0
[6 empty] | 23.087 | 25 | 20.219 | 0 | 0
[6 empty] | 23.097 | 25 | 20.219 | 0 | 0
[6 empty] | 23.107 | 25 | 20.219 | 0 | 0
[6 empty] | 23.117 | 25 | 20.219 | 0 | 0
[6 empty] | 23.127 | 25 | 20.219 | 0 | 0
[6 empty] | 23.137 | 25 | 20.219 | 0 | 0
[6 empty] | 23.147 | 25 | 20.219 | 0 | 0
[6 empty] | 23.157 | 25 | 20.219 | 0 | 0
[6 empty] | 23.167 | 25 | 20.219 | 0 | 0
[6 empty] | 23.177 | 25 | 20.219 | 0 | 0
[6 empty] | 23.187 | 25 | 20.219 | 0 | 0
[6 empty] | 23.197 | 25 | 20.219 | 0 | 0
[6 empty] | 23.207 | 25 | 20.219 | 0 | 0
[6 empty] | 23.217 | 25 | 20.219 | 0 | 0
[6 empty] | 23.227 | 25 | 20.219 | 0 | 0
[6 empty] | 23.237 | 25 | 20.219 | 0 | 0
[6 empty] | 23.247 | 25 | 20.219 | 0 | 0
[6 empty] | 23.257 | 25 | 20.22 | 0 | 0
[6 empty] | 23.267 | 25 | 20.22 | 0 | 0
[6 empty] | 23.277 | 25 | 20.22 | 0 | 0
[6 empty] | 23.287 | 25 | 20.22 | 0 | 0
[6 empty] | 23.297 | 25 | 20.22 | 0 | 0
[6 empty] | 23.307 | 25 | 20.22 | 0 | 0
[6 empty] | 23.317 | 25 | 20.22 | 0 | 0
[6 empty] | 23.327 | 25 | 20.22 | 0 | 0
[6 empty] | 23.337 | 25 | 20.22 | 0 | 0
[6 empty] | 23.347 | 25 | 20.22 | 0 | 0
[6 empty] | 23.357 | 25 | 20.22 | 0 | 0
[6 empty] | 23.367 | 25 | 20.22 | 0 | 0
[6 empty] | 23.377 | 25 | 20.22 | 0 | 0
[6 empty] | 23.387 | 25 | 20.22 | 0 | 0
[6 empty] | 23.397 | 25 | 20.22 | 0 | 0
[6 empty] | 23.407 | 25 | 20.22 | 0 | 0
[6 empty] | 23.417 | 25 | 20.22 | 0 | 0
[6 empty] | 23.427 | 25 | 20.22 | 0 | 0
[6 empty] | 23.437 | 25 | 20.22 | 0 | 0
[6 empty] | 23.447 | 25 | 20.22 | 0 | 0
[6 empty] | 23.457 | 25 | 20.22 | 0 | 0
[6 empty] | 23.467 | 25 | 20.22 | 0 | 0
[6 empty] | 23.477 | 25 | 20.22 | 0 | 0
[6 empty] | 23.487 | 25 | 20.22 | 0 | 0
[6 empty] | 23.497 | 25 | 20.22 | 0 | 0
[6 empty] | 23.507 | 25 | 20.22 | 0 | 0
[6 empty] | 23.517 | 25 | 20.22 | 0 | 0
[6 empty] | 23.527 | 25 | 20.22 | 0 | 0
[6 empty] | 23.537 | 25 | 20.22 | 0 | 0
[6 empty] | 23.547 | 25 | 20.22 | 0 | 0
[6 empty] | 23.557 | 25 | 20.22 | 0 | 0
[6 empty] | 23.567 | 25 | 20.22 | 0 | 0
[6 empty] | 23.577 | 25 | 20.22 | 0 | 0
[6 empty] | 23.587 | 25 | 20.22 | 0 | 0
[6 empty] | 23.597 | 25 | 20.22 | 0 | 0
[6 empty] | 23.607 | 25 | 20.22 | 0 | 0
[6 empty] | 23.617 | 25 | 20.22 | 0 | 0
[6 empty] | 23.627 | 25 | 20.22 | 0 | 0
[6 empty] | 23.637 | 25 | 20.22 | 0 | 0
[6 empty] | 23.647 | 25 | 20.22 | 0 | 0
[6 empty] | 23.657 | 25 | 20.22 | 0 | 0
[6 empty] | 23.667 | 25 | 20.22 | 0 | 0
[6 empty] | 23.677 | 25 | 20.22 | 0 | 0
[6 empty] | 23.687 | 25 | 20.22 | 0 | 0
[6 empty] | 23.697 | 25 | 20.22 | 0 | 0
[6 empty] | 23.707 | 25 | 20.22 | 0 | 0
[6 empty] | 23.717 | 25 | 20.22 | 0 | 0
[6 empty] | 23.727 | 25 | 20.22 | 0 | 0
[6 empty] | 23.737 | 25 | 20.22 | 0 | 0
[6 empty] | 23.747 | 25 | 20.22 | 0 | 0
[6 empty] | 23.757 | 25 | 20.22 | 0 | 0
[6 empty] | 23.767 | 25 | 20.22 | 0 | 0
[6 empty] | 23.777 | 25 | 20.22 | 0 | 0
[6 empty] | 23.787 | 25 | 20.22 | 0 | 0
[6 empty] | 23.797 | 25 | 20.22 | 0 | 0
[6 empty] | 23.807 | 25 | 20.22 | 0 | 0
[6 empty] | 23.817 | 25 | 20.22 | 0 | 0
[6 empty] | 23.827 | 25 | 20.22 | 0 | 0
[6 empty] | 23.837 | 25 | 20.22 | 0 | 0
[6 empty] | 23.847 | 25 | 20.22 | 0 | 0
[6 empty] | 23.857 | 25 | 20.22 | 0 | 0
[6 empty] | 23.867 | 25 | 20.22 | 0 | 0
[6 empty] | 23.877 | 25 | 20.22 | 0 | 0
[6 empty] | 23.887 | 25 | 20.22 | 0 | 0
[6 empty] | 23.897 | 25 | 20.221 | 0 | 0
[6 empty] | 23.907 | 25 | 20.221 | 0 | 0
[6 empty] | 23.917 | 25 | 20.221 | 0 | 0
[6 empty] | 23.927 | 25 | 20.221 | 0 | 0
[6 empty] | 23.937 | 25 | 20.221 | 0 | 0
[6 empty] | 23.947 | 25 | 20.221 | 0 | 0
[6 empty] | 23.957 | 25 | 20.221 | 0 | 0
[6 empty] | 23.967 | 25 | 20.221 | 0 | 0
[6 empty] | 23.977 | 25 | 20.221 | 0 | 0
[6 empty] | 23.987 | 25 | 20.221 | 0 | 0
[6 empty] | 23.997 | 25 | 20.221 | 0 | 0
[6 empty] | 24.007 | 25 | 20.221 | 0 | 0
[6 empty] | 24.017 | 25 | 20.221 | 0 | 0
[6 empty] | 24.027 | 25 | 20.221 | 0 | 0
[6 empty] | 24.037 | 25 | 20.221 | 0 | 0
[6 empty] | 24.047 | 25 | 20.221 | 0 | 0
[6 empty] | 24.057 | 25 | 20.221 | 0 | 0
[6 empty] | 24.067 | 25 | 20.221 | 0 | 0
[6 empty] | 24.077 | 25 | 20.221 | 0 | 0
[6 empty] | 24.087 | 25 | 20.221 | 0 | 0
[6 empty] | 24.097 | 25 | 20.221 | 0 | 0
[6 empty] | 24.107 | 25 | 20.221 | 0 | 0
[6 empty] | 24.117 | 25 | 20.221 | 0 | 0
[6 empty] | 24.127 | 25 | 20.221 | 0 | 0
[6 empty] | 24.137 | 25 | 20.221 | 0 | 0
[6 empty] | 24.147 | 25 | 20.221 | 0 | 0
[6 empty] | 24.157 | 25 | 20.221 | 0 | 0
[6 empty] | 24.167 | 25 | 20.221 | 0 | 0
[6 empty] | 24.177 | 25 | 20.221 | 0 | 0
[6 empty] | 24.187 | 25 | 20.221 | 0 | 0
[6 empty] | 24.197 | 25 | 20.221 | 0 | 0
[6 empty] | 24.207 | 25 | 20.221 | 0 | 0
[6 empty] | 24.217 | 25 | 20.221 | 0 | 0
[6 empty] | 24.227 | 25 | 20.221 | 0 | 0
[6 empty] | 24.237 | 25 | 20.221 | 0 | 0
[6 empty] | 24.247 | 25 | 20.221 | 0 | 0
[6 empty] | 24.257 | 25 | 20.221 | 0 | 0
[6 empty] | 24.267 | 25 | 20.221 | 0 | 0
[6 empty] | 24.277 | 25 | 20.221 | 0 | 0
[6 empty] | 24.287 | 25 | 20.221 | 0 | 0
[6 empty] | 24.297 | 25 | 20.221 | 0 | 0
[6 empty] | 24.307 | 25 | 20.221 | 0 | 0
[6 empty] | 24.317 | 25 | 20.221 | 0 | 0
[6 empty] | 24.327 | 25 | 20.221 | 0 | 0
[6 empty] | 24.337 | 25 | 20.221 | 0 | 0
[6 empty] | 24.347 | 25 | 20.221 | 0 | 0
[6 empty] | 24.357 | 25 | 20.221 | 0 | 0
[6 empty] | 24.367 | 25 | 20.221 | 0 | 0
[6 empty] | 24.377 | 25 | 20.221 | 0 | 0
[6 empty] | 24.387 | 25 | 20.221 | 0 | 0
[6 empty] | 24.397 | 25 | 20.221 | 0 | 0
[6 empty] | 24.407 | 25 | 20.221 | 0 | 0
[6 empty] | 24.417 | 25 | 20.221 | 0 | 0
[6 empty] | 24.427 | 25 | 20.221 | 0 | 0
[6 empty] | 24.437 | 25 | 20.221 | 0 | 0
[6 empty] | 24.447 | 25 | 20.221 | 0 | 0
[6 empty] | 24.457 | 25 | 20.221 | 0 | 0
[6 empty] | 24.467 | 25 | 20.221 | 0 | 0
[6 empty] | 24.477 | 25 | 20.221 | 0 | 0
[6 empty] | 24.487 | 25 | 20.221 | 0 | 0
[6 empty] | 24.497 | 25 | 20.221 | 0 | 0
[6 empty] | 24.507 | 25 | 20.221 | 0 | 0
[6 empty] | 24.517 | 25 | 20.221 | 0 | 0
[6 empty] | 24.527 | 25 | 20.222 | 0 | 0
[6 empty] | 24.537 | 25 | 20.222 | 0 | 0
[6 empty] | 24.547 | 25 | 20.222 | 0 | 0
[6 empty] | 24.557 | 25 | 20.222 | 0 | 0
[6 empty] | 24.567 | 25 | 20.222 | 0 | 0
[6 empty] | 24.577 | 25 | 20.222 | 0 | 0
[6 empty] | 24.587 | 25 | 20.222 | 0 | 0
[6 empty] | 24.597 | 25 | 20.222 | 0 | 0
[6 empty] | 24.607 | 25 | 20.222 | 0 | 0
[6 empty] | 24.617 | 25 | 20.222 | 0 | 0
[6 empty] | 24.627 | 25 | 20.222 | 0 | 0
[6 empty] | 24.637 | 25 | 20.222 | 0 | 0
[6 empty] | 24.647 | 25 | 20.222 | 0 | 0
[6 empty] | 24.657 | 25 | 20.222 | 0 | 0
[6 empty] | 24.667 | 25 | 20.222 | 0 | 0
[6 empty] | 24.677 | 25 | 20.222 | 0 | 0
[6 empty] | 24.687 | 25 | 20.222 | 0 | 0
[6 empty] | 24.697 | 25 | 20.222 | 0 | 0
[6 empty] | 24.707 | 25 | 20.222 | 0 | 0
[6 empty] | 24.717 | 25 | 20.222 | 0 | 0
[6 empty] | 24.727 | 25 | 20.222 | 0 | 0
[6 empty] | 24.737 | 25 | 20.222 | 0 | 0
[6 empty] | 24.747 | 25 | 20.222 | 0 | 0
[6 empty] | 24.757 | 25 | 20.222 | 0 | 0
[6 empty] | 24.767 | 25 | 20.222 | 0 | 0
[6 empty] | 24.777 | 25 | 20.222 | 0 | 0
[6 empty] | 24.787 | 25 | 20.222 | 0 | 0
[6 empty] | 24.797 | 25 | 20.222 | 0 | 0
[6 empty] | 24.807 | 25 | 20.222 | 0 | 0
[6 empty] | 24.817 | 25 | 20.222 | 0 | 0
[6 empty] | 24.827 | 25 | 20.222 | 0 | 0
[6 empty] | 24.837 | 25 | 20.222 | 0 | 0
[6 empty] | 24.847 | 25 | 20.222 | 0 | 0
[6 empty] | 24.857 | 25 | 20.222 | 0 | 0
[6 empty] | 24.867 | 25 | 20.222 | 0 | 0
[6 empty] | 24.877 | 25 | 20.222 | 0 | 0
[6 empty] | 24.887 | 25 | 20.222 | 0 | 0
[6 empty] | 24.897 | 25 | 20.222 | 0 | 0
[6 empty] | 24.907 | 25 | 20.222 | 0 | 0
[6 empty] | 24.917 | 25 | 20.222 | 0 | 0
[6 empty] | 24.927 | 25 | 20.222 | 0 | 0
[6 empty] | 24.937 | 25 | 20.222 | 0 | 0
[6 empty] | 24.947 | 25 | 20.222 | 0 | 0
[6 empty] | 24.957 | 25 | 20.222 | 0 | 0
[6 empty] | 24.967 | 25 | 20.222 | 0 | 0
[6 empty] | 24.977 | 25 | 20.222 | 0 | 0
[6 empty] | 24.987 | 25 | 20.222 | 0 | 0
[6 empty] | 24.997 | 25 | 20.222 | 0 | 0
[6 empty] | 25.007 | 25 | 20.222 | 0 | 0
[6 empty] | 25.017 | 25 | 20.222 | 0 | 0
[6 empty] | 25.027 | 25 | 20.222 | 0 | 0
[6 empty] | 25.037 | 25 | 20.222 | 0 | 0
[6 empty] | 25.047 | 25 | 20.222 | 0 | 0
[6 empty] | 25.057 | 25 | 20.222 | 0 | 0
[6 empty] | 25.067 | 25 | 20.222 | 0 | 0
[6 empty] | 25.077 | 25 | 20.222 | 0 | 0
[6 empty] | 25.087 | 25 | 20.222 | 0 | 0
[6 empty] | 25.097 | 25 | 20.222 | 0 | 0
[6 empty] | 25.107 | 25 | 20.222 | 0 | 0
[6 empty] | 25.117 | 25 | 20.222 | 0 | 0
[6 empty] | 25.127 | 25 | 20.222 | 0 | 0
[6 empty] | 25.137 | 25 | 20.222 | 0 | 0
[6 empty] | 25.147 | 25 | 20.222 | 0 | 0
[6 empty] | 25.157 | 25 | 20.222 | 0 | 0
[6 empty] | 25.167 | 25 | 20.223 | 0 | 0
[6 empty] | 25.177 | 25 | 20.223 | 0 | 0
[6 empty] | 25.187 | 25 | 20.223 | 0 | 0
[6 empty] | 25.197 | 25 | 20.223 | 0 | 0
[6 empty] | 25.207 | 25 | 20.223 | 0 | 0
[6 empty] | 25.217 | 25 | 20.223 | 0 | 0
[6 empty] | 25.227 | 25 | 20.223 | 0 | 0
[6 empty] | 25.237 | 25 | 20.223 | 0 | 0
[6 empty] | 25.246 | 25 | 20.223 | 0 | 0
[6 empty] | 25.256 | 25 | 20.223 | 0 | 0
[6 empty] | 25.266 | 25 | 20.223 | 0 | 0
[6 empty] | 25.276 | 25 | 20.223 | 0 | 0
[6 empty] | 25.286 | 25 | 20.223 | 0 | 0
[6 empty] | 25.296 | 25 | 20.223 | 0 | 0
[6 empty] | 25.306 | 25 | 20.223 | 0 | 0
[6 empty] | 25.316 | 25 | 20.223 | 0 | 0
[6 empty] | 25.326 | 25 | 20.223 | 0 | 0
[6 empty] | 25.336 | 25 | 20.223 | 0 | 0
[6 empty] | 25.346 | 25 | 20.223 | 0 | 0
[6 empty] | 25.356 | 25 | 20.223 | 0 | 0
[6 empty] | 25.366 | 25 | 20.223 | 0 | 0
[6 empty] | 25.376 | 25 | 20.223 | 0 | 0
[6 empty] | 25.386 | 25 | 20.223 | 0 | 0
[6 empty] | 25.396 | 25 | 20.223 | 0 | 0
[6 empty] | 25.406 | 25 | 20.223 | 0 | 0
[6 empty] | 25.416 | 25 | 20.223 | 0 | 0
[6 empty] | 25.426 | 25 | 20.223 | 0 | 0
[6 empty] | 25.436 | 25 | 20.223 | 0 | 0
[6 empty] | 25.446 | 25 | 20.223 | 0 | 0
[6 empty] | 25.456 | 25 | 20.223 | 0 | 0
[6 empty] | 25.466 | 25 | 20.223 | 0 | 0
[6 empty] | 25.476 | 25 | 20.223 | 0 | 0
[6 empty] | 25.486 | 25 | 20.223 | 0 | 0
[6 empty] | 25.496 | 25 | 20.223 | 0 | 0
[6 empty] | 25.506 | 25 | 20.223 | 0 | 0
[6 empty] | 25.516 | 25 | 20.223 | 0 | 0
[6 empty] | 25.526 | 25 | 20.223 | 0 | 0
[6 empty] | 25.536 | 25 | 20.223 | 0 | 0
[6 empty] | 25.546 | 25 | 20.223 | 0 | 0
[6 empty] | 25.556 | 25 | 20.223 | 0 | 0
[6 empty] | 25.566 | 25 | 20.223 | 0 | 0
[6 empty] | 25.576 | 25 | 20.223 | 0 | 0
[6 empty] | 25.586 | 25 | 20.223 | 0 | 0
[6 empty] | 25.596 | 25 | 20.223 | 0 | 0
[6 empty] | 25.606 | 25 | 20.223 | 0 | 0
[6 empty] | 25.616 | 25 | 20.223 | 0 | 0
[6 empty] | 25.626 | 25 | 20.223 | 0 | 0
[6 empty] | 25.636 | 25 | 20.223 | 0 | 0
[6 empty] | 25.646 | 25 | 20.223 | 0 | 0
[6 empty] | 25.656 | 25 | 20.223 | 0 | 0
[6 empty] | 25.666 | 25 | 20.223 | 0 | 0
[6 empty] | 25.676 | 25 | 20.223 | 0 | 0
[6 empty] | 25.686 | 25 | 20.223 | 0 | 0
[6 empty] | 25.696 | 25 | 20.223 | 0 | 0
[6 empty] | 25.706 | 25 | 20.223 | 0 | 0
[6 empty] | 25.716 | 25 | 20.223 | 0 | 0
[6 empty] | 25.726 | 25 | 20.223 | 0 | 0
[6 empty] | 25.736 | 25 | 20.223 | 0 | 0
[6 empty] | 25.746 | 25 | 20.223 | 0 | 0
[6 empty] | 25.756 | 25 | 20.223 | 0 | 0
[6 empty] | 25.766 | 25 | 20.223 | 0 | 0
[6 empty] | 25.776 | 25 | 20.223 | 0 | 0
[6 empty] | 25.786 | 25 | 20.223 | 0 | 0
[6 empty] | 25.796 | 25 | 20.223 | 0 | 0
[6 empty] | 25.806 | 25 | 20.224 | 0 | 0
[6 empty] | 25.816 | 25 | 20.224 | 0 | 0
[6 empty] | 25.826 | 25 | 20.224 | 0 | 0
[6 empty] | 25.836 | 25 | 20.224 | 0 | 0
[6 empty] | 25.846 | 25 | 20.224 | 0 | 0
[6 empty] | 25.856 | 25 | 20.224 | 0 | 0
[6 empty] | 25.866 | 25 | 20.224 | 0 | 0
[6 empty] | 25.876 | 25 | 20.224 | 0 | 0
[6 empty] | 25.886 | 25 | 20.224 | 0 | 0
[6 empty] | 25.896 | 25 | 20.224 | 0 | 0
[6 empty] | 25.906 | 25 | 20.224 | 0 | 0
[6 empty] | 25.916 | 25 | 20.224 | 0 | 0
[6 empty] | 25.926 | 25 | 20.224 | 0 | 0
[6 empty] | 25.936 | 25 | 20.224 | 0 | 0
[6 empty] | 25.946 | 25 | 20.224 | 0 | 0
[6 empty] | 25.956 | 25 | 20.224 | 0 | 0
[6 empty] | 25.966 | 25 | 20.224 | 0 | 0
[6 empty] | 25.976 | 25 | 20.224 | 0 | 0
[6 empty] | 25.986 | 25 | 20.224 | 0 | 0
[6 empty] | 25.996 | 25 | 20.224 | 0 | 0
[6 empty] | 26.006 | 25 | 20.224 | 0 | 0
[6 empty] | 26.016 | 25 | 20.224 | 0 | 0
[6 empty] | 26.026 | 25 | 20.224 | 0 | 0
[6 empty] | 26.036 | 25 | 20.224 | 0 | 0
[6 empty] | 26.046 | 25 | 20.224 | 0 | 0
[6 empty] | 26.056 | 25 | 20.224 | 0 | 0
[6 empty] | 26.066 | 25 | 20.224 | 0 | 0
[6 empty] | 26.076 | 25 | 20.224 | 0 | 0
[6 empty] | 26.086 | 25 | 20.224 | 0 | 0
[6 empty] | 26.096 | 25 | 20.224 | 0 | 0
[6 empty] | 26.106 | 25 | 20.224 | 0 | 0
[6 empty] | 26.116 | 25 | 20.224 | 0 | 0
[6 empty] | 26.126 | 25 | 20.224 | 0 | 0
[6 empty] | 26.136 | 25 | 20.224 | 0 | 0
[6 empty] | 26.146 | 25 | 20.224 | 0 | 0
[6 empty] | 26.156 | 25 | 20.224 | 0 | 0
[6 empty] | 26.166 | 25 | 20.224 | 0 | 0
[6 empty] | 26.176 | 25 | 20.224 | 0 | 0
[6 empty] | 26.186 | 25 | 20.224 | 0 | 0
[6 empty] | 26.196 | 25 | 20.224 | 0 | 0
[6 empty] | 26.206 | 25 | 20.224 | 0 | 0
[6 empty] | 26.216 | 25 | 20.224 | 0 | 0
[6 empty] | 26.226 | 25 | 20.224 | 0 | 0
[6 empty] | 26.236 | 25 | 20.224 | 0 | 0
[6 empty] | 26.246 | 25 | 20.224 | 0 | 0
[6 empty] | 26.256 | 25 | 20.224 | 0 | 0
[6 empty] | 26.266 | 25 | 20.224 | 0 | 0
[6 empty] | 26.276 | 25 | 20.224 | 0 | 0
[6 empty] | 26.286 | 25 | 20.224 | 0 | 0
[6 empty] | 26.296 | 25 | 20.224 | 0 | 0
[6 empty] | 26.306 | 25 | 20.224 | 0 | 0
[6 empty] | 26.316 | 25 | 20.224 | 0 | 0
[6 empty] | 26.326 | 25 | 20.224 | 0 | 0
[6 empty] | 26.336 | 25 | 20.224 | 0 | 0
[6 empty] | 26.346 | 25 | 20.224 | 0 | 0
[6 empty] | 26.356 | 25 | 20.224 | 0 | 0
[6 empty] | 26.366 | 25 | 20.224 | 0 | 0
[6 empty] | 26.376 | 25 | 20.224 | 0 | 0
[6 empty] | 26.386 | 25 | 20.224 | 0 | 0
[6 empty] | 26.396 | 25 | 20.224 | 0 | 0
[6 empty] | 26.406 | 25 | 20.224 | 0 | 0
[6 empty] | 26.416 | 25 | 20.224 | 0 | 0
[6 empty] | 26.426 | 25 | 20.224 | 0 | 0
[6 empty] | 26.436 | 25 | 20.225 | 0 | 0
[6 empty] | 26.446 | 25 | 20.225 | 0 | 0
[6 empty] | 26.456 | 25 | 20.225 | 0 | 0
[6 empty] | 26.466 | 25 | 20.225 | 0 | 0
[6 empty] | 26.476 | 25 | 20.225 | 0 | 0
[6 empty] | 26.486 | 25 | 20.225 | 0 | 0
[6 empty] | 26.496 | 25 | 20.225 | 0 | 0
[6 empty] | 26.506 | 25 | 20.225 | 0 | 0
[6 empty] | 26.516 | 25 | 20.225 | 0 | 0
[6 empty] | 26.526 | 25 | 20.225 | 0 | 0
[6 empty] | 26.536 | 25 | 20.225 | 0 | 0
[6 empty] | 26.546 | 25 | 20.225 | 0 | 0
[6 empty] | 26.556 | 25 | 20.225 | 0 | 0
[6 empty] | 26.566 | 25 | 20.225 | 0 | 0
[6 empty] | 26.576 | 25 | 20.225 | 0 | 0
[6 empty] | 26.586 | 25 | 20.225 | 0 | 0
[6 empty] | 26.596 | 25 | 20.225 | 0 | 0
[6 empty] | 26.606 | 25 | 20.225 | 0 | 0
[6 empty] | 26.616 | 25 | 20.225 | 0 | 0
[6 empty] | 26.626 | 25 | 20.225 | 0 | 0
[6 empty] | 26.636 | 25 | 20.225 | 0 | 0
[6 empty] | 26.646 | 25 | 20.225 | 0 | 0
[6 empty] | 26.656 | 25 | 20.225 | 0 | 0
[6 empty] | 26.666 | 25 | 20.225 | 0 | 0
[6 empty] | 26.676 | 25 | 20.225 | 0 | 0
[6 empty] | 26.686 | 25 | 20.225 | 0 | 0
[6 empty] | 26.696 | 25 | 20.225 | 0 | 0
[6 empty] | 26.706 | 25 | 20.225 | 0 | 0
[6 empty] | 26.716 | 25 | 20.225 | 0 | 0
[6 empty] | 26.726 | 25 | 20.225 | 0 | 0
[6 empty] | 26.736 | 25 | 20.225 | 0 | 0
[6 empty] | 26.746 | 25 | 20.225 | 0 | 0
[6 empty] | 26.756 | 25 | 20.225 | 0 | 0
[6 empty] | 26.766 | 25 | 20.225 | 0 | 0
[6 empty] | 26.776 | 25 | 20.225 | 0 | 0
[6 empty] | 26.786 | 25 | 20.225 | 0 | 0
[6 empty] | 26.796 | 25 | 20.225 | 0 | 0
[6 empty] | 26.806 | 25 | 20.225 | 0 | 0
[6 empty] | 26.816 | 25 | 20.225 | 0 | 0
[6 empty] | 26.826 | 25 | 20.225 | 0 | 0
[6 empty] | 26.836 | 25 | 20.225 | 0 | 0
[6 empty] | 26.846 | 25 | 20.225 | 0 | 0
[6 empty] | 26.856 | 25 | 20.225 | 0 | 0
[6 empty] | 26.866 | 25 | 20.225 | 0 | 0
[6 empty] | 26.876 | 25 | 20.225 | 0 | 0
[6 empty] | 26.886 | 25 | 20.225 | 0 | 0
[6 empty] | 26.896 | 25 | 20.225 | 0 | 0
[6 empty] | 26.906 | 25 | 20.225 | 0 | 0
[6 empty] | 26.916 | 25 | 20.225 | 0 | 0
[6 empty] | 26.926 | 25 | 20.225 | 0 | 0
[6 empty] | 26.936 | 25 | 20.225 | 0 | 0
[6 empty] | 26.946 | 25 | 20.225 | 0 | 0
[6 empty] | 26.956 | 25 | 20.225 | 0 | 0
[6 empty] | 26.966 | 25 | 20.225 | 0 | 0
[6 empty] | 26.976 | 25 | 20.225 | 0 | 0
[6 empty] | 26.986 | 25 | 20.225 | 0 | 0
[6 empty] | 26.996 | 25 | 20.225 | 0 | 0
[6 empty] | 27.006 | 25 | 20.225 | 0 | 0
[6 empty] | 27.016 | 25 | 20.225 | 0 | 0
[6 empty] | 27.026 | 25 | 20.225 | 0 | 0
[6 empty] | 27.036 | 25 | 20.225 | 0 | 0
[6 empty] | 27.046 | 25 | 20.225 | 0 | 0
[6 empty] | 27.056 | 25 | 20.225 | 0 | 0
[6 empty] | 27.066 | 25 | 20.225 | 0 | 0
[6 empty] | 27.076 | 25 | 20.226 | 0 | 0
[6 empty] | 27.086 | 25 | 20.226 | 0 | 0
[6 empty] | 27.096 | 25 | 20.226 | 0 | 0
[6 empty] | 27.106 | 25 | 20.226 | 0 | 0
[6 empty] | 27.116 | 25 | 20.226 | 0 | 0
[6 empty] | 27.126 | 25 | 20.226 | 0 | 0
[6 empty] | 27.136 | 25 | 20.226 | 0 | 0
[6 empty] | 27.146 | 25 | 20.226 | 0 | 0
[6 empty] | 27.156 | 25 | 20.226 | 0 | 0
[6 empty] | 27.166 | 25 | 20.226 | 0 | 0
[6 empty] | 27.176 | 25 | 20.226 | 0 | 0
[6 empty] | 27.186 | 25 | 20.226 | 0 | 0
[6 empty] | 27.196 | 25 | 20.226 | 0 | 0
[6 empty] | 27.206 | 25 | 20.226 | 0 | 0
[6 empty] | 27.216 | 25 | 20.226 | 0 | 0
[6 empty] | 27.226 | 25 | 20.226 | 0 | 0
[6 empty] | 27.236 | 25 | 20.226 | 0 | 0
[6 empty] | 27.246 | 25 | 20.226 | 0 | 0
[6 empty] | 27.256 | 25 | 20.226 | 0 | 0
[6 empty] | 27.266 | 25 | 20.226 | 0 | 0
[6 empty] | 27.276 | 25 | 20.226 | 0 | 0
[6 empty] | 27.286 | 25 | 20.226 | 0 | 0
[6 empty] | 27.296 | 25 | 20.226 | 0 | 0
[6 empty] | 27.306 | 25 | 20.226 | 0 | 0
[6 empty] | 27.316 | 25 | 20.226 | 0 | 0
[6 empty] | 27.326 | 25 | 20.226 | 0 | 0
[6 empty] | 27.336 | 25 | 20.226 | 0 | 0
[6 empty] | 27.346 | 25 | 20.226 | 0 | 0
[6 empty] | 27.356 | 25 | 20.226 | 0 | 0
[6 empty] | 27.366 | 25 | 20.226 | 0 | 0
[6 empty] | 27.376 | 25 | 20.226 | 0 | 0
[6 empty] | 27.386 | 25 | 20.226 | 0 | 0
[6 empty] | 27.396 | 25 | 20.226 | 0 | 0
[6 empty] | 27.406 | 25 | 20.226 | 0 | 0
[6 empty] | 27.416 | 25 | 20.226 | 0 | 0
[6 empty] | 27.426 | 25 | 20.226 | 0 | 0
[6 empty] | 27.436 | 25 | 20.226 | 0 | 0
[6 empty] | 27.446 | 25 | 20.226 | 0 | 0
[6 empty] | 27.456 | 25 | 20.226 | 0 | 0
[6 empty] | 27.466 | 25 | 20.226 | 0 | 0
[6 empty] | 27.476 | 25 | 20.226 | 0 | 0
[6 empty] | 27.486 | 25 | 20.226 | 0 | 0
[6 empty] | 27.496 | 25 | 20.226 | 0 | 0
[6 empty] | 27.506 | 25 | 20.226 | 0 | 0
[6 empty] | 27.516 | 25 | 20.226 | 0 | 0
[6 empty] | 27.526 | 25 | 20.226 | 0 | 0
[6 empty] | 27.536 | 25 | 20.226 | 0 | 0
[6 empty] | 27.546 | 25 | 20.226 | 0 | 0
[6 empty] | 27.556 | 25 | 20.226 | 0 | 0
[6 empty] | 27.566 | 25 | 20.226 | 0 | 0
[6 empty] | 27.576 | 25 | 20.226 | 0 | 0
[6 empty] | 27.586 | 25 | 20.226 | 0 | 0
[6 empty] | 27.596 | 25 | 20.226 | 0 | 0
[6 empty] | 27.606 | 25 | 20.226 | 0 | 0
[6 empty] | 27.616 | 25 | 20.226 | 0 | 0
[6 empty] | 27.626 | 25 | 20.226 | 0 | 0
[6 empty] | 27.636 | 25 | 20.226 | 0 | 0
[6 empty] | 27.646 | 25 | 20.226 | 0 | 0
[6 empty] | 27.656 | 25 | 20.226 | 0 | 0
[6 empty] | 27.666 | 25 | 20.226 | 0 | 0
[6 empty] | 27.676 | 25 | 20.226 | 0 | 0
[6 empty] | 27.686 | 25 | 20.226 | 0 | 0
[6 empty] | 27.696 | 25 | 20.226 | 0 | 0
[6 empty] | 27.706 | 25 | 20.227 | 0 | 0
[6 empty] | 27.716 | 25 | 20.227 | 0 | 0
[6 empty] | 27.726 | 25 | 20.227 | 0 | 0
[6 empty] | 27.736 | 25 | 20.227 | 0 | 0
[6 empty] | 27.746 | 25 | 20.227 | 0 | 0
[6 empty] | 27.756 | 25 | 20.227 | 0 | 0
[6 empty] | 27.766 | 25 | 20.227 | 0 | 0
[6 empty] | 27.776 | 25 | 20.227 | 0 | 0
[6 empty] | 27.786 | 25 | 20.227 | 0 | 0
[6 empty] | 27.796 | 25 | 20.227 | 0 | 0
[6 empty] | 27.806 | 25 | 20.227 | 0 | 0
[6 empty] | 27.816 | 25 | 20.227 | 0 | 0
[6 empty] | 27.826 | 25 | 20.227 | 0 | 0
[6 empty] | 27.836 | 25 | 20.227 | 0 | 0
[6 empty] | 27.846 | 25 | 20.227 | 0 | 0
[6 empty] | 27.856 | 25 | 20.227 | 0 | 0
[6 empty] | 27.866 | 25 | 20.227 | 0 | 0
[6 empty] | 27.876 | 25 | 20.227 | 0 | 0
[6 empty] | 27.886 | 25 | 20.227 | 0 | 0
[6 empty] | 27.896 | 25 | 20.227 | 0 | 0
[6 empty] | 27.906 | 25 | 20.227 | 0 | 0
[6 empty] | 27.916 | 25 | 20.227 | 0 | 0
[6 empty] | 27.926 | 25 | 20.227 | 0 | 0
[6 empty] | 27.936 | 25 | 20.227 | 0 | 0
[6 empty] | 27.946 | 25 | 20.227 | 0 | 0
[6 empty] | 27.956 | 25 | 20.227 | 0 | 0
[6 empty] | 27.966 | 25 | 20.227 | 0 | 0
[6 empty] | 27.976 | 25 | 20.227 | 0 | 0
[6 empty] | 27.986 | 25 | 20.227 | 0 | 0
[6 empty] | 27.996 | 25 | 20.227 | 0 | 0
[6 empty] | 28.006 | 25 | 20.227 | 0 | 0
[6 empty] | 28.016 | 25 | 20.227 | 0 | 0
[6 empty] | 28.026 | 25 | 20.227 | 0 | 0
[6 empty] | 28.036 | 25 | 20.227 | 0 | 0
[6 empty] | 28.046 | 25 | 20.227 | 0 | 0
[6 empty] | 28.056 | 25 | 20.227 | 0 | 0
[6 empty] | 28.066 | 25 | 20.227 | 0 | 0
[6 empty] | 28.076 | 25 | 20.227 | 0 | 0
[6 empty] | 28.086 | 25 | 20.227 | 0 | 0
[6 empty] | 28.096 | 25 | 20.227 | 0 | 0
[6 empty] | 28.106 | 25 | 20.227 | 0 | 0
[6 empty] | 28.116 | 25 | 20.227 | 0 | 0
[6 empty] | 28.126 | 25 | 20.227 | 0 | 0
[6 empty] | 28.136 | 25 | 20.227 | 0 | 0
[6 empty] | 28.146 | 25 | 20.227 | 0 | 0
[6 empty] | 28.156 | 25 | 20.227 | 0 | 0
[6 empty] | 28.166 | 25 | 20.227 | 0 | 0
[6 empty] | 28.176 | 25 | 20.227 | 0 | 0
[6 empty] | 28.186 | 25 | 20.227 | 0 | 0
[6 empty] | 28.196 | 25 | 20.227 | 0 | 0
[6 empty] | 28.206 | 25 | 20.227 | 0 | 0
[6 empty] | 28.216 | 25 | 20.227 | 0 | 0
[6 empty] | 28.226 | 25 | 20.227 | 0 | 0
[6 empty] | 28.236 | 25 | 20.227 | 0 | 0
[6 empty] | 28.246 | 25 | 20.227 | 0 | 0
[6 empty] | 28.256 | 25 | 20.227 | 0 | 0
[6 empty] | 28.266 | 25 | 20.227 | 0 | 0
[6 empty] | 28.276 | 25 | 20.227 | 0 | 0
[6 empty] | 28.286 | 25 | 20.227 | 0 | 0
[6 empty] | 28.296 | 25 | 20.227 | 0 | 0
[6 empty] | 28.306 | 25 | 20.227 | 0 | 0
[6 empty] | 28.316 | 25 | 20.227 | 0 | 0
[6 empty] | 28.326 | 25 | 20.227 | 0 | 0
[6 empty] | 28.336 | 25 | 20.227 | 0 | 0
[6 empty] | 28.346 | 25 | 20.228 | 0 | 0
[6 empty] | 28.356 | 25 | 20.228 | 0 | 0
[6 empty] | 28.366 | 25 | 20.228 | 0 | 0
[6 empty] | 28.376 | 25 | 20.228 | 0 | 0
[6 empty] | 28.386 | 25 | 20.228 | 0 | 0
[6 empty] | 28.396 | 25 | 20.228 | 0 | 0
[6 empty] | 28.406 | 25 | 20.228 | 0 | 0
[6 empty] | 28.416 | 25 | 20.228 | 0 | 0
[6 empty] | 28.426 | 25 | 20.228 | 0 | 0
[6 empty] | 28.436 | 25 | 20.228 | 0 | 0
[6 empty] | 28.446 | 25 | 20.228 | 0 | 0
[6 empty] | 28.456 | 25 | 20.228 | 0 | 0
[6 empty] | 28.466 | 25 | 20.228 | 0 | 0
[6 empty] | 28.476 | 25 | 20.228 | 0 | 0
[6 empty] | 28.486 | 25 | 20.228 | 0 | 0
[6 empty] | 28.496 | 25 | 20.228 | 0 | 0
[6 empty] | 28.506 | 25 | 20.228 | 0 | 0
[6 empty] | 28.516 | 25 | 20.228 | 0 | 0
[6 empty] | 28.526 | 25 | 20.228 | 0 | 0
[6 empty] | 28.536 | 25 | 20.228 | 0 | 0
[6 empty] | 28.546 | 25 | 20.228 | 0 | 0
[6 empty] | 28.556 | 25 | 20.228 | 0 | 0
[6 empty] | 28.566 | 25 | 20.228 | 0 | 0
[6 empty] | 28.576 | 25 | 20.228 | 0 | 0
[6 empty] | 28.586 | 25 | 20.228 | 0 | 0
[6 empty] | 28.596 | 25 | 20.228 | 0 | 0
[6 empty] | 28.606 | 25 | 20.228 | 0 | 0
[6 empty] | 28.616 | 25 | 20.228 | 0 | 0
[6 empty] | 28.626 | 25 | 20.228 | 0 | 0
[6 empty] | 28.636 | 25 | 20.228 | 0 | 0
[6 empty] | 28.646 | 25 | 20.228 | 0 | 0
[6 empty] | 28.656 | 25 | 20.228 | 0 | 0
[6 empty] | 28.666 | 25 | 20.228 | 0 | 0
[6 empty] | 28.676 | 25 | 20.228 | 0 | 0
[6 empty] | 28.686 | 25 | 20.228 | 0 | 0
[6 empty] | 28.696 | 25 | 20.228 | 0 | 0
[6 empty] | 28.706 | 25 | 20.228 | 0 | 0
[6 empty] | 28.716 | 25 | 20.228 | 0 | 0
[6 empty] | 28.726 | 25 | 20.228 | 0 | 0
[6 empty] | 28.736 | 25 | 20.228 | 0 | 0
[6 empty] | 28.746 | 25 | 20.228 | 0 | 0
[6 empty] | 28.756 | 25 | 20.228 | 0 | 0
[6 empty] | 28.766 | 25 | 20.228 | 0 | 0
[6 empty] | 28.776 | 25 | 20.228 | 0 | 0
[6 empty] | 28.786 | 25 | 20.228 | 0 | 0
[6 empty] | 28.796 | 25 | 20.228 | 0 | 0
[6 empty] | 28.806 | 25 | 20.228 | 0 | 0
[6 empty] | 28.816 | 25 | 20.228 | 0 | 0
[6 empty] | 28.826 | 25 | 20.228 | 0 | 0
[6 empty] | 28.836 | 25 | 20.228 | 0 | 0
[6 empty] | 28.846 | 25 | 20.228 | 0 | 0
[6 empty] | 28.856 | 25 | 20.228 | 0 | 0
[6 empty] | 28.866 | 25 | 20.228 | 0 | 0
[6 empty] | 28.876 | 25 | 20.228 | 0 | 0
[6 empty] | 28.886 | 25 | 20.228 | 0 | 0
[6 empty] | 28.896 | 25 | 20.228 | 0 | 0
[6 empty] | 28.906 | 25 | 20.228 | 0 | 0
[6 empty] | 28.916 | 25 | 20.228 | 0 | 0
[6 empty] | 28.926 | 25 | 20.228 | 0 | 0
[6 empty] | 28.936 | 25 | 20.228 | 0 | 0
[6 empty] | 28.946 | 25 | 20.228 | 0 | 0
[6 empty] | 28.956 | 25 | 20.228 | 0 | 0
[6 empty] | 28.966 | 25 | 20.228 | 0 | 0
[6 empty] | 28.976 | 25 | 20.228 | 0 | 0
[6 empty] | 28.986 | 25 | 20.229 | 0 | 0
[6 empty] | 28.996 | 25 | 20.229 | 0 | 0
[6 empty] | 29.006 | 25 | 20.229 | 0 | 0
[6 empty] | 29.016 | 25 | 20.229 | 0 | 0
[6 empty] | 29.026 | 25 | 20.229 | 0 | 0
[6 empty] | 29.036 | 25 | 20.229 | 0 | 0
[6 empty] | 29.046 | 25 | 20.229 | 0 | 0
[6 empty] | 29.056 | 25 | 20.229 | 0 | 0
[6 empty] | 29.066 | 25 | 20.229 | 0 | 0
[6 empty] | 29.076 | 25 | 20.229 | 0 | 0
[6 empty] | 29.086 | 25 | 20.229 | 0 | 0
[6 empty] | 29.096 | 25 | 20.229 | 0 | 0
[6 empty] | 29.106 | 25 | 20.229 | 0 | 0
[6 empty] | 29.116 | 25 | 20.229 | 0 | 0
[6 empty] | 29.126 | 25 | 20.229 | 0 | 0
[6 empty] | 29.136 | 25 | 20.229 | 0 | 0
[6 empty] | 29.146 | 25 | 20.229 | 0 | 0
[6 empty] | 29.156 | 25 | 20.229 | 0 | 0
[6 empty] | 29.166 | 25 | 20.229 | 0 | 0
[6 empty] | 29.176 | 25 | 20.229 | 0 | 0
[6 empty] | 29.186 | 25 | 20.229 | 0 | 0
[6 empty] | 29.196 | 25 | 20.229 | 0 | 0
[6 empty] | 29.206 | 25 | 20.229 | 0 | 0
[6 empty] | 29.216 | 25 | 20.229 | 0 | 0
[6 empty] | 29.226 | 25 | 20.229 | 0 | 0
[6 empty] | 29.236 | 25 | 20.229 | 0 | 0
[6 empty] | 29.246 | 25 | 20.229 | 0 | 0
[6 empty] | 29.256 | 25 | 20.229 | 0 | 0
[6 empty] | 29.266 | 25 | 20.229 | 0 | 0
[6 empty] | 29.276 | 25 | 20.229 | 0 | 0
[6 empty] | 29.286 | 25 | 20.229 | 0 | 0
[6 empty] | 29.296 | 25 | 20.229 | 0 | 0
[6 empty] | 29.306 | 25 | 20.229 | 0 | 0
[6 empty] | 29.316 | 25 | 20.229 | 0 | 0
[6 empty] | 29.326 | 25 | 20.229 | 0 | 0
[6 empty] | 29.336 | 25 | 20.229 | 0 | 0
[6 empty] | 29.346 | 25 | 20.229 | 0 | 0
[6 empty] | 29.356 | 25 | 20.229 | 0 | 0
[6 empty] | 29.366 | 25 | 20.229 | 0 | 0
[6 empty] | 29.376 | 25 | 20.229 | 0 | 0
[6 empty] | 29.386 | 25 | 20.229 | 0 | 0
[6 empty] | 29.396 | 25 | 20.229 | 0 | 0
[6 empty] | 29.406 | 25 | 20.229 | 0 | 0
[6 empty] | 29.416 | 25 | 20.229 | 0 | 0
[6 empty] | 29.426 | 25 | 20.229 | 0 | 0
[6 empty] | 29.436 | 25 | 20.229 | 0 | 0
[6 empty] | 29.446 | 25 | 20.229 | 0 | 0
[6 empty] | 29.456 | 25 | 20.229 | 0 | 0
[6 empty] | 29.466 | 25 | 20.229 | 0 | 0
[6 empty] | 29.476 | 25 | 20.229 | 0 | 0
[6 empty] | 29.486 | 25 | 20.229 | 0 | 0
[6 empty] | 29.496 | 25 | 20.229 | 0 | 0
[6 empty] | 29.506 | 25 | 20.229 | 0 | 0
[6 empty] | 29.516 | 25 | 20.229 | 0 | 0
[6 empty] | 29.526 | 25 | 20.229 | 0 | 0
[6 empty] | 29.536 | 25 | 20.229 | 0 | 0
[6 empty] | 29.546 | 25 | 20.229 | 0 | 0
[6 empty] | 29.556 | 25 | 20.229 | 0 | 0
[6 empty] | 29.566 | 25 | 20.229 | 0 | 0
[6 empty] | 29.576 | 25 | 20.229 | 0 | 0
[6 empty] | 29.586 | 25 | 20.229 | 0 | 0
[6 empty] | 29.596 | 25 | 20.229 | 0 | 0
[6 empty] | 29.606 | 25 | 20.229 | 0 | 0
[6 empty] | 29.616 | 25 | 20.23 | 0 | 0
[6 empty] | 29.626 | 25 | 20.23 | 0 | 0
[6 empty] | 29.636 | 25 | 20.23 | 0 | 0
[6 empty] | 29.646 | 25 | 20.23 | 0 | 0
[6 empty] | 29.656 | 25 | 20.23 | 0 | 0
[6 empty] | 29.666 | 25 | 20.23 | 0 | 0
[6 empty] | 29.676 | 25 | 20.23 | 0 | 0
[6 empty] | 29.686 | 25 | 20.23 | 0 | 0
[6 empty] | 29.696 | 25 | 20.23 | 0 | 0
[6 empty] | 29.706 | 25 | 20.23 | 0 | 0
[6 empty] | 29.716 | 25 | 20.23 | 0 | 0
[6 empty] | 29.726 | 25 | 20.23 | 0 | 0
[6 empty] | 29.736 | 25 | 20.23 | 0 | 0
[6 empty] | 29.746 | 25 | 20.23 | 0 | 0
[6 empty] | 29.756 | 25 | 20.23 | 0 | 0
[6 empty] | 29.766 | 25 | 20.23 | 0 | 0
[6 empty] | 29.776 | 25 | 20.23 | 0 | 0
[6 empty] | 29.786 | 25 | 20.23 | 0 | 0
[6 empty] | 29.796 | 25 | 20.23 | 0 | 0
[6 empty] | 29.806 | 25 | 20.23 | 0 | 0
[6 empty] | 29.816 | 25 | 20.23 | 0 | 0
[6 empty] | 29.826 | 25 | 20.23 | 0 | 0
[6 empty] | 29.836 | 25 | 20.23 | 0 | 0
[6 empty] | 29.846 | 25 | 20.23 | 0 | 0
[6 empty] | 29.856 | 25 | 20.23 | 0 | 0
[6 empty] | 29.866 | 25 | 20.23 | 0 | 0
[6 empty] | 29.876 | 25 | 20.23 | 0 | 0
[6 empty] | ####################################################### |  |  |  | 0

## INPLANE
| SCRIPT |  |  |  |  | RAYSTATION
BEGIN_SCAN_DATA | [5 empty] | #######################################################
| FORMAT=CC-Export V1.9 |  |  |  |  | #RayStationVersion: 13.1.0.144
| FILE_CREATION_DATE= |  |  |  |  | #PatientName: RayTraining 40x40x40 Phantom
| LAST_MODIFIED=11-Jan-2023 11:56:47 |  |  |  |  | #PatientId: Watertank cube
| BEGIN_SCAN 1 |  |  |  |  | #CoordinateSystem: IEC patient coordinate system
|  | TASK_NAME=RayStation Phantom Profiles & PDDs |  |  |  | #######################################################
|  | PROGRAM=RayStation |  |  |  | #LineName: 1
|  | COMMENT=RayStation Script Export |  |  |  | #DoseName: Beam dose: 15x10 (Validation, CT 1)
|  | MEAS_DATE=11-Jan-2023 11:56:47 |  |  |  | #DoseEngine: Clinical: Collapsed Cone v5.7
|  | LINAC=RayStation |  |  |  | #TreatmentMachine: RSL_TrueBeam Commission time: 11/20/2018 14:29:57
|  | MODALITY=X |  |  |  | #X [cm] | Y [cm] | Z [cm] | Dose [Gy] | cGy
|  | ISOCENTER=1000.00 |  |  |  | -0.146 | 55.043 | 20.092 | 0 | 0
|  | INPLANE_AXIS=Y |  |  |  | -0.146 | 55.033 | 20.092 | 0 | 0
|  | CROSSPLANE_AXIS=X |  |  |  | -0.146 | 55.023 | 20.092 | 0 | 0
|  | DEPTH_AXIS=Z |  |  |  | -0.146 | 55.013 | 20.092 | 0 | 0
|  | INPLANE_AXIS_DIR=GUN_TARGET |  |  |  | -0.146 | 55.003 | 20.092 | 0 | 0
|  | CROSSPLANE_AXIS_DIR=LEFT_RIGHT |  |  |  | -0.146 | 54.993 | 20.092 | 0 | 0
|  | DEPTH_AXIS_DIR=UP_DOWN |  |  |  | -0.146 | 54.983 | 20.092 | 0 | 0
|  | ENERGY=6.3 |  |  |  | -0.146 | 54.974 | 20.092 | 0 | 0
|  | NOMINAL_DMAX=20.00 |  |  |  | -0.146 | 54.964 | 20.092 | 0 | 0
|  | SSD=1000.0 |  |  |  | -0.146 | 54.954 | 20.092 | 0 | 0
|  | SCD=1000.00 |  |  |  | -0.146 | 54.944 | 20.092 | 0 | 0
|  | WEDGE_ANGLE=0.00 |  |  |  | -0.146 | 54.934 | 20.092 | 0 | 0
|  | FIELD_INPLANE=150.0 |  |  |  | -0.146 | 54.924 | 20.092 | 0 | 0
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -0.146 | 54.914 | 20.092 | 0 | 0
|  | FIELD_TYPE=IRREGULAR |  |  |  | -0.146 | 54.904 | 20.092 | 0 | 0
|  | GANTRY=0.00 |  |  |  | -0.146 | 54.894 | 20.092 | 0 | 0
|  | GANTRY_UPRIGHT_POSITION=0 |  |  |  | -0.146 | 54.884 | 20.092 | 0 | 0
|  | GANTRY_ROTATION=CW |  |  |  | -0.146 | 54.874 | 20.092 | 0 | 0
|  | COLL_ANGLE=0.00 |  |  |  | -0.146 | 54.864 | 20.092 | 0 | 0
|  | REF_FIELD_DEPTH=50.00 |  |  |  | -0.146 | 54.854 | 20.092 | 0 | 0
|  | REF_FIELD_DEFINED=ISOCENTER |  |  |  | -0.146 | 54.844 | 20.092 | 0 | 0
|  | FIELD_INPLANE=150.0 |  |  |  | -0.146 | 54.834 | 20.092 | 0 | 0
|  | FIELD_CROSSPLANE=100.0 |  |  |  | -0.146 | 54.824 | 20.092 | 0 | 0
|  | SCAN_CURVETYPE=INPLANE_PROFILE |  |  |  | -0.146 | 54.814 | 20.092 | 0 | 0
|  | SCAN_DEPTH=100.0 |  |  |  | -0.146 | 54.804 | 20.092 | 0 | 0
|  | SCAN_OFFAXIS_INPLANE=0.00 |  |  |  | -0.146 | 54.794 | 20.092 | 0 | 0
|  | SCAN_OFFAXIS_CROSSPLANE=0.00 |  |  |  | -0.146 | 54.784 | 20.092 | 0 | 0
|  | SCAN_ANGLE=0.00 |  |  |  | -0.146 | 54.774 | 20.092 | 0 | 0
|  | SCAN_DIAGONAL=NOT_DIAGONAL |  |  |  | -0.146 | 54.764 | 20.092 | 0 | 0
|  | SCAN_DIRECTION=NEGATIVE |  |  |  | -0.146 | 54.754 | 20.092 | 0 | 0
|  | MEAS_MEDIUM=WATER |  |  |  | -0.146 | 54.744 | 20.092 | 0 | 0
|  | MEAS_UNIT=cGy |  |  |  | -0.146 | 54.734 | 20.092 | 0 | 0
|  | INCLINATION_ANGLE=0.00 |  |  |  | -0.146 | 54.724 | 20.092 | 0 | 0
|  | PRESSURE=1013.20 |  |  |  | -0.146 | 54.714 | 20.092 | 0 | 0
|  | TEMPERATURE=20.00 |  |  |  | -0.146 | 54.704 | 20.092 | 0 | 0
|  | NORM_TEMPERATURE=20.00 |  |  |  | -0.146 | 54.694 | 20.092 | 0 | 0
|  | CORRECTION_FACTOR=1.0000; |  |  |  | -0.146 | 54.684 | 20.092 | 0 | 0
|  | BEGIN_DATA |  |  |  | -0.146 | 54.674 | 20.092 | 0 | 0
|  |  | -5 | 0.0012718 | 1 | -0.146 | 54.664 | 20.092 | 0 | 0
|  |  | -4.95 | 0.0012833 | 1 | -0.146 | 54.654 | 20.092 | 0 | 0
|  |  | -4.9 | 0.0012947 | 1 | -0.146 | 54.644 | 20.092 | 0 | 0
|  |  | -4.85 | 0.0013062 | 1 | -0.146 | 54.634 | 20.092 | 0 | 0
|  |  | -4.8 | 0.0013177 | 1 | -0.146 | 54.624 | 20.092 | 0 | 0
|  |  | -4.75 | 0.0013473 | 1 | -0.146 | 54.614 | 20.092 | 0 | 0
|  |  | -4.7 | 0.0013769 | 1 | -0.146 | 54.604 | 20.092 | 0 | 0
|  |  | -4.65 | 0.0014064 | 1 | -0.146 | 54.594 | 20.092 | 0 | 0
|  |  | -4.6 | 0.001436 | 1 | -0.146 | 54.584 | 20.092 | 0 | 0
|  |  | -4.55 | 0.0014519 | 1 | -0.146 | 54.574 | 20.092 | 0 | 0
|  |  | -4.5 | 0.0014678 | 1 | -0.146 | 54.564 | 20.092 | 0 | 0
|  |  | -4.45 | 0.0014836 | 1 | -0.146 | 54.554 | 20.092 | 0 | 0
|  |  | -4.4 | 0.0014995 | 1 | -0.146 | 54.544 | 20.092 | 0 | 0
|  |  | -4.35 | 0.0015241 | 1 | -0.146 | 54.534 | 20.092 | 0 | 0
|  |  | -4.3 | 0.0015486 | 1 | -0.146 | 54.524 | 20.092 | 0 | 0
|  |  | -4.25 | 0.0015732 | 1 | -0.146 | 54.514 | 20.092 | 0 | 0
|  |  | -4.2 | 0.0015977 | 1 | -0.146 | 54.504 | 20.092 | 0 | 0
|  |  | -4.15 | 0.0016264 | 1 | -0.146 | 54.494 | 20.091 | 0 | 0
|  |  | -4.1 | 0.0016551 | 1 | -0.146 | 54.484 | 20.091 | 0 | 0
|  |  | -4.05 | 0.0016838 | 1 | -0.146 | 54.474 | 20.091 | 0 | 0
|  |  | -4 | 0.0017124 | 1 | -0.146 | 54.464 | 20.091 | 0 | 0
|  |  | -3.95 | 0.0017292 | 1 | -0.146 | 54.454 | 20.091 | 0 | 0
|  |  | -3.9 | 0.001746 | 1 | -0.146 | 54.444 | 20.091 | 0 | 0
|  |  | -3.85 | 0.0017628 | 1 | -0.146 | 54.434 | 20.091 | 0 | 0
|  |  | -3.8 | 0.0017796 | 1 | -0.146 | 54.424 | 20.091 | 0 | 0
|  |  | -3.75 | 0.0018182 | 1 | -0.146 | 54.414 | 20.091 | 0 | 0
|  |  | -3.7 | 0.0018569 | 1 | -0.146 | 54.404 | 20.091 | 0 | 0
|  |  | -3.65 | 0.0018956 | 1 | -0.146 | 54.394 | 20.091 | 0 | 0
|  |  | -3.6 | 0.0019342 | 1 | -0.146 | 54.384 | 20.091 | 0 | 0
|  |  | -3.55 | 0.001956 | 1 | -0.146 | 54.374 | 20.091 | 0 | 0
|  |  | -3.5 | 0.0019778 | 1 | -0.146 | 54.364 | 20.091 | 0 | 0
|  |  | -3.45 | 0.0019996 | 1 | -0.146 | 54.354 | 20.091 | 0 | 0
|  |  | -3.4 | 0.0020214 | 1 | -0.146 | 54.344 | 20.091 | 0 | 0
|  |  | -3.35 | 0.0020512 | 1 | -0.146 | 54.334 | 20.091 | 0 | 0
|  |  | -3.3 | 0.0020809 | 1 | -0.146 | 54.324 | 20.091 | 0 | 0
|  |  | -3.25 | 0.0021107 | 1 | -0.146 | 54.314 | 20.091 | 0 | 0
|  |  | -3.2 | 0.0021404 | 1 | -0.146 | 54.304 | 20.091 | 0 | 0
|  |  | -3.15 | 0.0021803 | 1 | -0.146 | 54.294 | 20.091 | 0 | 0
|  |  | -3.1 | 0.0022203 | 1 | -0.146 | 54.284 | 20.091 | 0 | 0
|  |  | -3.05 | 0.0022602 | 1 | -0.146 | 54.274 | 20.091 | 0 | 0
|  |  | -3 | 0.0023001 | 1 | -0.146 | 54.264 | 20.091 | 0 | 0
|  |  | -2.95 | 0.002321 | 1 | -0.146 | 54.254 | 20.091 | 0 | 0
|  |  | -2.9 | 0.0023419 | 1 | -0.146 | 54.244 | 20.091 | 0 | 0
|  |  | -2.85 | 0.0023628 | 1 | -0.146 | 54.234 | 20.091 | 0 | 0
|  |  | -2.8 | 0.0023837 | 1 | -0.146 | 54.224 | 20.091 | 0 | 0
|  |  | -2.75 | 0.0024327 | 1 | -0.146 | 54.214 | 20.091 | 0 | 0
|  |  | -2.7 | 0.0024817 | 1 | -0.146 | 54.204 | 20.091 | 0 | 0
|  |  | -2.65 | 0.0025307 | 1 | -0.146 | 54.194 | 20.091 | 0 | 0
|  |  | -2.6 | 0.0025797 | 1 | -0.146 | 54.184 | 20.091 | 0 | 0
|  |  | -2.55 | 0.0026094 | 1 | -0.146 | 54.174 | 20.091 | 0 | 0
|  |  | -2.5 | 0.0026391 | 1 | -0.146 | 54.164 | 20.091 | 0 | 0
|  |  | -2.45 | 0.0026689 | 1 | -0.146 | 54.154 | 20.091 | 0 | 0
|  |  | -2.4 | 0.0026986 | 1 | -0.146 | 54.144 | 20.091 | 0 | 0
|  |  | -2.35 | 0.0027332 | 1 | -0.146 | 54.134 | 20.091 | 0 | 0
|  |  | -2.3 | 0.0027678 | 1 | -0.146 | 54.124 | 20.091 | 0 | 0
|  |  | -2.25 | 0.0028025 | 1 | -0.146 | 54.114 | 20.091 | 0 | 0
|  |  | -2.2 | 0.0028371 | 1 | -0.146 | 54.104 | 20.091 | 0 | 0
|  |  | -2.15 | 0.0028908 | 1 | -0.146 | 54.094 | 20.091 | 0 | 0
|  |  | -2.1 | 0.0029446 | 1 | -0.146 | 54.084 | 20.091 | 0 | 0
|  |  | -2.05 | 0.0029983 | 1 | -0.146 | 54.074 | 20.091 | 0 | 0
|  |  | -2 | 0.003052 | 1 | -0.146 | 54.064 | 20.091 | 0 | 0
|  |  | -1.95 | 0.0030778 | 1 | -0.146 | 54.054 | 20.091 | 0 | 0
|  |  | -1.9 | 0.0031036 | 1 | -0.146 | 54.044 | 20.091 | 0 | 0
|  |  | -1.85 | 0.0031293 | 1 | -0.146 | 54.034 | 20.091 | 0 | 0
|  |  | -1.8 | 0.0031551 | 1 | -0.146 | 54.024 | 20.091 | 0 | 0
|  |  | -1.75 | 0.0032158 | 1 | -0.146 | 54.014 | 20.091 | 0 | 0
|  |  | -1.7 | 0.0032766 | 1 | -0.146 | 54.004 | 20.091 | 0 | 0
|  |  | -1.65 | 0.0033373 | 1 | -0.146 | 53.994 | 20.091 | 0 | 0
|  |  | -1.6 | 0.003398 | 1 | -0.146 | 53.984 | 20.091 | 0 | 0
|  |  | -1.55 | 0.0034385 | 1 | -0.146 | 53.974 | 20.091 | 0 | 0
|  |  | -1.5 | 0.0034789 | 1 | -0.146 | 53.964 | 20.091 | 0 | 0
|  |  | -1.45 | 0.0035194 | 1 | -0.146 | 53.954 | 20.091 | 0 | 0
|  |  | -1.4 | 0.0035598 | 1 | -0.146 | 53.944 | 20.091 | 0 | 0
|  |  | -1.35 | 0.0036005 | 1 | -0.146 | 53.934 | 20.091 | 0 | 0
|  |  | -1.3 | 0.0036412 | 1 | -0.146 | 53.924 | 20.091 | 0 | 0
|  |  | -1.25 | 0.0036818 | 1 | -0.146 | 53.914 | 20.091 | 0 | 0
|  |  | -1.2 | 0.0037225 | 1 | -0.146 | 53.904 | 20.091 | 0 | 0
|  |  | -1.15 | 0.003793 | 1 | -0.146 | 53.894 | 20.091 | 0 | 0
|  |  | -1.1 | 0.0038634 | 1 | -0.146 | 53.884 | 20.091 | 0 | 0
|  |  | -1.05 | 0.0039338 | 1 | -0.146 | 53.874 | 20.091 | 0 | 0
|  |  | -1 | 0.0040043 | 1 | -0.146 | 53.864 | 20.091 | 0 | 0
|  |  | -0.95 | 0.0040366 | 1 | -0.146 | 53.854 | 20.091 | 0 | 0
|  |  | -0.9 | 0.0040689 | 1 | -0.146 | 53.844 | 20.091 | 0 | 0
|  |  | -0.85 | 0.0041013 | 1 | -0.146 | 53.834 | 20.091 | 0 | 0
|  |  | -0.8 | 0.0041336 | 1 | -0.146 | 53.824 | 20.091 | 0 | 0
|  |  | -0.75 | 0.0042065 | 1 | -0.146 | 53.814 | 20.091 | 0 | 0
|  |  | -0.7 | 0.0042793 | 1 | -0.146 | 53.804 | 20.09 | 0 | 0
|  |  | -0.65 | 0.0043522 | 1 | -0.146 | 53.794 | 20.09 | 0 | 0
|  |  | -0.6 | 0.004425 | 1 | -0.146 | 53.784 | 20.09 | 0 | 0
|  |  | -0.55 | 0.0044792 | 1 | -0.146 | 53.774 | 20.09 | 0 | 0
|  |  | -0.5 | 0.0045333 | 1 | -0.146 | 53.764 | 20.09 | 0 | 0
|  |  | -0.45 | 0.0045874 | 1 | -0.146 | 53.754 | 20.09 | 0 | 0
|  |  | -0.4 | 0.0046416 | 1 | -0.146 | 53.744 | 20.09 | 0 | 0
|  |  | -0.35 | 0.0046885 | 1 | -0.146 | 53.734 | 20.09 | 0 | 0
|  |  | -0.3 | 0.0047354 | 1 | -0.146 | 53.724 | 20.09 | 0 | 0
|  |  | -0.25 | 0.0047823 | 1 | -0.146 | 53.714 | 20.09 | 0 | 0
|  |  | -0.2 | 0.0048292 | 1 | -0.146 | 53.704 | 20.09 | 0 | 0
|  |  | -0.15 | 0.0049194 | 1 | -0.146 | 53.694 | 20.09 | 0 | 0
|  |  | -0.1 | 0.0050097 | 1 | -0.146 | 53.684 | 20.09 | 0 | 0
|  |  | -0.05 | 0.0050999 | 1 | -0.146 | 53.674 | 20.09 | 0 | 0
|  |  | 0 | 0.0051901 | 1 | -0.146 | 53.664 | 20.09 | 0 | 0
|  |  | 0.05 | 0.0052316 | 1 | -0.146 | 53.654 | 20.09 | 0 | 0
|  |  | 0.1 | 0.0052732 | 1 | -0.146 | 53.644 | 20.09 | 0 | 0
|  |  | 0.15 | 0.0053148 | 1 | -0.146 | 53.634 | 20.09 | 0 | 0
|  |  | 0.2 | 0.0053563 | 1 | -0.146 | 53.624 | 20.09 | 0 | 0
|  |  | 0.25 | 0.0054414 | 1 | -0.146 | 53.614 | 20.09 | 0 | 0
|  |  | 0.3 | 0.0055265 | 1 | -0.146 | 53.604 | 20.09 | 0 | 0
|  |  | 0.35 | 0.0056115 | 1 | -0.146 | 53.594 | 20.09 | 0 | 0
|  |  | 0.4 | 0.0056966 | 1 | -0.146 | 53.584 | 20.09 | 0 | 0
|  |  | 0.45 | 0.0057696 | 1 | -0.146 | 53.574 | 20.09 | 0 | 0
|  |  | 0.5 | 0.0058425 | 1 | -0.146 | 53.564 | 20.09 | 0 | 0
|  |  | 0.55 | 0.0059155 | 1 | -0.146 | 53.554 | 20.09 | 0 | 0
|  |  | 0.6 | 0.0059885 | 1 | -0.146 | 53.544 | 20.09 | 0 | 0
|  |  | 0.65 | 0.0060416 | 1 | -0.146 | 53.534 | 20.09 | 0 | 0
|  |  | 0.7 | 0.0060948 | 1 | -0.146 | 53.524 | 20.09 | 0 | 0
|  |  | 0.75 | 0.0061479 | 1 | -0.146 | 53.514 | 20.09 | 0 | 0
|  |  | 0.8 | 0.0062011 | 1 | -0.146 | 53.504 | 20.09 | 0 | 0
|  |  | 0.85 | 0.0063143 | 1 | -0.146 | 53.494 | 20.09 | 0 | 0
|  |  | 0.9 | 0.0064276 | 1 | -0.146 | 53.484 | 20.09 | 0 | 0
|  |  | 0.95 | 0.0065409 | 1 | -0.146 | 53.474 | 20.09 | 0 | 0
|  |  | 1 | 0.0066541 | 1 | -0.146 | 53.464 | 20.09 | 0 | 0
|  |  | 1.05 | 0.0067067 | 1 | -0.146 | 53.454 | 20.09 | 0 | 0
|  |  | 1.1 | 0.0067593 | 1 | -0.146 | 53.444 | 20.09 | 0 | 0
|  |  | 1.15 | 0.0068119 | 1 | -0.146 | 53.434 | 20.09 | 0 | 0
|  |  | 1.2 | 0.0068646 | 1 | -0.146 | 53.424 | 20.09 | 0 | 0
|  |  | 1.25 | 0.0069626 | 1 | -0.146 | 53.414 | 20.09 | 0 | 0
|  |  | 1.3 | 0.0070607 | 1 | -0.146 | 53.404 | 20.09 | 0 | 0
|  |  | 1.35 | 0.0071587 | 1 | -0.146 | 53.394 | 20.09 | 0 | 0
|  |  | 1.4 | 0.0072568 | 1 | -0.146 | 53.384 | 20.09 | 0 | 0
|  |  | 1.45 | 0.0073532 | 1 | -0.146 | 53.374 | 20.09 | 0 | 0
|  |  | 1.5 | 0.0074497 | 1 | -0.146 | 53.364 | 20.09 | 0 | 0
|  |  | 1.55 | 0.0075462 | 1 | -0.146 | 53.354 | 20.09 | 0 | 0
|  |  | 1.6 | 0.0076426 | 1 | -0.146 | 53.344 | 20.09 | 0 | 0
|  |  | 1.65 | 0.007704 | 1 | -0.146 | 53.334 | 20.09 | 0 | 0
|  |  | 1.7 | 0.0077654 | 1 | -0.146 | 53.324 | 20.09 | 0 | 0
|  |  | 1.75 | 0.0078268 | 1 | -0.146 | 53.314 | 20.09 | 0 | 0
|  |  | 1.8 | 0.0078882 | 1 | -0.146 | 53.304 | 20.09 | 0 | 0
|  |  | 1.85 | 0.0080269 | 1 | -0.146 | 53.294 | 20.09 | 0 | 0
|  |  | 1.9 | 0.0081656 | 1 | -0.146 | 53.284 | 20.09 | 0 | 0
|  |  | 1.95 | 0.0083043 | 1 | -0.146 | 53.274 | 20.09 | 0 | 0
|  |  | 2 | 0.0084431 | 1 | -0.146 | 53.264 | 20.09 | 0 | 0
|  |  | 2.05 | 0.0085109 | 1 | -0.146 | 53.254 | 20.09 | 0 | 0
|  |  | 2.1 | 0.0085787 | 1 | -0.146 | 53.244 | 20.09 | 0 | 0
|  |  | 2.15 | 0.0086466 | 1 | -0.146 | 53.234 | 20.09 | 0 | 0
|  |  | 2.2 | 0.0087144 | 1 | -0.146 | 53.224 | 20.09 | 0 | 0
|  |  | 2.25 | 0.0088239 | 1 | -0.146 | 53.214 | 20.09 | 0 | 0
|  |  | 2.3 | 0.0089333 | 1 | -0.146 | 53.204 | 20.09 | 0 | 0
|  |  | 2.35 | 0.0090427 | 1 | -0.146 | 53.194 | 20.09 | 0 | 0
|  |  | 2.4 | 0.0091522 | 1 | -0.146 | 53.184 | 20.09 | 0 | 0
|  |  | 2.45 | 0.0092767 | 1 | -0.146 | 53.174 | 20.09 | 0 | 0
|  |  | 2.5 | 0.0094013 | 1 | -0.146 | 53.164 | 20.09 | 0 | 0
|  |  | 2.55 | 0.0095259 | 1 | -0.146 | 53.154 | 20.09 | 0 | 0
|  |  | 2.6 | 0.0096504 | 1 | -0.146 | 53.144 | 20.09 | 0 | 0
|  |  | 2.65 | 0.0097195 | 1 | -0.146 | 53.134 | 20.09 | 0 | 0
|  |  | 2.7 | 0.0097886 | 1 | -0.146 | 53.124 | 20.09 | 0 | 0
|  |  | 2.75 | 0.0098576 | 1 | -0.146 | 53.114 | 20.089 | 0 | 0
|  |  | 2.8 | 0.0099267 | 1 | -0.146 | 53.104 | 20.089 | 0 | 0
|  |  | 2.85 | 0.010093 | 1 | -0.146 | 53.094 | 20.089 | 0 | 0
|  |  | 2.9 | 0.010258 | 1 | -0.146 | 53.084 | 20.089 | 0 | 0
|  |  | 2.95 | 0.010424 | 1 | -0.146 | 53.074 | 20.089 | 0 | 0
|  |  | 3 | 0.01059 | 1 | -0.146 | 53.064 | 20.089 | 0 | 0
|  |  | 3.05 | 0.010672 | 1 | -0.146 | 53.054 | 20.089 | 0 | 0
|  |  | 3.1 | 0.010755 | 1 | -0.146 | 53.044 | 20.089 | 0 | 0
|  |  | 3.15 | 0.010837 | 1 | -0.146 | 53.034 | 20.089 | 0 | 0
|  |  | 3.2 | 0.010919 | 1 | -0.146 | 53.024 | 20.089 | 0 | 0
|  |  | 3.25 | 0.011035 | 1 | -0.146 | 53.014 | 20.089 | 0 | 0
|  |  | 3.3 | 0.011151 | 1 | -0.146 | 53.004 | 20.089 | 0 | 0
|  |  | 3.35 | 0.011267 | 1 | -0.146 | 52.994 | 20.089 | 0 | 0
|  |  | 3.4 | 0.011383 | 1 | -0.146 | 52.984 | 20.089 | 0 | 0
|  |  | 3.45 | 0.011537 | 1 | -0.146 | 52.974 | 20.089 | 0 | 0
|  |  | 3.5 | 0.01169 | 1 | -0.146 | 52.964 | 20.089 | 0 | 0
|  |  | 3.55 | 0.011843 | 1 | -0.146 | 52.954 | 20.089 | 0 | 0
|  |  | 3.6 | 0.011996 | 1 | -0.146 | 52.944 | 20.089 | 0 | 0
|  |  | 3.65 | 0.01207 | 1 | -0.146 | 52.934 | 20.089 | 0 | 0
|  |  | 3.7 | 0.012144 | 1 | -0.146 | 52.924 | 20.089 | 0 | 0
|  |  | 3.75 | 0.012219 | 1 | -0.146 | 52.914 | 20.089 | 0 | 0
|  |  | 3.8 | 0.012293 | 1 | -0.146 | 52.904 | 20.089 | 0 | 0
|  |  | 3.85 | 0.012479 | 1 | -0.146 | 52.894 | 20.089 | 0 | 0
|  |  | 3.9 | 0.012666 | 1 | -0.146 | 52.884 | 20.089 | 0 | 0
|  |  | 3.95 | 0.012853 | 1 | -0.146 | 52.874 | 20.089 | 0 | 0
|  |  | 4 | 0.013039 | 1 | -0.146 | 52.864 | 20.089 | 0 | 0
|  |  | 4.05 | 0.013144 | 1 | -0.146 | 52.854 | 20.089 | 0 | 0
|  |  | 4.1 | 0.013249 | 1 | -0.146 | 52.844 | 20.089 | 0 | 0
|  |  | 4.15 | 0.013353 | 1 | -0.146 | 52.834 | 20.089 | 0 | 0
|  |  | 4.2 | 0.013458 | 1 | -0.146 | 52.824 | 20.089 | 0 | 0
|  |  | 4.25 | 0.013585 | 1 | -0.146 | 52.814 | 20.089 | 0 | 0
|  |  | 4.3 | 0.013711 | 1 | -0.146 | 52.804 | 20.089 | 0 | 0
|  |  | 4.35 | 0.013838 | 1 | -0.146 | 52.794 | 20.089 | 0 | 0
|  |  | 4.4 | 0.013965 | 1 | -0.146 | 52.784 | 20.089 | 0 | 0
|  |  | 4.45 | 0.014157 | 1 | -0.146 | 52.774 | 20.089 | 0 | 0
|  |  | 4.5 | 0.014349 | 1 | -0.146 | 52.764 | 20.089 | 0 | 0
|  |  | 4.55 | 0.014541 | 1 | -0.146 | 52.754 | 20.089 | 0 | 0
|  |  | 4.6 | 0.014733 | 1 | -0.146 | 52.744 | 20.089 | 0 | 0
|  |  | 4.65 | 0.014823 | 1 | -0.146 | 52.734 | 20.089 | 0 | 0
|  |  | 4.7 | 0.014912 | 1 | -0.146 | 52.724 | 20.089 | 0 | 0
|  |  | 4.75 | 0.015002 | 1 | -0.146 | 52.714 | 20.089 | 0 | 0
|  |  | 4.8 | 0.015092 | 1 | -0.146 | 52.704 | 20.089 | 0 | 0
|  |  | 4.85 | 0.015305 | 1 | -0.146 | 52.694 | 20.089 | 0 | 0
|  |  | 4.9 | 0.015519 | 1 | -0.146 | 52.684 | 20.089 | 0 | 0
|  |  | 4.95 | 0.015732 | 1 | -0.146 | 52.674 | 20.089 | 0 | 0
|  |  | 5 | 0.015946 | 1 | -0.146 | 52.664 | 20.089 | 0 | 0
|  |  | 5.05 | 0.01608 | 1 | -0.146 | 52.654 | 20.089 | 0 | 0
|  |  | 5.1 | 0.016214 | 1 | -0.146 | 52.644 | 20.089 | 0 | 0
|  |  | 5.15 | 0.016348 | 1 | -0.146 | 52.634 | 20.089 | 0 | 0
|  |  | 5.2 | 0.016483 | 1 | -0.146 | 52.624 | 20.089 | 0 | 0
|  |  | 5.25 | 0.016616 | 1 | -0.146 | 52.614 | 20.089 | 0 | 0
|  |  | 5.3 | 0.01675 | 1 | -0.146 | 52.604 | 20.089 | 0 | 0
|  |  | 5.35 | 0.016884 | 1 | -0.146 | 52.594 | 20.089 | 0 | 0
|  |  | 5.4 | 0.017017 | 1 | -0.146 | 52.584 | 20.089 | 0 | 0
|  |  | 5.45 | 0.017253 | 1 | -0.146 | 52.574 | 20.089 | 0 | 0
|  |  | 5.5 | 0.017488 | 1 | -0.146 | 52.564 | 20.089 | 0 | 0
|  |  | 5.55 | 0.017724 | 1 | -0.146 | 52.554 | 20.089 | 0 | 0
|  |  | 5.6 | 0.017959 | 1 | -0.146 | 52.544 | 20.089 | 0 | 0
|  |  | 5.65 | 0.018062 | 1 | -0.146 | 52.534 | 20.089 | 0 | 0
|  |  | 5.7 | 0.018165 | 1 | -0.146 | 52.524 | 20.089 | 0 | 0
|  |  | 5.75 | 0.018269 | 1 | -0.146 | 52.514 | 20.089 | 0 | 0
|  |  | 5.8 | 0.018372 | 1 | -0.146 | 52.504 | 20.089 | 0 | 0
|  |  | 5.85 | 0.018605 | 1 | -0.146 | 52.494 | 20.089 | 0 | 0
|  |  | 5.9 | 0.018838 | 1 | -0.146 | 52.484 | 20.089 | 0 | 0
|  |  | 5.95 | 0.019071 | 1 | -0.146 | 52.474 | 20.089 | 0 | 0
|  |  | 6 | 0.019304 | 1 | -0.146 | 52.464 | 20.089 | 0 | 0
|  |  | 6.05 | 0.019473 | 1 | -0.146 | 52.454 | 20.089 | 0 | 0
|  |  | 6.1 | 0.019642 | 1 | -0.146 | 52.444 | 20.089 | 0 | 0
|  |  | 6.15 | 0.019811 | 1 | -0.146 | 52.434 | 20.089 | 0 | 0
|  |  | 6.2 | 0.01998 | 1 | -0.146 | 52.424 | 20.088 | 0 | 0
|  |  | 6.25 | 0.02012 | 1 | -0.146 | 52.414 | 20.088 | 0 | 0
|  |  | 6.3 | 0.020261 | 1 | -0.146 | 52.404 | 20.088 | 0 | 0
|  |  | 6.35 | 0.020402 | 1 | -0.146 | 52.394 | 20.088 | 0 | 0
|  |  | 6.4 | 0.020542 | 1 | -0.146 | 52.384 | 20.088 | 0 | 0
|  |  | 6.45 | 0.020822 | 1 | -0.146 | 52.374 | 20.088 | 0 | 0
|  |  | 6.5 | 0.021102 | 1 | -0.146 | 52.364 | 20.088 | 0 | 0
|  |  | 6.55 | 0.021382 | 1 | -0.146 | 52.354 | 20.088 | 0 | 0
|  |  | 6.6 | 0.021662 | 1 | -0.146 | 52.344 | 20.088 | 0 | 0
|  |  | 6.65 | 0.02178 | 1 | -0.146 | 52.334 | 20.088 | 0 | 0
|  |  | 6.7 | 0.021899 | 1 | -0.146 | 52.324 | 20.088 | 0 | 0
|  |  | 6.75 | 0.022017 | 1 | -0.146 | 52.314 | 20.088 | 0 | 0
|  |  | 6.8 | 0.022135 | 1 | -0.146 | 52.304 | 20.088 | 0 | 0
|  |  | 6.85 | 0.022384 | 1 | -0.146 | 52.294 | 20.088 | 0 | 0
|  |  | 6.9 | 0.022633 | 1 | -0.146 | 52.284 | 20.088 | 0 | 0
|  |  | 6.95 | 0.022882 | 1 | -0.146 | 52.274 | 20.088 | 0 | 0
|  |  | 7 | 0.02313 | 1 | -0.146 | 52.264 | 20.088 | 0 | 0
|  |  | 7.05 | 0.023338 | 1 | -0.146 | 52.254 | 20.088 | 0 | 0
|  |  | 7.1 | 0.023546 | 1 | -0.146 | 52.244 | 20.088 | 0 | 0
|  |  | 7.15 | 0.023754 | 1 | -0.146 | 52.234 | 20.088 | 0 | 0
|  |  | 7.2 | 0.023961 | 1 | -0.146 | 52.224 | 20.088 | 0 | 0
|  |  | 7.25 | 0.024106 | 1 | -0.146 | 52.214 | 20.088 | 0 | 0
|  |  | 7.3 | 0.024251 | 1 | -0.146 | 52.204 | 20.088 | 0 | 0
|  |  | 7.35 | 0.024397 | 1 | -0.146 | 52.194 | 20.088 | 0 | 0
|  |  | 7.4 | 0.024542 | 1 | -0.146 | 52.184 | 20.088 | 0 | 0
|  |  | 7.45 | 0.024864 | 1 | -0.146 | 52.174 | 20.088 | 0 | 0
|  |  | 7.5 | 0.025185 | 1 | -0.146 | 52.164 | 20.088 | 0 | 0
|  |  | 7.55 | 0.025507 | 1 | -0.146 | 52.154 | 20.088 | 0 | 0
|  |  | 7.6 | 0.025829 | 1 | -0.146 | 52.144 | 20.088 | 0 | 0
|  |  | 7.65 | 0.025966 | 1 | -0.146 | 52.134 | 20.088 | 0 | 0
|  |  | 7.7 | 0.026103 | 1 | -0.146 | 52.124 | 20.088 | 0 | 0
|  |  | 7.75 | 0.02624 | 1 | -0.146 | 52.114 | 20.088 | 0 | 0
|  |  | 7.8 | 0.026377 | 1 | -0.146 | 52.104 | 20.088 | 0 | 0
|  |  | 7.85 | 0.026638 | 1 | -0.146 | 52.094 | 20.088 | 0 | 0
|  |  | 7.9 | 0.026899 | 1 | -0.146 | 52.084 | 20.088 | 0 | 0
|  |  | 7.95 | 0.02716 | 1 | -0.146 | 52.074 | 20.088 | 0 | 0
|  |  | 8 | 0.027421 | 1 | -0.146 | 52.064 | 20.088 | 0 | 0
|  |  | 8.05 | 0.027674 | 1 | -0.146 | 52.054 | 20.088 | 0 | 0
|  |  | 8.1 | 0.027926 | 1 | -0.146 | 52.044 | 20.088 | 0 | 0
|  |  | 8.15 | 0.028179 | 1 | -0.146 | 52.034 | 20.088 | 0 | 0
|  |  | 8.2 | 0.028432 | 1 | -0.146 | 52.024 | 20.088 | 0 | 0
|  |  | 8.25 | 0.02858 | 1 | -0.146 | 52.014 | 20.088 | 0 | 0
|  |  | 8.3 | 0.028729 | 1 | -0.146 | 52.004 | 20.088 | 0 | 0
|  |  | 8.35 | 0.028878 | 1 | -0.146 | 51.994 | 20.088 | 0 | 0
|  |  | 8.4 | 0.029026 | 1 | -0.146 | 51.984 | 20.088 | 0 | 0
|  |  | 8.45 | 0.029388 | 1 | -0.146 | 51.975 | 20.088 | 0 | 0
|  |  | 8.5 | 0.02975 | 1 | -0.146 | 51.965 | 20.088 | 0 | 0
|  |  | 8.55 | 0.030112 | 1 | -0.146 | 51.955 | 20.088 | 0 | 0
|  |  | 8.6 | 0.030474 | 1 | -0.146 | 51.945 | 20.088 | 0 | 0
|  |  | 8.65 | 0.03063 | 1 | -0.146 | 51.935 | 20.088 | 0 | 0
|  |  | 8.7 | 0.030786 | 1 | -0.146 | 51.925 | 20.088 | 0 | 0
|  |  | 8.75 | 0.030941 | 1 | -0.146 | 51.915 | 20.088 | 0 | 0
|  |  | 8.8 | 0.031097 | 1 | -0.146 | 51.905 | 20.088 | 0 | 0
|  |  | 8.85 | 0.031367 | 1 | -0.146 | 51.895 | 20.088 | 0 | 0
|  |  | 8.9 | 0.031636 | 1 | -0.146 | 51.885 | 20.088 | 0 | 0
|  |  | 8.95 | 0.031906 | 1 | -0.146 | 51.875 | 20.088 | 0 | 0
|  |  | 9 | 0.032175 | 1 | -0.146 | 51.865 | 20.088 | 0 | 0
|  |  | 9.05 | 0.032468 | 1 | -0.146 | 51.855 | 20.088 | 0 | 0
|  |  | 9.1 | 0.032761 | 1 | -0.146 | 51.845 | 20.088 | 0 | 0
|  |  | 9.15 | 0.033053 | 1 | -0.146 | 51.835 | 20.088 | 0 | 0
|  |  | 9.2 | 0.033346 | 1 | -0.146 | 51.825 | 20.088 | 0 | 0
|  |  | 9.25 | 0.1029 | 1 | -0.146 | 51.815 | 20.088 | 0 | 0
|  |  | 9.3 | 0.17246 | 1 | -0.146 | 51.805 | 20.088 | 0 | 0
|  |  | 9.35 | 0.24202 | 1 | -0.146 | 51.795 | 20.088 | 0 | 0
|  |  | 9.4 | 0.31157 | 1 | -0.146 | 51.785 | 20.088 | 0 | 0
|  |  | 9.45 | 0.34656 | 1 | -0.146 | 51.775 | 20.088 | 0 | 0
|  |  | 9.5 | 0.38155 | 1 | -0.146 | 51.765 | 20.088 | 0 | 0
|  |  | 9.55 | 0.41654 | 1 | -0.146 | 51.755 | 20.088 | 0 | 0
|  |  | 9.6 | 0.45153 | 1 | -0.146 | 51.745 | 20.088 | 0 | 0
|  |  | 9.65 | 0.48656 | 1 | -0.146 | 51.735 | 20.087 | 0 | 0
|  |  | 9.7 | 0.52158 | 1 | -0.146 | 51.725 | 20.087 | 0 | 0
|  |  | 9.75 | 0.55661 | 1 | -0.146 | 51.715 | 20.087 | 0 | 0
|  |  | 9.8 | 0.59163 | 1 | -0.146 | 51.705 | 20.087 | 0 | 0
|  |  | 9.85 | 0.59785 | 1 | -0.146 | 51.695 | 20.087 | 0 | 0
|  |  | 9.9 | 0.60407 | 1 | -0.146 | 51.685 | 20.087 | 0 | 0
|  |  | 9.95 | 0.61029 | 1 | -0.146 | 51.675 | 20.087 | 0 | 0
|  |  | 10 | 0.61651 | 1 | -0.146 | 51.665 | 20.087 | 0 | 0
|  |  | 10.05 | 0.62273 | 1 | -0.146 | 51.655 | 20.087 | 0 | 0
|  |  | 10.1 | 0.62896 | 1 | -0.146 | 51.645 | 20.087 | 0 | 0
|  |  | 10.15 | 0.63518 | 1 | -0.146 | 51.635 | 20.087 | 0 | 0
|  |  | 10.2 | 0.6414 | 1 | -0.146 | 51.625 | 20.087 | 0 | 0
|  |  | 10.25 | 0.64763 | 1 | -0.146 | 51.615 | 20.087 | 0 | 0
|  |  | 10.3 | 0.65386 | 1 | -0.146 | 51.605 | 20.087 | 0 | 0
|  |  | 10.35 | 0.66009 | 1 | -0.146 | 51.595 | 20.087 | 0 | 0
|  |  | 10.4 | 0.66632 | 1 | -0.146 | 51.585 | 20.087 | 0 | 0
|  |  | 10.45 | 0.67255 | 1 | -0.146 | 51.575 | 20.087 | 0 | 0
|  |  | 10.5 | 0.67878 | 1 | -0.146 | 51.565 | 20.087 | 0 | 0
|  |  | 10.55 | 0.68502 | 1 | -0.146 | 51.555 | 20.087 | 0 | 0
|  |  | 10.6 | 0.69125 | 1 | -0.146 | 51.545 | 20.087 | 0 | 0
|  |  | 10.65 | 0.69915 | 1 | -0.146 | 51.535 | 20.087 | 0 | 0
|  |  | 10.7 | 0.70704 | 1 | -0.146 | 51.525 | 20.087 | 0 | 0
|  |  | 10.75 | 0.71493 | 1 | -0.146 | 51.515 | 20.087 | 0 | 0
|  |  | 10.8 | 0.72283 | 1 | -0.146 | 51.505 | 20.087 | 0 | 0
|  |  | 10.85 | 0.73073 | 1 | -0.146 | 51.495 | 20.087 | 0 | 0
|  |  | 10.9 | 0.73863 | 1 | -0.146 | 51.485 | 20.087 | 0 | 0
|  |  | 10.95 | 0.74653 | 1 | -0.146 | 51.475 | 20.087 | 0 | 0
|  |  | 11 | 0.75443 | 1 | -0.146 | 51.465 | 20.087 | 0 | 0
|  |  | 11.05 | 0.76234 | 1 | -0.146 | 51.455 | 20.087 | 0 | 0
|  |  | 11.1 | 0.77025 | 1 | -0.146 | 51.445 | 20.087 | 0 | 0
|  |  | 11.15 | 0.77815 | 1 | -0.146 | 51.435 | 20.087 | 0 | 0
|  |  | 11.2 | 0.78606 | 1 | -0.146 | 51.425 | 20.087 | 0 | 0
|  |  | 11.25 | 0.79397 | 1 | -0.146 | 51.415 | 20.087 | 0 | 0
|  |  | 11.3 | 0.80189 | 1 | -0.146 | 51.405 | 20.087 | 0 | 0
|  |  | 11.35 | 0.8098 | 1 | -0.146 | 51.395 | 20.087 | 0 | 0
|  |  | 11.4 | 0.81771 | 1 | -0.146 | 51.385 | 20.087 | 0 | 0
|  |  | 11.45 | 0.82702 | 1 | -0.146 | 51.375 | 20.087 | 0 | 0
|  |  | 11.5 | 0.83634 | 1 | -0.146 | 51.365 | 20.087 | 0 | 0
|  |  | 11.55 | 0.84565 | 1 | -0.146 | 51.355 | 20.087 | 0 | 0
|  |  | 11.6 | 0.85497 | 1 | -0.146 | 51.345 | 20.087 | 0 | 0
|  |  | 11.65 | 0.86429 | 1 | -0.146 | 51.335 | 20.087 | 0 | 0
|  |  | 11.7 | 0.87361 | 1 | -0.146 | 51.325 | 20.087 | 0 | 0
|  |  | 11.75 | 0.88293 | 1 | -0.146 | 51.315 | 20.087 | 0 | 0
|  |  | 11.8 | 0.89225 | 1 | -0.146 | 51.305 | 20.087 | 0 | 0
|  |  | 11.85 | 0.90157 | 1 | -0.146 | 51.295 | 20.087 | 0 | 0
|  |  | 11.9 | 0.9109 | 1 | -0.146 | 51.285 | 20.087 | 0 | 0
|  |  | 11.95 | 0.92023 | 1 | -0.146 | 51.275 | 20.087 | 0 | 0
|  |  | 12 | 0.92955 | 1 | -0.146 | 51.265 | 20.087 | 0 | 0
|  |  | 12.05 | 0.93889 | 1 | -0.146 | 51.255 | 20.087 | 0 | 0
|  |  | 12.1 | 0.94822 | 1 | -0.146 | 51.245 | 20.087 | 0 | 0
|  |  | 12.15 | 0.95756 | 1 | -0.146 | 51.235 | 20.087 | 0 | 0
|  |  | 12.2 | 0.96689 | 1 | -0.146 | 51.225 | 20.087 | 0 | 0
|  |  | 12.25 | 0.97711 | 1 | -0.146 | 51.215 | 20.087 | 0 | 0
|  |  | 12.3 | 0.98732 | 1 | -0.146 | 51.205 | 20.087 | 0 | 0
|  |  | 12.35 | 0.99753 | 1 | -0.146 | 51.195 | 20.087 | 0 | 0
|  |  | 12.4 | 1.0077 | 1 | -0.146 | 51.185 | 20.087 | 0 | 0
|  |  | 12.45 | 1.018 | 1 | -0.146 | 51.175 | 20.087 | 0 | 0
|  |  | 12.5 | 1.0282 | 1 | -0.146 | 51.165 | 20.087 | 0 | 0
|  |  | 12.55 | 1.0384 | 1 | -0.146 | 51.155 | 20.087 | 0 | 0
|  |  | 12.6 | 1.0486 | 1 | -0.146 | 51.145 | 20.087 | 0 | 0
|  |  | 12.65 | 1.0589 | 1 | -0.146 | 51.135 | 20.087 | 0 | 0
|  |  | 12.7 | 1.0691 | 1 | -0.146 | 51.125 | 20.087 | 0 | 0
|  |  | 12.75 | 1.0793 | 1 | -0.146 | 51.115 | 20.087 | 0 | 0
|  |  | 12.8 | 1.0896 | 1 | -0.146 | 51.105 | 20.087 | 0 | 0
|  |  | 12.85 | 1.0998 | 1 | -0.146 | 51.095 | 20.087 | 0 | 0
|  |  | 12.9 | 1.11 | 1 | -0.146 | 51.085 | 20.087 | 0 | 0
|  |  | 12.95 | 1.1203 | 1 | -0.146 | 51.075 | 20.087 | 0 | 0
|  |  | 13 | 1.1305 | 1 | -0.146 | 51.065 | 20.087 | 0 | 0
|  |  | 13.05 | 1.1437 | 1 | -0.146 | 51.055 | 20.087 | 0 | 0
|  |  | 13.1 | 1.157 | 1 | -0.146 | 51.045 | 20.086 | 0 | 0
|  |  | 13.15 | 1.1702 | 1 | -0.146 | 51.035 | 20.086 | 0 | 0
|  |  | 13.2 | 1.1834 | 1 | -0.146 | 51.025 | 20.086 | 0 | 0
|  |  | 13.25 | 1.1967 | 1 | -0.146 | 51.015 | 20.086 | 0 | 0
|  |  | 13.3 | 1.2099 | 1 | -0.146 | 51.005 | 20.086 | 0 | 0
|  |  | 13.35 | 1.2232 | 1 | -0.146 | 50.995 | 20.086 | 0 | 0
|  |  | 13.4 | 1.2364 | 1 | -0.146 | 50.985 | 20.086 | 0 | 0
|  |  | 13.45 | 1.2497 | 1 | -0.146 | 50.975 | 20.086 | 0 | 0
|  |  | 13.5 | 1.2629 | 1 | -0.146 | 50.965 | 20.086 | 0 | 0
|  |  | 13.55 | 1.2762 | 1 | -0.146 | 50.955 | 20.086 | 0 | 0
|  |  | 13.6 | 1.2894 | 1 | -0.146 | 50.945 | 20.086 | 0 | 0
|  |  | 13.65 | 1.3027 | 1 | -0.146 | 50.935 | 20.086 | 0 | 0
|  |  | 13.7 | 1.316 | 1 | -0.146 | 50.925 | 20.086 | 0 | 0
|  |  | 13.75 | 1.3292 | 1 | -0.146 | 50.915 | 20.086 | 0 | 0
|  |  | 13.8 | 1.3425 | 1 | -0.146 | 50.905 | 20.086 | 0 | 0
|  |  | 13.85 | 1.3733 | 1 | -0.146 | 50.895 | 20.086 | 0 | 0
|  |  | 13.9 | 1.4042 | 1 | -0.146 | 50.885 | 20.086 | 0 | 0
|  |  | 13.95 | 1.435 | 1 | -0.146 | 50.875 | 20.086 | 0 | 0
|  |  | 14 | 1.4658 | 1 | -0.146 | 50.865 | 20.086 | 0 | 0
|  |  | 14.05 | 1.4967 | 1 | -0.146 | 50.855 | 20.086 | 0 | 0
|  |  | 14.1 | 1.5276 | 1 | -0.146 | 50.845 | 20.086 | 0 | 0
|  |  | 14.15 | 1.5584 | 1 | -0.146 | 50.835 | 20.086 | 0 | 0
|  |  | 14.2 | 1.5893 | 1 | -0.146 | 50.825 | 20.086 | 0 | 0
|  |  | 14.25 | 1.6202 | 1 | -0.146 | 50.815 | 20.086 | 0 | 0
|  |  | 14.3 | 1.651 | 1 | -0.146 | 50.805 | 20.086 | 0 | 0
|  |  | 14.35 | 1.6819 | 1 | -0.146 | 50.795 | 20.086 | 0 | 0
|  |  | 14.4 | 1.7128 | 1 | -0.146 | 50.785 | 20.086 | 0 | 0
|  |  | 14.45 | 1.7437 | 1 | -0.146 | 50.775 | 20.086 | 0 | 0
|  |  | 14.5 | 1.7746 | 1 | -0.146 | 50.765 | 20.086 | 0 | 0
|  |  | 14.55 | 1.8055 | 1 | -0.146 | 50.755 | 20.086 | 0 | 0
|  |  | 14.6 | 1.8364 | 1 | -0.146 | 50.745 | 20.086 | 0 | 0
|  |  | 14.65 | 1.8699 | 1 | -0.146 | 50.735 | 20.086 | 0 | 0
|  |  | 14.7 | 1.9033 | 1 | -0.146 | 50.725 | 20.086 | 0 | 0
|  |  | 14.75 | 1.9368 | 1 | -0.146 | 50.715 | 20.086 | 0 | 0
|  |  | 14.8 | 1.9702 | 1 | -0.146 | 50.705 | 20.086 | 0 | 0
|  |  | 14.85 | 2.0037 | 1 | -0.146 | 50.695 | 20.086 | 0 | 0
|  |  | 14.9 | 2.0372 | 1 | -0.146 | 50.685 | 20.086 | 0 | 0
|  |  | 14.95 | 2.0706 | 1 | -0.146 | 50.675 | 20.086 | 0 | 0
|  |  | 15 | 2.1041 | 1 | -0.146 | 50.665 | 20.086 | 0 | 0
|  |  | 15.05 | 2.1376 | 1 | -0.146 | 50.655 | 20.086 | 0 | 0
|  |  | 15.1 | 2.1711 | 1 | -0.146 | 50.645 | 20.086 | 0 | 0
|  |  | 15.15 | 2.2046 | 1 | -0.146 | 50.635 | 20.086 | 0 | 0
|  |  | 15.2 | 2.2381 | 1 | -0.146 | 50.625 | 20.086 | 0 | 0
|  |  | 15.25 | 2.2716 | 1 | -0.146 | 50.615 | 20.086 | 0 | 0
|  |  | 15.3 | 2.3051 | 1 | -0.146 | 50.605 | 20.086 | 0 | 0
|  |  | 15.35 | 2.3386 | 1 | -0.146 | 50.595 | 20.086 | 0 | 0
|  |  | 15.4 | 2.3721 | 1 | -0.146 | 50.585 | 20.086 | 0 | 0
|  |  | 15.45 | 2.4098 | 1 | -0.146 | 50.575 | 20.086 | 0 | 0
|  |  | 15.5 | 2.4474 | 1 | -0.146 | 50.565 | 20.086 | 0 | 0
|  |  | 15.55 | 2.4851 | 1 | -0.146 | 50.555 | 20.086 | 0 | 0
|  |  | 15.6 | 2.5228 | 1 | -0.146 | 50.545 | 20.086 | 0 | 0
|  |  | 15.65 | 2.5605 | 1 | -0.146 | 50.535 | 20.086 | 0 | 0
|  |  | 15.7 | 2.5982 | 1 | -0.146 | 50.525 | 20.086 | 0 | 0
|  |  | 15.75 | 2.6359 | 1 | -0.146 | 50.515 | 20.086 | 0 | 0
|  |  | 15.8 | 2.6736 | 1 | -0.146 | 50.505 | 20.086 | 0 | 0
|  |  | 15.85 | 2.7113 | 1 | -0.146 | 50.495 | 20.086 | 0 | 0
|  |  | 15.9 | 2.749 | 1 | -0.146 | 50.485 | 20.086 | 0 | 0
|  |  | 15.95 | 2.7867 | 1 | -0.146 | 50.475 | 20.086 | 0 | 0
|  |  | 16 | 2.8244 | 1 | -0.146 | 50.465 | 20.086 | 0 | 0
|  |  | 16.05 | 2.8622 | 1 | -0.146 | 50.455 | 20.086 | 0 | 0
|  |  | 16.1 | 2.8999 | 1 | -0.146 | 50.445 | 20.086 | 0 | 0
|  |  | 16.15 | 2.9377 | 1 | -0.146 | 50.435 | 20.086 | 0 | 0
|  |  | 16.2 | 2.9754 | 1 | -0.146 | 50.425 | 20.086 | 0 | 0
|  |  | 16.25 | 3.0248 | 1 | -0.146 | 50.415 | 20.086 | 0 | 0
|  |  | 16.3 | 3.0743 | 1 | -0.146 | 50.405 | 20.086 | 0 | 0
|  |  | 16.35 | 3.1237 | 1 | -0.146 | 50.395 | 20.086 | 0 | 0
|  |  | 16.4 | 3.1731 | 1 | -0.146 | 50.385 | 20.086 | 0 | 0
|  |  | 16.45 | 3.2226 | 1 | -0.146 | 50.375 | 20.086 | 0 | 0
|  |  | 16.5 | 3.272 | 1 | -0.146 | 50.365 | 20.086 | 0 | 0
|  |  | 16.55 | 3.3215 | 1 | -0.146 | 50.355 | 20.085 | 0 | 0
|  |  | 16.6 | 3.371 | 1 | -0.146 | 50.345 | 20.085 | 0 | 0
|  |  | 16.65 | 3.4204 | 1 | -0.146 | 50.335 | 20.085 | 0 | 0
|  |  | 16.7 | 3.4699 | 1 | -0.146 | 50.325 | 20.085 | 0 | 0
|  |  | 16.75 | 3.5194 | 1 | -0.146 | 50.315 | 20.085 | 0 | 0
|  |  | 16.8 | 3.5689 | 1 | -0.146 | 50.305 | 20.085 | 0 | 0
|  |  | 16.85 | 3.6184 | 1 | -0.146 | 50.295 | 20.085 | 0 | 0
|  |  | 16.9 | 3.6679 | 1 | -0.146 | 50.285 | 20.085 | 0 | 0
|  |  | 16.95 | 3.7174 | 1 | -0.146 | 50.275 | 20.085 | 0 | 0
|  |  | 17 | 3.7669 | 1 | -0.146 | 50.265 | 20.085 | 0 | 0
|  |  | 17.05 | 3.8254 | 1 | -0.146 | 50.255 | 20.085 | 0 | 0
|  |  | 17.1 | 3.8839 | 1 | -0.146 | 50.245 | 20.085 | 0 | 0
|  |  | 17.15 | 3.9424 | 1 | -0.146 | 50.235 | 20.085 | 0 | 0
|  |  | 17.2 | 4.001 | 1 | -0.146 | 50.225 | 20.085 | 0 | 0
|  |  | 17.25 | 4.0595 | 1 | -0.146 | 50.215 | 20.085 | 0 | 0
|  |  | 17.3 | 4.118 | 1 | -0.146 | 50.205 | 20.085 | 0 | 0
|  |  | 17.35 | 4.1766 | 1 | -0.146 | 50.195 | 20.085 | 0 | 0
|  |  | 17.4 | 4.2351 | 1 | -0.146 | 50.185 | 20.085 | 0 | 0
|  |  | 17.45 | 4.2937 | 1 | -0.146 | 50.175 | 20.085 | 0 | 0
|  |  | 17.5 | 4.3522 | 1 | -0.146 | 50.165 | 20.085 | 0 | 0
|  |  | 17.55 | 4.4108 | 1 | -0.146 | 50.155 | 20.085 | 0 | 0
|  |  | 17.6 | 4.4694 | 1 | -0.146 | 50.145 | 20.085 | 0 | 0
|  |  | 17.65 | 4.5279 | 1 | -0.146 | 50.135 | 20.085 | 0 | 0
|  |  | 17.7 | 4.5865 | 1 | -0.146 | 50.125 | 20.085 | 0 | 0
|  |  | 17.75 | 4.6451 | 1 | -0.146 | 50.115 | 20.085 | 0 | 0
|  |  | 17.8 | 4.7037 | 1 | -0.146 | 50.105 | 20.085 | 0 | 0
|  |  | 17.85 | 4.7896 | 1 | -0.146 | 50.095 | 20.085 | 0 | 0
|  |  | 17.9 | 4.8755 | 1 | -0.146 | 50.085 | 20.085 | 0 | 0
|  |  | 17.95 | 4.9614 | 1 | -0.146 | 50.075 | 20.085 | 0 | 0
|  |  | 18 | 5.0473 | 1 | -0.146 | 50.065 | 20.085 | 0 | 0
|  |  | 18.05 | 5.1332 | 1 | -0.146 | 50.055 | 20.085 | 0 | 0
|  |  | 18.1 | 5.2192 | 1 | -0.146 | 50.045 | 20.085 | 0 | 0
|  |  | 18.15 | 5.3051 | 1 | -0.146 | 50.035 | 20.085 | 0 | 0
|  |  | 18.2 | 5.3911 | 1 | -0.146 | 50.025 | 20.085 | 0 | 0
|  |  | 18.25 | 5.4775 | 1 | -0.146 | 50.015 | 20.085 | 0 | 0
|  |  | 18.3 | 5.5638 | 1 | -0.146 | 50.005 | 20.085 | 0 | 0
|  |  | 18.35 | 5.6502 | 1 | -0.146 | 49.995 | 20.085 | 0 | 0
|  |  | 18.4 | 5.7366 | 1 | -0.146 | 49.985 | 20.085 | 0 | 0
|  |  | 18.45 | 5.8231 | 1 | -0.146 | 49.975 | 20.085 | 0 | 0
|  |  | 18.5 | 5.9095 | 1 | -0.146 | 49.965 | 20.085 | 0 | 0
|  |  | 18.55 | 5.9959 | 1 | -0.146 | 49.955 | 20.085 | 0 | 0
|  |  | 18.6 | 6.0823 | 1 | -0.146 | 49.945 | 20.085 | 0 | 0
|  |  | 18.65 | 6.2732 | 1 | -0.146 | 49.935 | 20.085 | 0 | 0
|  |  | 18.7 | 6.464 | 1 | -0.146 | 49.925 | 20.085 | 0 | 0
|  |  | 18.75 | 6.6548 | 1 | -0.146 | 49.915 | 20.085 | 0 | 0
|  |  | 18.8 | 6.8456 | 1 | -0.146 | 49.905 | 20.085 | 0 | 0
|  |  | 18.85 | 7.0374 | 1 | -0.146 | 49.895 | 20.085 | 0 | 0
|  |  | 18.9 | 7.2292 | 1 | -0.146 | 49.885 | 20.085 | 0 | 0
|  |  | 18.95 | 7.421 | 1 | -0.146 | 49.875 | 20.085 | 0 | 0
|  |  | 19 | 7.6128 | 1 | -0.146 | 49.865 | 20.085 | 0 | 0
|  |  | 19.05 | 8.062 | 1 | -0.146 | 49.855 | 20.085 | 0 | 0
|  |  | 19.1 | 8.5112 | 1 | -0.146 | 49.845 | 20.085 | 0 | 0
|  |  | 19.15 | 8.9604 | 1 | -0.146 | 49.835 | 20.085 | 0 | 0
|  |  | 19.2 | 9.4096 | 1 | -0.146 | 49.825 | 20.085 | 0 | 0
|  |  | 19.25 | 10.901 | 1 | -0.146 | 49.815 | 20.085 | 0 | 0
|  |  | 19.3 | 12.392 | 1 | -0.146 | 49.805 | 20.085 | 0 | 0
|  |  | 19.35 | 13.883 | 1 | -0.146 | 49.795 | 20.085 | 0 | 0
|  |  | 19.4 | 15.374 | 1 | -0.146 | 49.785 | 20.085 | 0 | 0
|  |  | 19.45 | 20.112 | 1 | -0.146 | 49.775 | 20.085 | 0 | 0
|  |  | 19.5 | 24.851 | 1 | -0.146 | 49.765 | 20.085 | 0 | 0
|  |  | 19.55 | 29.589 | 1 | -0.146 | 49.755 | 20.085 | 0 | 0
|  |  | 19.6 | 34.327 | 1 | -0.146 | 49.745 | 20.085 | 0 | 0
|  |  | 19.65 | 36.914 | 1 | -0.146 | 49.735 | 20.085 | 0 | 0
|  |  | 19.7 | 39.502 | 1 | -0.146 | 49.725 | 20.085 | 0 | 0
|  |  | 19.75 | 42.09 | 1 | -0.146 | 49.715 | 20.085 | 0 | 0
|  |  | 19.8 | 44.677 | 1 | -0.146 | 49.705 | 20.085 | 0 | 0
|  |  | 19.85 | 45.44 | 1 | -0.146 | 49.695 | 20.085 | 0 | 0
|  |  | 19.9 | 46.203 | 1 | -0.146 | 49.685 | 20.085 | 0 | 0
|  |  | 19.95 | 46.966 | 1 | -0.146 | 49.675 | 20.085 | 0 | 0
|  |  | 20 | 47.729 | 1 | -0.146 | 49.665 | 20.084 | 0 | 0
|  |  | 20.05 | 48.046 | 1 | -0.146 | 49.655 | 20.084 | 0 | 0
|  |  | 20.1 | 48.363 | 1 | -0.146 | 49.645 | 20.084 | 0 | 0
|  |  | 20.15 | 48.681 | 1 | -0.146 | 49.635 | 20.084 | 0 | 0
|  |  | 20.2 | 48.998 | 1 | -0.146 | 49.625 | 20.084 | 0 | 0
|  |  | 20.25 | 49.168 | 1 | -0.146 | 49.615 | 20.084 | 0 | 0
|  |  | 20.3 | 49.338 | 1 | -0.146 | 49.605 | 20.084 | 0 | 0
|  |  | 20.35 | 49.508 | 1 | -0.146 | 49.595 | 20.084 | 0 | 0
|  |  | 20.4 | 49.678 | 1 | -0.146 | 49.585 | 20.084 | 0 | 0
|  |  | 20.45 | 49.809 | 1 | -0.146 | 49.575 | 20.084 | 0 | 0
|  |  | 20.5 | 49.94 | 1 | -0.146 | 49.565 | 20.084 | 0 | 0
|  |  | 20.55 | 50.072 | 1 | -0.146 | 49.555 | 20.084 | 0 | 0
|  |  | 20.6 | 50.203 | 1 | -0.146 | 49.545 | 20.084 | 0 | 0
|  |  | 20.65 | 50.301 | 1 | -0.146 | 49.535 | 20.084 | 0 | 0
|  |  | 20.7 | 50.399 | 1 | -0.146 | 49.525 | 20.084 | 0 | 0
|  |  | 20.75 | 50.497 | 1 | -0.146 | 49.515 | 20.084 | 0 | 0
|  |  | 20.8 | 50.594 | 1 | -0.146 | 49.505 | 20.084 | 0 | 0
|  |  | 20.85 | 50.688 | 1 | -0.146 | 49.495 | 20.084 | 0 | 0
|  |  | 20.9 | 50.782 | 1 | -0.146 | 49.485 | 20.084 | 0 | 0
|  |  | 20.95 | 50.876 | 1 | -0.146 | 49.475 | 20.084 | 0 | 0
|  |  | 21 | 50.969 | 1 | -0.146 | 49.465 | 20.084 | 0 | 0
|  |  | 21.05 | 51.021 | 1 | -0.146 | 49.455 | 20.084 | 0 | 0
|  |  | 21.1 | 51.072 | 1 | -0.146 | 49.445 | 20.084 | 0 | 0
|  |  | 21.15 | 51.124 | 1 | -0.146 | 49.435 | 20.084 | 0 | 0
|  |  | 21.2 | 51.175 | 1 | -0.146 | 49.425 | 20.084 | 0 | 0
|  |  | 21.25 | 51.227 | 1 | -0.146 | 49.415 | 20.084 | 0 | 0
|  |  | 21.3 | 51.278 | 1 | -0.146 | 49.405 | 20.084 | 0 | 0
|  |  | 21.35 | 51.329 | 1 | -0.146 | 49.395 | 20.084 | 0 | 0
|  |  | 21.4 | 51.381 | 1 | -0.146 | 49.385 | 20.084 | 0 | 0
|  |  | 21.45 | 51.432 | 1 | -0.146 | 49.375 | 20.084 | 0 | 0
|  |  | 21.5 | 51.483 | 1 | -0.146 | 49.365 | 20.084 | 0 | 0
|  |  | 21.55 | 51.535 | 1 | -0.146 | 49.355 | 20.084 | 0 | 0
|  |  | 21.6 | 51.586 | 1 | -0.146 | 49.345 | 20.084 | 0 | 0
|  |  | 21.65 | 51.637 | 1 | -0.146 | 49.335 | 20.084 | 0 | 0
|  |  | 21.7 | 51.688 | 1 | -0.146 | 49.325 | 20.084 | 0 | 0
|  |  | 21.75 | 51.74 | 1 | -0.146 | 49.315 | 20.084 | 0 | 0
|  |  | 21.8 | 51.791 | 1 | -0.146 | 49.305 | 20.084 | 0 | 0
|  |  | 21.85 | 51.817 | 1 | -0.146 | 49.295 | 20.084 | 0 | 0
|  |  | 21.9 | 51.842 | 1 | -0.146 | 49.285 | 20.084 | 0 | 0
|  |  | 21.95 | 51.868 | 1 | -0.146 | 49.275 | 20.084 | 0 | 0
|  |  | 22 | 51.894 | 1 | -0.146 | 49.265 | 20.084 | 0 | 0
|  |  | 22.05 | 51.919 | 1 | -0.146 | 49.255 | 20.084 | 0 | 0
|  |  | 22.1 | 51.945 | 1 | -0.146 | 49.245 | 20.084 | 0 | 0
|  |  | 22.15 | 51.971 | 1 | -0.146 | 49.235 | 20.084 | 0 | 0
|  |  | 22.2 | 51.996 | 1 | -0.146 | 49.225 | 20.084 | 0 | 0
|  |  | 22.25 | 52.022 | 1 | -0.146 | 49.215 | 20.084 | 0 | 0
|  |  | 22.3 | 52.048 | 1 | -0.146 | 49.205 | 20.084 | 0 | 0
|  |  | 22.35 | 52.073 | 1 | -0.146 | 49.195 | 20.084 | 0 | 0
|  |  | 22.4 | 52.099 | 1 | -0.146 | 49.185 | 20.084 | 0 | 0
|  |  | 22.45 | 52.124 | 1 | -0.146 | 49.175 | 20.084 | 0 | 0
|  |  | 22.5 | 52.15 | 1 | -0.146 | 49.165 | 20.084 | 0 | 0
|  |  | 22.55 | 52.175 | 1 | -0.146 | 49.155 | 20.084 | 0 | 0
|  |  | 22.6 | 52.201 | 1 | -0.146 | 49.145 | 20.084 | 0 | 0
|  |  | 22.65 | 52.218 | 1 | -0.146 | 49.135 | 20.084 | 0 | 0
|  |  | 22.7 | 52.235 | 1 | -0.146 | 49.125 | 20.084 | 0 | 0
|  |  | 22.75 | 52.252 | 1 | -0.146 | 49.115 | 20.084 | 0 | 0
|  |  | 22.8 | 52.269 | 1 | -0.146 | 49.105 | 20.084 | 0 | 0
|  |  | 22.85 | 52.287 | 1 | -0.146 | 49.095 | 20.084 | 0 | 0
|  |  | 22.9 | 52.304 | 1 | -0.146 | 49.085 | 20.084 | 0 | 0
|  |  | 22.95 | 52.321 | 1 | -0.146 | 49.075 | 20.084 | 0 | 0
|  |  | 23 | 52.338 | 1 | -0.146 | 49.065 | 20.084 | 0 | 0
|  |  | 23.05 | 52.355 | 1 | -0.146 | 49.055 | 20.084 | 0 | 0
|  |  | 23.1 | 52.372 | 1 | -0.146 | 49.045 | 20.084 | 0 | 0
|  |  | 23.15 | 52.389 | 1 | -0.146 | 49.035 | 20.084 | 0 | 0
|  |  | 23.2 | 52.406 | 1 | -0.146 | 49.025 | 20.084 | 0 | 0
|  |  | 23.25 | 52.423 | 1 | -0.146 | 49.015 | 20.084 | 0 | 0
|  |  | 23.3 | 52.44 | 1 | -0.146 | 49.005 | 20.084 | 0 | 0
|  |  | 23.35 | 52.457 | 1 | -0.146 | 48.995 | 20.084 | 0 | 0
|  |  | 23.4 | 52.474 | 1 | -0.146 | 48.985 | 20.084 | 0 | 0
|  |  | 23.45 | 52.48 | 1 | -0.146 | 48.976 | 20.083 | 0 | 0
|  |  | 23.5 | 52.487 | 1 | -0.146 | 48.966 | 20.083 | 0 | 0
|  |  | 23.55 | 52.493 | 1 | -0.146 | 48.956 | 20.083 | 0 | 0
|  |  | 23.6 | 52.5 | 1 | -0.146 | 48.946 | 20.083 | 0 | 0
|  |  | 23.65 | 52.506 | 1 | -0.146 | 48.936 | 20.083 | 0 | 0
|  |  | 23.7 | 52.513 | 1 | -0.146 | 48.926 | 20.083 | 0 | 0
|  |  | 23.75 | 52.52 | 1 | -0.146 | 48.916 | 20.083 | 0 | 0
|  |  | 23.8 | 52.526 | 1 | -0.146 | 48.906 | 20.083 | 0 | 0
|  |  | 23.85 | 52.532 | 1 | -0.146 | 48.896 | 20.083 | 0 | 0
|  |  | 23.9 | 52.539 | 1 | -0.146 | 48.886 | 20.083 | 0 | 0
|  |  | 23.95 | 52.545 | 1 | -0.146 | 48.876 | 20.083 | 0 | 0
|  |  | 24 | 52.552 | 1 | -0.146 | 48.866 | 20.083 | 0 | 0
|  |  | 24.05 | 52.558 | 1 | -0.146 | 48.856 | 20.083 | 0 | 0
|  |  | 24.1 | 52.564 | 1 | -0.146 | 48.846 | 20.083 | 0 | 0
|  |  | 24.15 | 52.571 | 1 | -0.146 | 48.836 | 20.083 | 0 | 0
|  |  | 24.2 | 52.577 | 1 | -0.146 | 48.826 | 20.083 | 0 | 0
|  |  | 24.25 | 52.577 | 1 | -0.146 | 48.816 | 20.083 | 0 | 0
|  |  | 24.3 | 52.578 | 1 | -0.146 | 48.806 | 20.083 | 0 | 0
|  |  | 24.35 | 52.578 | 1 | -0.146 | 48.796 | 20.083 | 0 | 0
|  |  | 24.4 | 52.578 | 1 | -0.146 | 48.786 | 20.083 | 0 | 0
|  |  | 24.45 | 52.578 | 1 | -0.146 | 48.776 | 20.083 | 0 | 0
|  |  | 24.5 | 52.579 | 1 | -0.146 | 48.766 | 20.083 | 0 | 0
|  |  | 24.55 | 52.579 | 1 | -0.146 | 48.756 | 20.083 | 0 | 0
|  |  | 24.6 | 52.579 | 1 | -0.146 | 48.746 | 20.083 | 0 | 0
|  |  | 24.65 | 52.579 | 1 | -0.146 | 48.736 | 20.083 | 0 | 0
|  |  | 24.7 | 52.579 | 1 | -0.146 | 48.726 | 20.083 | 0 | 0
|  |  | 24.75 | 52.579 | 1 | -0.146 | 48.716 | 20.083 | 0 | 0
|  |  | 24.8 | 52.579 | 1 | -0.146 | 48.706 | 20.083 | 0 | 0
|  |  | 24.85 | 52.579 | 1 | -0.146 | 48.696 | 20.083 | 0 | 0
|  |  | 24.9 | 52.579 | 1 | -0.146 | 48.686 | 20.083 | 0 | 0
|  |  | 24.95 | 52.579 | 1 | -0.146 | 48.676 | 20.083 | 0 | 0
|  |  | 25 | 52.579 | 1 | -0.146 | 48.666 | 20.083 | 0 | 0
|  |  | 25.05 | 52.579 | 1 | -0.146 | 48.656 | 20.083 | 0 | 0
|  |  | 25.1 | 52.579 | 1 | -0.146 | 48.646 | 20.083 | 0 | 0
|  |  | 25.15 | 52.579 | 1 | -0.146 | 48.636 | 20.083 | 0 | 0
|  |  | 25.2 | 52.579 | 1 | -0.146 | 48.626 | 20.083 | 0 | 0
|  |  | 25.25 | 52.579 | 1 | -0.146 | 48.616 | 20.083 | 0 | 0
|  |  | 25.3 | 52.579 | 1 | -0.146 | 48.606 | 20.083 | 0 | 0
|  |  | 25.35 | 52.579 | 1 | -0.146 | 48.596 | 20.083 | 0 | 0
|  |  | 25.4 | 52.579 | 1 | -0.146 | 48.586 | 20.083 | 0 | 0
|  |  | 25.45 | 52.579 | 1 | -0.146 | 48.576 | 20.083 | 0 | 0
|  |  | 25.5 | 52.579 | 1 | -0.146 | 48.566 | 20.083 | 0 | 0
|  |  | 25.55 | 52.579 | 1 | -0.146 | 48.556 | 20.083 | 0 | 0
|  |  | 25.6 | 52.578 | 1 | -0.146 | 48.546 | 20.083 | 0 | 0
|  |  | 25.65 | 52.578 | 1 | -0.146 | 48.536 | 20.083 | 0 | 0
|  |  | 25.7 | 52.578 | 1 | -0.146 | 48.526 | 20.083 | 0 | 0
|  |  | 25.75 | 52.578 | 1 | -0.146 | 48.516 | 20.083 | 0 | 0
|  |  | 25.8 | 52.577 | 1 | -0.146 | 48.506 | 20.083 | 0 | 0
|  |  | 25.85 | 52.571 | 1 | -0.146 | 48.496 | 20.083 | 0 | 0
|  |  | 25.9 | 52.565 | 1 | -0.146 | 48.486 | 20.083 | 0 | 0
|  |  | 25.95 | 52.559 | 1 | -0.146 | 48.476 | 20.083 | 0 | 0
|  |  | 26 | 52.552 | 1 | -0.146 | 48.466 | 20.083 | 0 | 0
|  |  | 26.05 | 52.546 | 1 | -0.146 | 48.456 | 20.083 | 0 | 0
|  |  | 26.1 | 52.54 | 1 | -0.146 | 48.446 | 20.083 | 0 | 0
|  |  | 26.15 | 52.533 | 1 | -0.146 | 48.436 | 20.083 | 0 | 0
|  |  | 26.2 | 52.527 | 1 | -0.146 | 48.426 | 20.083 | 0 | 0
|  |  | 26.25 | 52.52 | 1 | -0.146 | 48.416 | 20.083 | 0 | 0
|  |  | 26.3 | 52.514 | 1 | -0.146 | 48.406 | 20.083 | 0 | 0
|  |  | 26.35 | 52.507 | 1 | -0.146 | 48.396 | 20.083 | 0 | 0
|  |  | 26.4 | 52.501 | 1 | -0.146 | 48.386 | 20.083 | 0 | 0
|  |  | 26.45 | 52.494 | 1 | -0.146 | 48.376 | 20.083 | 0 | 0
|  |  | 26.5 | 52.488 | 1 | -0.146 | 48.366 | 20.083 | 0 | 0
|  |  | 26.55 | 52.481 | 1 | -0.146 | 48.356 | 20.083 | 0 | 0
|  |  | 26.6 | 52.475 | 1 | -0.146 | 48.346 | 20.083 | 0 | 0
|  |  | 26.65 | 52.458 | 1 | -0.146 | 48.336 | 20.083 | 0 | 0
|  |  | 26.7 | 52.441 | 1 | -0.146 | 48.326 | 20.083 | 0 | 0
|  |  | 26.75 | 52.424 | 1 | -0.146 | 48.316 | 20.083 | 0 | 0
|  |  | 26.8 | 52.407 | 1 | -0.146 | 48.306 | 20.083 | 0 | 0
|  |  | 26.85 | 52.39 | 1 | -0.146 | 48.296 | 20.083 | 0 | 0
|  |  | 26.9 | 52.372 | 1 | -0.146 | 48.286 | 20.082 | 0 | 0
|  |  | 26.95 | 52.355 | 1 | -0.146 | 48.276 | 20.082 | 0 | 0
|  |  | 27 | 52.338 | 1 | -0.146 | 48.266 | 20.082 | 0 | 0
|  |  | 27.05 | 52.321 | 1 | -0.146 | 48.256 | 20.082 | 0 | 0
|  |  | 27.1 | 52.304 | 1 | -0.146 | 48.246 | 20.082 | 0 | 0
|  |  | 27.15 | 52.287 | 1 | -0.146 | 48.236 | 20.082 | 0 | 0
|  |  | 27.2 | 52.269 | 1 | -0.146 | 48.226 | 20.082 | 0 | 0
|  |  | 27.25 | 52.252 | 1 | -0.146 | 48.216 | 20.082 | 0 | 0
|  |  | 27.3 | 52.235 | 1 | -0.146 | 48.206 | 20.082 | 0 | 0
|  |  | 27.35 | 52.218 | 1 | -0.146 | 48.196 | 20.082 | 0 | 0
|  |  | 27.4 | 52.2 | 1 | -0.146 | 48.186 | 20.082 | 0 | 0
|  |  | 27.45 | 52.175 | 1 | -0.146 | 48.176 | 20.082 | 0 | 0
|  |  | 27.5 | 52.15 | 1 | -0.146 | 48.166 | 20.082 | 0 | 0
|  |  | 27.55 | 52.125 | 1 | -0.146 | 48.156 | 20.082 | 0 | 0
|  |  | 27.6 | 52.1 | 1 | -0.146 | 48.146 | 20.082 | 0 | 0
|  |  | 27.65 | 52.074 | 1 | -0.146 | 48.136 | 20.082 | 0 | 0
|  |  | 27.7 | 52.049 | 1 | -0.146 | 48.126 | 20.082 | 0 | 0
|  |  | 27.75 | 52.024 | 1 | -0.146 | 48.116 | 20.082 | 0 | 0
|  |  | 27.8 | 51.998 | 1 | -0.146 | 48.106 | 20.082 | 0 | 0
|  |  | 27.85 | 51.973 | 1 | -0.146 | 48.096 | 20.082 | 0 | 0
|  |  | 27.9 | 51.948 | 1 | -0.146 | 48.086 | 20.082 | 0 | 0
|  |  | 27.95 | 51.922 | 1 | -0.146 | 48.076 | 20.082 | 0 | 0
|  |  | 28 | 51.897 | 1 | -0.146 | 48.066 | 20.082 | 0 | 0
|  |  | 28.05 | 51.871 | 1 | -0.146 | 48.056 | 20.082 | 0 | 0
|  |  | 28.1 | 51.846 | 1 | -0.146 | 48.046 | 20.082 | 0 | 0
|  |  | 28.15 | 51.821 | 1 | -0.146 | 48.036 | 20.082 | 0 | 0
|  |  | 28.2 | 51.795 | 1 | -0.146 | 48.026 | 20.082 | 0 | 0
|  |  | 28.25 | 51.744 | 1 | -0.146 | 48.016 | 20.082 | 0 | 0
|  |  | 28.3 | 51.692 | 1 | -0.146 | 48.006 | 20.082 | 0 | 0
|  |  | 28.35 | 51.641 | 1 | -0.146 | 47.996 | 20.082 | 0 | 0
|  |  | 28.4 | 51.589 | 1 | -0.146 | 47.986 | 20.082 | 0 | 0
|  |  | 28.45 | 51.538 | 1 | -0.146 | 47.976 | 20.082 | 0 | 0
|  |  | 28.5 | 51.486 | 1 | -0.146 | 47.966 | 20.082 | 0 | 0
|  |  | 28.55 | 51.434 | 1 | -0.146 | 47.956 | 20.082 | 0 | 0
|  |  | 28.6 | 51.383 | 1 | -0.146 | 47.946 | 20.082 | 0 | 0
|  |  | 28.65 | 51.328 | 1 | -0.146 | 47.936 | 20.082 | 0 | 0
|  |  | 28.7 | 51.273 | 1 | -0.146 | 47.926 | 20.082 | 0 | 0
|  |  | 28.75 | 51.218 | 1 | -0.146 | 47.916 | 20.082 | 0 | 0
|  |  | 28.8 | 51.162 | 1 | -0.146 | 47.906 | 20.082 | 0 | 0
|  |  | 28.85 | 51.107 | 1 | -0.146 | 47.896 | 20.082 | 0 | 0
|  |  | 28.9 | 51.052 | 1 | -0.146 | 47.886 | 20.082 | 0 | 0
|  |  | 28.95 | 50.997 | 1 | -0.146 | 47.876 | 20.082 | 0 | 0
|  |  | 29 | 50.942 | 1 | -0.146 | 47.866 | 20.082 | 0 | 0
|  |  | 29.05 | 50.851 | 1 | -0.146 | 47.856 | 20.082 | 0 | 0
|  |  | 29.1 | 50.761 | 1 | -0.146 | 47.846 | 20.082 | 0 | 0
|  |  | 29.15 | 50.67 | 1 | -0.146 | 47.836 | 20.082 | 0 | 0
|  |  | 29.2 | 50.579 | 1 | -0.146 | 47.826 | 20.082 | 0 | 0
|  |  | 29.25 | 50.485 | 1 | -0.146 | 47.816 | 20.082 | 0 | 0
|  |  | 29.3 | 50.391 | 1 | -0.146 | 47.806 | 20.082 | 0 | 0
|  |  | 29.35 | 50.297 | 1 | -0.146 | 47.796 | 20.082 | 0 | 0
|  |  | 29.4 | 50.203 | 1 | -0.146 | 47.786 | 20.082 | 0 | 0
|  |  | 29.45 | 50.072 | 1 | -0.146 | 47.776 | 20.082 | 0 | 0
|  |  | 29.5 | 49.941 | 1 | -0.146 | 47.766 | 20.082 | 0 | 0
|  |  | 29.55 | 49.809 | 1 | -0.146 | 47.756 | 20.082 | 0 | 0
|  |  | 29.6 | 49.678 | 1 | -0.146 | 47.746 | 20.082 | 0 | 0
|  |  | 29.65 | 49.508 | 1 | -0.146 | 47.736 | 20.082 | 0 | 0
|  |  | 29.7 | 49.338 | 1 | -0.146 | 47.726 | 20.082 | 0 | 0
|  |  | 29.75 | 49.167 | 1 | -0.146 | 47.716 | 20.082 | 0 | 0
|  |  | 29.8 | 48.997 | 1 | -0.146 | 47.706 | 20.082 | 0 | 0
|  |  | 29.85 | 48.68 | 1 | -0.146 | 47.696 | 20.082 | 0 | 0
|  |  | 29.9 | 48.362 | 1 | -0.146 | 47.686 | 20.082 | 0 | 0
|  |  | 29.95 | 48.045 | 1 | -0.146 | 47.676 | 20.082 | 0 | 0
|  |  | 30 | 47.728 | 1 | -0.146 | 47.666 | 20.082 | 0 | 0
|  |  | 30.05 | 46.965 | 1 | -0.146 | 47.656 | 20.082 | 0 | 0
|  |  | 30.1 | 46.202 | 1 | -0.146 | 47.646 | 20.082 | 0 | 0
|  |  | 30.15 | 45.44 | 1 | -0.146 | 47.636 | 20.082 | 0 | 0
|  |  | 30.2 | 44.677 | 1 | -0.146 | 47.626 | 20.082 | 0 | 0
|  |  | 30.25 | 42.089 | 1 | -0.146 | 47.616 | 20.082 | 0 | 0
|  |  | 30.3 | 39.502 | 1 | -0.146 | 47.606 | 20.082 | 0 | 0
|  |  | 30.35 | 36.914 | 1 | -0.146 | 47.596 | 20.082 | 0 | 0
|  |  | 30.4 | 34.326 | 1 | -0.146 | 47.586 | 20.081 | 0 | 0
|  |  | 30.45 | 29.59 | 1 | -0.146 | 47.576 | 20.081 | 0 | 0
|  |  | 30.5 | 24.855 | 1 | -0.146 | 47.566 | 20.081 | 0 | 0
|  |  | 30.55 | 20.119 | 1 | -0.146 | 47.556 | 20.081 | 0 | 0
|  |  | 30.6 | 15.383 | 1 | -0.146 | 47.546 | 20.081 | 0 | 0
|  |  | 30.65 | 13.89 | 1 | -0.146 | 47.536 | 20.081 | 0 | 0
|  |  | 30.7 | 12.397 | 1 | -0.146 | 47.526 | 20.081 | 0 | 0
|  |  | 30.75 | 10.904 | 1 | -0.146 | 47.516 | 20.081 | 0 | 0
|  |  | 30.8 | 9.4105 | 1 | -0.146 | 47.506 | 20.081 | 0 | 0
|  |  | 30.85 | 8.9603 | 1 | -0.146 | 47.496 | 20.081 | 0 | 0
|  |  | 30.9 | 8.5101 | 1 | -0.146 | 47.486 | 20.081 | 0 | 0
|  |  | 30.95 | 8.0598 | 1 | -0.146 | 47.476 | 20.081 | 0 | 0
|  |  | 31 | 7.6096 | 1 | -0.146 | 47.466 | 20.081 | 0 | 0
|  |  | 31.05 | 7.4182 | 1 | -0.146 | 47.456 | 20.081 | 0 | 0
|  |  | 31.1 | 7.2268 | 1 | -0.146 | 47.446 | 20.081 | 0 | 0
|  |  | 31.15 | 7.0355 | 1 | -0.146 | 47.436 | 20.081 | 0 | 0
|  |  | 31.2 | 6.8441 | 1 | -0.146 | 47.426 | 20.081 | 0 | 0
|  |  | 31.25 | 6.6528 | 1 | -0.146 | 47.416 | 20.081 | 0 | 0
|  |  | 31.3 | 6.4614 | 1 | -0.146 | 47.406 | 20.081 | 0 | 0
|  |  | 31.35 | 6.2701 | 1 | -0.146 | 47.396 | 20.081 | 0 | 0
|  |  | 31.4 | 6.0788 | 1 | -0.146 | 47.386 | 20.081 | 0 | 0
|  |  | 31.45 | 5.9928 | 1 | -0.146 | 47.376 | 20.081 | 0 | 0
|  |  | 31.5 | 5.9068 | 1 | -0.146 | 47.366 | 20.081 | 0 | 0
|  |  | 31.55 | 5.8209 | 1 | -0.146 | 47.356 | 20.081 | 0 | 0
|  |  | 31.6 | 5.7349 | 1 | -0.146 | 47.346 | 20.081 | 0 | 0
|  |  | 31.65 | 5.6489 | 1 | -0.146 | 47.336 | 20.081 | 0 | 0
|  |  | 31.7 | 5.563 | 1 | -0.146 | 47.326 | 20.081 | 0 | 0
|  |  | 31.75 | 5.477 | 1 | -0.146 | 47.316 | 20.081 | 0 | 0
|  |  | 31.8 | 5.391 | 1 | -0.146 | 47.306 | 20.081 | 0 | 0
|  |  | 31.85 | 5.3051 | 1 | -0.146 | 47.296 | 20.081 | 0 | 0
|  |  | 31.9 | 5.2192 | 1 | -0.146 | 47.286 | 20.081 | 0 | 0
|  |  | 31.95 | 5.1332 | 1 | -0.146 | 47.276 | 20.081 | 0 | 0
|  |  | 32 | 5.0473 | 1 | -0.146 | 47.266 | 20.081 | 0 | 0
|  |  | 32.05 | 4.9614 | 1 | -0.146 | 47.256 | 20.081 | 0 | 0
|  |  | 32.1 | 4.8755 | 1 | -0.146 | 47.246 | 20.081 | 0 | 0
|  |  | 32.15 | 4.7896 | 1 | -0.146 | 47.236 | 20.081 | 0 | 0
|  |  | 32.2 | 4.7037 | 1 | -0.146 | 47.226 | 20.081 | 0 | 0
|  |  | 32.25 | 4.6451 | 1 | -0.146 | 47.216 | 20.081 | 0 | 0
|  |  | 32.3 | 4.5866 | 1 | -0.146 | 47.206 | 20.081 | 0 | 0
|  |  | 32.35 | 4.528 | 1 | -0.146 | 47.196 | 20.081 | 0 | 0
|  |  | 32.4 | 4.4694 | 1 | -0.146 | 47.186 | 20.081 | 0 | 0
|  |  | 32.45 | 4.4108 | 1 | -0.146 | 47.176 | 20.081 | 0 | 0
|  |  | 32.5 | 4.3522 | 1 | -0.146 | 47.166 | 20.081 | 0 | 0
|  |  | 32.55 | 4.2937 | 1 | -0.146 | 47.156 | 20.081 | 0 | 0
|  |  | 32.6 | 4.2351 | 1 | -0.146 | 47.146 | 20.081 | 0 | 0
|  |  | 32.65 | 4.1765 | 1 | -0.146 | 47.136 | 20.081 | 0 | 0
|  |  | 32.7 | 4.118 | 1 | -0.146 | 47.126 | 20.081 | 0 | 0
|  |  | 32.75 | 4.0594 | 1 | -0.146 | 47.116 | 20.081 | 0 | 0
|  |  | 32.8 | 4.0009 | 1 | -0.146 | 47.106 | 20.081 | 0 | 0
|  |  | 32.85 | 3.9424 | 1 | -0.146 | 47.096 | 20.081 | 0 | 0
|  |  | 32.9 | 3.8839 | 1 | -0.146 | 47.086 | 20.081 | 0 | 0
|  |  | 32.95 | 3.8253 | 1 | -0.146 | 47.076 | 20.081 | 0 | 0
|  |  | 33 | 3.7668 | 1 | -0.146 | 47.066 | 20.081 | 0 | 0
|  |  | 33.05 | 3.7173 | 1 | -0.146 | 47.056 | 20.081 | 0 | 0
|  |  | 33.1 | 3.6678 | 1 | -0.146 | 47.046 | 20.081 | 0 | 0
|  |  | 33.15 | 3.6183 | 1 | -0.146 | 47.036 | 20.081 | 0 | 0
|  |  | 33.2 | 3.5688 | 1 | -0.146 | 47.026 | 20.081 | 0 | 0
|  |  | 33.25 | 3.5194 | 1 | -0.146 | 47.016 | 20.081 | 0 | 0
|  |  | 33.3 | 3.4699 | 1 | -0.146 | 47.006 | 20.081 | 0 | 0
|  |  | 33.35 | 3.4204 | 1 | -0.146 | 46.996 | 20.081 | 0 | 0
|  |  | 33.4 | 3.3709 | 1 | -0.146 | 46.986 | 20.081 | 0 | 0
|  |  | 33.45 | 3.3215 | 1 | -0.146 | 46.976 | 20.081 | 0 | 0
|  |  | 33.5 | 3.272 | 1 | -0.146 | 46.966 | 20.081 | 0 | 0
|  |  | 33.55 | 3.2226 | 1 | -0.146 | 46.956 | 20.081 | 0 | 0
|  |  | 33.6 | 3.1731 | 1 | -0.146 | 46.946 | 20.081 | 0 | 0
|  |  | 33.65 | 3.1237 | 1 | -0.146 | 46.936 | 20.081 | 0 | 0
|  |  | 33.7 | 3.0743 | 1 | -0.146 | 46.926 | 20.081 | 0 | 0
|  |  | 33.75 | 3.0248 | 1 | -0.146 | 46.916 | 20.081 | 0 | 0
|  |  | 33.8 | 2.9754 | 1 | -0.146 | 46.906 | 20.081 | 0 | 0
|  |  | 33.85 | 2.9377 | 1 | -0.146 | 46.896 | 20.08 | 0 | 0
|  |  | 33.9 | 2.8999 | 1 | -0.146 | 46.886 | 20.08 | 0 | 0
|  |  | 33.95 | 2.8622 | 1 | -0.146 | 46.876 | 20.08 | 0 | 0
|  |  | 34 | 2.8244 | 1 | -0.146 | 46.866 | 20.08 | 0 | 0
|  |  | 34.05 | 2.7867 | 1 | -0.146 | 46.856 | 20.08 | 0 | 0
|  |  | 34.1 | 2.749 | 1 | -0.146 | 46.846 | 20.08 | 0 | 0
|  |  | 34.15 | 2.7113 | 1 | -0.146 | 46.836 | 20.08 | 0 | 0
|  |  | 34.2 | 2.6736 | 1 | -0.146 | 46.826 | 20.08 | 0 | 0
|  |  | 34.25 | 2.6359 | 1 | -0.146 | 46.816 | 20.08 | 0 | 0
|  |  | 34.3 | 2.5982 | 1 | -0.146 | 46.806 | 20.08 | 0 | 0
|  |  | 34.35 | 2.5605 | 1 | -0.146 | 46.796 | 20.08 | 0 | 0
|  |  | 34.4 | 2.5228 | 1 | -0.146 | 46.786 | 20.08 | 0 | 0
|  |  | 34.45 | 2.4851 | 1 | -0.146 | 46.776 | 20.08 | 0 | 0
|  |  | 34.5 | 2.4474 | 1 | -0.146 | 46.766 | 20.08 | 0 | 0
|  |  | 34.55 | 2.4097 | 1 | -0.146 | 46.756 | 20.08 | 0 | 0
|  |  | 34.6 | 2.3721 | 1 | -0.146 | 46.746 | 20.08 | 0 | 0
|  |  | 34.65 | 2.3386 | 1 | -0.146 | 46.736 | 20.08 | 0 | 0
|  |  | 34.7 | 2.3051 | 1 | -0.146 | 46.726 | 20.08 | 0 | 0
|  |  | 34.75 | 2.2716 | 1 | -0.146 | 46.716 | 20.08 | 0 | 0
|  |  | 34.8 | 2.2381 | 1 | -0.146 | 46.706 | 20.08 | 0 | 0
|  |  | 34.85 | 2.2046 | 1 | -0.146 | 46.696 | 20.08 | 0 | 0
|  |  | 34.9 | 2.1711 | 1 | -0.146 | 46.686 | 20.08 | 0 | 0
|  |  | 34.95 | 2.1376 | 1 | -0.146 | 46.676 | 20.08 | 0 | 0
|  |  | 35 | 2.1041 | 1 | -0.146 | 46.666 | 20.08 | 0 | 0
|  |  | 35.05 | 2.0706 | 1 | -0.146 | 46.656 | 20.08 | 0 | 0
|  |  | 35.1 | 2.0372 | 1 | -0.146 | 46.646 | 20.08 | 0 | 0
|  |  | 35.15 | 2.0037 | 1 | -0.146 | 46.636 | 20.08 | 0 | 0
|  |  | 35.2 | 1.9702 | 1 | -0.146 | 46.626 | 20.08 | 0 | 0
|  |  | 35.25 | 1.9368 | 1 | -0.146 | 46.616 | 20.08 | 0 | 0
|  |  | 35.3 | 1.9033 | 1 | -0.146 | 46.606 | 20.08 | 0 | 0
|  |  | 35.35 | 1.8699 | 1 | -0.146 | 46.596 | 20.08 | 0 | 0
|  |  | 35.4 | 1.8364 | 1 | -0.146 | 46.586 | 20.08 | 0 | 0
|  |  | 35.45 | 1.8055 | 1 | -0.146 | 46.576 | 20.08 | 0 | 0
|  |  | 35.5 | 1.7746 | 1 | -0.146 | 46.566 | 20.08 | 0 | 0
|  |  | 35.55 | 1.7437 | 1 | -0.146 | 46.556 | 20.08 | 0 | 0
|  |  | 35.6 | 1.7128 | 1 | -0.146 | 46.546 | 20.08 | 0 | 0
|  |  | 35.65 | 1.682 | 1 | -0.146 | 46.536 | 20.08 | 0 | 0
|  |  | 35.7 | 1.6511 | 1 | -0.146 | 46.526 | 20.08 | 0 | 0
|  |  | 35.75 | 1.6202 | 1 | -0.146 | 46.516 | 20.08 | 0 | 0
|  |  | 35.8 | 1.5893 | 1 | -0.146 | 46.506 | 20.08 | 0 | 0
|  |  | 35.85 | 1.5585 | 1 | -0.146 | 46.496 | 20.08 | 0 | 0
|  |  | 35.9 | 1.5276 | 1 | -0.146 | 46.486 | 20.08 | 0 | 0
|  |  | 35.95 | 1.4967 | 1 | -0.146 | 46.476 | 20.08 | 0 | 0
|  |  | 36 | 1.4659 | 1 | -0.146 | 46.466 | 20.08 | 0 | 0
|  |  | 36.05 | 1.435 | 1 | -0.146 | 46.456 | 20.08 | 0 | 0
|  |  | 36.1 | 1.4042 | 1 | -0.146 | 46.446 | 20.08 | 0 | 0
|  |  | 36.15 | 1.3734 | 1 | -0.146 | 46.436 | 20.08 | 0 | 0
|  |  | 36.2 | 1.3425 | 1 | -0.146 | 46.426 | 20.08 | 0 | 0
|  |  | 36.25 | 1.3293 | 1 | -0.146 | 46.416 | 20.08 | 0 | 0
|  |  | 36.3 | 1.316 | 1 | -0.146 | 46.406 | 20.08 | 0 | 0
|  |  | 36.35 | 1.3027 | 1 | -0.146 | 46.396 | 20.08 | 0 | 0
|  |  | 36.4 | 1.2895 | 1 | -0.146 | 46.386 | 20.08 | 0 | 0
|  |  | 36.45 | 1.2762 | 1 | -0.146 | 46.376 | 20.08 | 0 | 0
|  |  | 36.5 | 1.263 | 1 | -0.146 | 46.366 | 20.08 | 0 | 0
|  |  | 36.55 | 1.2497 | 1 | -0.146 | 46.356 | 20.08 | 0 | 0
|  |  | 36.6 | 1.2364 | 1 | -0.146 | 46.346 | 20.08 | 0 | 0
|  |  | 36.65 | 1.2232 | 1 | -0.146 | 46.336 | 20.08 | 0 | 0
|  |  | 36.7 | 1.2099 | 1 | -0.146 | 46.326 | 20.08 | 0 | 0
|  |  | 36.75 | 1.1967 | 1 | -0.146 | 46.316 | 20.08 | 0 | 0
|  |  | 36.8 | 1.1835 | 1 | -0.146 | 46.306 | 20.08 | 0 | 0
|  |  | 36.85 | 1.1702 | 1 | -0.146 | 46.296 | 20.08 | 0 | 0
|  |  | 36.9 | 1.157 | 1 | -0.146 | 46.286 | 20.08 | 0 | 0
|  |  | 36.95 | 1.1437 | 1 | -0.146 | 46.276 | 20.08 | 0 | 0
|  |  | 37 | 1.1305 | 1 | -0.146 | 46.266 | 20.08 | 0 | 0
|  |  | 37.05 | 1.1203 | 1 | -0.146 | 46.256 | 20.08 | 0 | 0
|  |  | 37.1 | 1.11 | 1 | -0.146 | 46.246 | 20.08 | 0 | 0
|  |  | 37.15 | 1.0998 | 1 | -0.146 | 46.236 | 20.08 | 0 | 0
|  |  | 37.2 | 1.0895 | 1 | -0.146 | 46.226 | 20.08 | 0 | 0
|  |  | 37.25 | 1.0793 | 1 | -0.146 | 46.216 | 20.08 | 0 | 0
|  |  | 37.3 | 1.0691 | 1 | -0.146 | 46.206 | 20.079 | 0 | 0
|  |  | 37.35 | 1.0589 | 1 | -0.146 | 46.196 | 20.079 | 0 | 0
|  |  | 37.4 | 1.0486 | 1 | -0.146 | 46.186 | 20.079 | 0 | 0
|  |  | 37.45 | 1.0384 | 1 | -0.146 | 46.176 | 20.079 | 0 | 0
|  |  | 37.5 | 1.0282 | 1 | -0.146 | 46.166 | 20.079 | 0 | 0
|  |  | 37.55 | 1.018 | 1 | -0.146 | 46.156 | 20.079 | 0 | 0
|  |  | 37.6 | 1.0077 | 1 | -0.146 | 46.146 | 20.079 | 0 | 0
|  |  | 37.65 | 0.99753 | 1 | -0.146 | 46.136 | 20.079 | 0 | 0
|  |  | 37.7 | 0.98732 | 1 | -0.146 | 46.126 | 20.079 | 0 | 0
|  |  | 37.75 | 0.9771 | 1 | -0.146 | 46.116 | 20.079 | 0 | 0
|  |  | 37.8 | 0.96689 | 1 | -0.146 | 46.106 | 20.079 | 0 | 0
|  |  | 37.85 | 0.95755 | 1 | -0.146 | 46.096 | 20.079 | 0 | 0
|  |  | 37.9 | 0.94822 | 1 | -0.146 | 46.086 | 20.079 | 0 | 0
|  |  | 37.95 | 0.93889 | 1 | -0.146 | 46.076 | 20.079 | 0 | 0
|  |  | 38 | 0.92955 | 1 | -0.146 | 46.066 | 20.079 | 0 | 0
|  |  | 38.05 | 0.92023 | 1 | -0.146 | 46.056 | 20.079 | 0 | 0
|  |  | 38.1 | 0.9109 | 1 | -0.146 | 46.046 | 20.079 | 0 | 0
|  |  | 38.15 | 0.90157 | 1 | -0.146 | 46.036 | 20.079 | 0 | 0
|  |  | 38.2 | 0.89225 | 1 | -0.146 | 46.026 | 20.079 | 0 | 0
|  |  | 38.25 | 0.88293 | 1 | -0.146 | 46.016 | 20.079 | 0 | 0
|  |  | 38.3 | 0.87361 | 1 | -0.146 | 46.006 | 20.079 | 0 | 0
|  |  | 38.35 | 0.86429 | 1 | -0.146 | 45.996 | 20.079 | 0 | 0
|  |  | 38.4 | 0.85497 | 1 | -0.146 | 45.986 | 20.079 | 0 | 0
|  |  | 38.45 | 0.84565 | 1 | -0.146 | 45.977 | 20.079 | 0 | 0
|  |  | 38.5 | 0.83634 | 1 | -0.146 | 45.967 | 20.079 | 0 | 0
|  |  | 38.55 | 0.82702 | 1 | -0.146 | 45.957 | 20.079 | 0 | 0
|  |  | 38.6 | 0.81771 | 1 | -0.146 | 45.947 | 20.079 | 0 | 0
|  |  | 38.65 | 0.8098 | 1 | -0.146 | 45.937 | 20.079 | 0 | 0
|  |  | 38.7 | 0.80188 | 1 | -0.146 | 45.927 | 20.079 | 0 | 0
|  |  | 38.75 | 0.79397 | 1 | -0.146 | 45.917 | 20.079 | 0 | 0
|  |  | 38.8 | 0.78606 | 1 | -0.146 | 45.907 | 20.079 | 0 | 0
|  |  | 38.85 | 0.77815 | 1 | -0.146 | 45.897 | 20.079 | 0 | 0
|  |  | 38.9 | 0.77025 | 1 | -0.146 | 45.887 | 20.079 | 0 | 0
|  |  | 38.95 | 0.76234 | 1 | -0.146 | 45.877 | 20.079 | 0 | 0
|  |  | 39 | 0.75443 | 1 | -0.146 | 45.867 | 20.079 | 0 | 0
|  |  | 39.05 | 0.74653 | 1 | -0.146 | 45.857 | 20.079 | 0 | 0
|  |  | 39.1 | 0.73863 | 1 | -0.146 | 45.847 | 20.079 | 0 | 0
|  |  | 39.15 | 0.73073 | 1 | -0.146 | 45.837 | 20.079 | 0 | 0
|  |  | 39.2 | 0.72283 | 1 | -0.146 | 45.827 | 20.079 | 0 | 0
|  |  | 39.25 | 0.71493 | 1 | -0.146 | 45.817 | 20.079 | 0 | 0
|  |  | 39.3 | 0.70704 | 1 | -0.146 | 45.807 | 20.079 | 0 | 0
|  |  | 39.35 | 0.69914 | 1 | -0.146 | 45.797 | 20.079 | 0 | 0
|  |  | 39.4 | 0.69125 | 1 | -0.146 | 45.787 | 20.079 | 0 | 0
|  |  | 39.45 | 0.68501 | 1 | -0.146 | 45.777 | 20.079 | 0 | 0
|  |  | 39.5 | 0.67878 | 1 | -0.146 | 45.767 | 20.079 | 0 | 0
|  |  | 39.55 | 0.67255 | 1 | -0.146 | 45.757 | 20.079 | 0 | 0
|  |  | 39.6 | 0.66631 | 1 | -0.146 | 45.747 | 20.079 | 0 | 0
|  |  | 39.65 | 0.66008 | 1 | -0.146 | 45.737 | 20.079 | 0 | 0
|  |  | 39.7 | 0.65386 | 1 | -0.146 | 45.727 | 20.079 | 0 | 0
|  |  | 39.75 | 0.64763 | 1 | -0.146 | 45.717 | 20.079 | 0 | 0
|  |  | 39.8 | 0.6414 | 1 | -0.146 | 45.707 | 20.079 | 0 | 0
|  |  | 39.85 | 0.63518 | 1 | -0.146 | 45.697 | 20.079 | 0 | 0
|  |  | 39.9 | 0.62895 | 1 | -0.146 | 45.687 | 20.079 | 0 | 0
|  |  | 39.95 | 0.62273 | 1 | -0.146 | 45.677 | 20.079 | 0 | 0
|  |  | 40 | 0.6165 | 1 | -0.146 | 45.667 | 20.079 | 0 | 0
|  |  | 40.05 | 0.61028 | 1 | -0.146 | 45.657 | 20.079 | 0 | 0
|  |  | 40.1 | 0.60407 | 1 | -0.146 | 45.647 | 20.079 | 0 | 0
|  |  | 40.15 | 0.59785 | 1 | -0.146 | 45.637 | 20.079 | 0 | 0
|  |  | 40.2 | 0.59163 | 1 | -0.146 | 45.627 | 20.079 | 0 | 0
|  |  | 40.25 | 0.5566 | 1 | -0.146 | 45.617 | 20.079 | 0 | 0
|  |  | 40.3 | 0.52158 | 1 | -0.146 | 45.607 | 20.079 | 0 | 0
|  |  | 40.35 | 0.48656 | 1 | -0.146 | 45.597 | 20.079 | 0 | 0
|  |  | 40.4 | 0.45153 | 1 | -0.146 | 45.587 | 20.079 | 0 | 0
|  |  | 40.45 | 0.41654 | 1 | -0.146 | 45.577 | 20.079 | 0 | 0
|  |  | 40.5 | 0.38155 | 1 | -0.146 | 45.567 | 20.079 | 0 | 0
|  |  | 40.55 | 0.34656 | 1 | -0.146 | 45.557 | 20.079 | 0 | 0
|  |  | 40.6 | 0.31157 | 1 | -0.146 | 45.547 | 20.079 | 0 | 0
|  |  | 40.65 | 0.24202 | 1 | -0.146 | 45.537 | 20.079 | 0 | 0
|  |  | 40.7 | 0.17246 | 1 | -0.146 | 45.527 | 20.079 | 0 | 0
|  |  | 40.75 | 0.10291 | 1 | -0.146 | 45.517 | 20.078 | 0 | 0
|  |  | 40.8 | 0.033352 | 1 | -0.146 | 45.507 | 20.078 | 0 | 0
|  |  | 40.85 | 0.033056 | 1 | -0.146 | 45.497 | 20.078 | 0 | 0
|  |  | 40.9 | 0.032761 | 1 | -0.146 | 45.487 | 20.078 | 0 | 0
|  |  | 40.95 | 0.032465 | 1 | -0.146 | 45.477 | 20.078 | 0 | 0
|  |  | 41 | 0.032169 | 1 | -0.146 | 45.467 | 20.078 | 0 | 0
|  |  | 41.05 | 0.031902 | 1 | -0.146 | 45.457 | 20.078 | 0 | 0
|  |  | 41.1 | 0.031636 | 1 | -0.146 | 45.447 | 20.078 | 0 | 0
|  |  | 41.15 | 0.031369 | 1 | -0.146 | 45.437 | 20.078 | 0 | 0
|  |  | 41.2 | 0.031102 | 1 | -0.146 | 45.427 | 20.078 | 0 | 0
|  |  | 41.25 | 0.030945 | 1 | -0.146 | 45.417 | 20.078 | 0 | 0
|  |  | 41.3 | 0.030788 | 1 | -0.146 | 45.407 | 20.078 | 0 | 0
|  |  | 41.35 | 0.030632 | 1 | -0.146 | 45.397 | 20.078 | 0 | 0
|  |  | 41.4 | 0.030475 | 1 | -0.146 | 45.387 | 20.078 | 0 | 0
|  |  | 41.45 | 0.030113 | 1 | -0.146 | 45.377 | 20.078 | 0 | 0
|  |  | 41.5 | 0.029751 | 1 | -0.146 | 45.367 | 20.078 | 0 | 0
|  |  | 41.55 | 0.029389 | 1 | -0.146 | 45.357 | 20.078 | 0 | 0
|  |  | 41.6 | 0.029027 | 1 | -0.146 | 45.347 | 20.078 | 0 | 0
|  |  | 41.65 | 0.028879 | 1 | -0.146 | 45.337 | 20.078 | 0 | 0
|  |  | 41.7 | 0.02873 | 1 | -0.146 | 45.327 | 20.078 | 0 | 0
|  |  | 41.75 | 0.028582 | 1 | -0.146 | 45.317 | 20.078 | 0 | 0
|  |  | 41.8 | 0.028433 | 1 | -0.146 | 45.307 | 20.078 | 0 | 0
|  |  | 41.85 | 0.02818 | 1 | -0.146 | 45.297 | 20.078 | 0 | 0
|  |  | 41.9 | 0.027927 | 1 | -0.146 | 45.287 | 20.078 | 0 | 0
|  |  | 41.95 | 0.027673 | 1 | -0.146 | 45.277 | 20.078 | 0 | 0
|  |  | 42 | 0.02742 | 1 | -0.146 | 45.267 | 20.078 | 0 | 0
|  |  | 42.05 | 0.027161 | 1 | -0.146 | 45.257 | 20.078 | 0 | 0
|  |  | 42.1 | 0.026901 | 1 | -0.146 | 45.247 | 20.078 | 0 | 0
|  |  | 42.15 | 0.026642 | 1 | -0.146 | 45.237 | 20.078 | 0 | 0
|  |  | 42.2 | 0.026382 | 1 | -0.146 | 45.227 | 20.078 | 0 | 0
|  |  | 42.25 | 0.026245 | 1 | -0.146 | 45.217 | 20.078 | 0 | 0
|  |  | 42.3 | 0.026107 | 1 | -0.146 | 45.207 | 20.078 | 0 | 0
|  |  | 42.35 | 0.02597 | 1 | -0.146 | 45.197 | 20.078 | 0 | 0
|  |  | 42.4 | 0.025832 | 1 | -0.146 | 45.187 | 20.078 | 0 | 0
|  |  | 42.45 | 0.02551 | 1 | -0.146 | 45.177 | 20.078 | 0 | 0
|  |  | 42.5 | 0.025187 | 1 | -0.146 | 45.167 | 20.078 | 0 | 0
|  |  | 42.55 | 0.024865 | 1 | -0.146 | 45.157 | 20.078 | 0 | 0
|  |  | 42.6 | 0.024542 | 1 | -0.146 | 45.147 | 20.078 | 0 | 0
|  |  | 42.65 | 0.024397 | 1 | -0.146 | 45.137 | 20.078 | 0 | 0
|  |  | 42.7 | 0.024251 | 1 | -0.146 | 45.127 | 20.078 | 0 | 0
|  |  | 42.75 | 0.024106 | 1 | -0.146 | 45.117 | 20.078 | 0 | 0
|  |  | 42.8 | 0.02396 | 1 | -0.146 | 45.107 | 20.078 | 0 | 0
|  |  | 42.85 | 0.023752 | 1 | -0.146 | 45.097 | 20.078 | 0 | 0
|  |  | 42.9 | 0.023544 | 1 | -0.146 | 45.087 | 20.078 | 0 | 0
|  |  | 42.95 | 0.023336 | 1 | -0.146 | 45.077 | 20.078 | 0 | 0
|  |  | 43 | 0.023127 | 1 | -0.146 | 45.067 | 20.078 | 0 | 0
|  |  | 43.05 | 0.022878 | 1 | -0.146 | 45.057 | 20.078 | 0 | 0
|  |  | 43.1 | 0.022628 | 1 | -0.146 | 45.047 | 20.078 | 0 | 0
|  |  | 43.15 | 0.022378 | 1 | -0.146 | 45.037 | 20.078 | 0 | 0
|  |  | 43.2 | 0.022128 | 1 | -0.146 | 45.027 | 20.078 | 0 | 0
|  |  | 43.25 | 0.022011 | 1 | -0.146 | 45.017 | 20.078 | 0 | 0
|  |  | 43.3 | 0.021893 | 1 | -0.146 | 45.007 | 20.078 | 0 | 0
|  |  | 43.35 | 0.021776 | 1 | -0.146 | 44.997 | 20.078 | 0 | 0
|  |  | 43.4 | 0.021658 | 1 | -0.146 | 44.987 | 20.078 | 0 | 0
|  |  | 43.45 | 0.021379 | 1 | -0.146 | 44.977 | 20.078 | 0 | 0
|  |  | 43.5 | 0.0211 | 1 | -0.146 | 44.967 | 20.078 | 0 | 0
|  |  | 43.55 | 0.020821 | 1 | -0.146 | 44.957 | 20.078 | 0 | 0
|  |  | 43.6 | 0.020542 | 1 | -0.146 | 44.947 | 20.078 | 0 | 0
|  |  | 43.65 | 0.020401 | 1 | -0.146 | 44.937 | 20.078 | 0 | 0
|  |  | 43.7 | 0.020261 | 1 | -0.146 | 44.927 | 20.078 | 0 | 0
|  |  | 43.75 | 0.020121 | 1 | -0.146 | 44.917 | 20.078 | 0 | 0
|  |  | 43.8 | 0.019981 | 1 | -0.146 | 44.907 | 20.078 | 0 | 0
|  |  | 43.85 | 0.019813 | 1 | -0.146 | 44.897 | 20.078 | 0 | 0
|  |  | 43.9 | 0.019644 | 1 | -0.146 | 44.887 | 20.078 | 0 | 0
|  |  | 43.95 | 0.019476 | 1 | -0.146 | 44.877 | 20.078 | 0 | 0
|  |  | 44 | 0.019308 | 1 | -0.146 | 44.867 | 20.078 | 0 | 0
|  |  | 44.05 | 0.019075 | 1 | -0.146 | 44.857 | 20.078 | 0 | 0
|  |  | 44.1 | 0.018842 | 1 | -0.146 | 44.847 | 20.078 | 0 | 0
|  |  | 44.15 | 0.018608 | 1 | -0.146 | 44.837 | 20.078 | 0 | 0
|  |  | 44.2 | 0.018375 | 1 | -0.146 | 44.827 | 20.077 | 0 | 0
|  |  | 44.25 | 0.018271 | 1 | -0.146 | 44.817 | 20.077 | 0 | 0
|  |  | 44.3 | 0.018168 | 1 | -0.146 | 44.807 | 20.077 | 0 | 0
|  |  | 44.35 | 0.018064 | 1 | -0.146 | 44.797 | 20.077 | 0 | 0
|  |  | 44.4 | 0.01796 | 1 | -0.146 | 44.787 | 20.077 | 0 | 0
|  |  | 44.45 | 0.017724 | 1 | -0.146 | 44.777 | 20.077 | 0 | 0
|  |  | 44.5 | 0.017488 | 1 | -0.146 | 44.767 | 20.077 | 0 | 0
|  |  | 44.55 | 0.017252 | 1 | -0.146 | 44.757 | 20.077 | 0 | 0
|  |  | 44.6 | 0.017017 | 1 | -0.146 | 44.747 | 20.077 | 0 | 0
|  |  | 44.65 | 0.016883 | 1 | -0.146 | 44.737 | 20.077 | 0 | 0
|  |  | 44.7 | 0.01675 | 1 | -0.146 | 44.727 | 20.077 | 0 | 0
|  |  | 44.75 | 0.016616 | 1 | -0.146 | 44.717 | 20.077 | 0 | 0
|  |  | 44.8 | 0.016482 | 1 | -0.146 | 44.707 | 20.077 | 0 | 0
|  |  | 44.85 | 0.016348 | 1 | -0.146 | 44.697 | 20.077 | 0 | 0
|  |  | 44.9 | 0.016215 | 1 | -0.146 | 44.687 | 20.077 | 0 | 0
|  |  | 44.95 | 0.016081 | 1 | -0.146 | 44.677 | 20.077 | 0 | 0
|  |  | 45 | 0.015947 | 1 | -0.146 | 44.667 | 20.077 | 0 | 0
|  |  | 45.05 | 0.015733 | 1 | -0.146 | 44.657 | 20.077 | 0 | 0
|  |  | 45.1 | 0.015519 | 1 | -0.146 | 44.647 | 20.077 | 0 | 0
|  |  | 45.15 | 0.015305 | 1 | -0.146 | 44.637 | 20.077 | 0 | 0
|  |  | 45.2 | 0.015091 | 1 | -0.146 | 44.627 | 20.077 | 0 | 0
|  |  | 45.25 | 0.015002 | 1 | -0.146 | 44.617 | 20.077 | 0 | 0
|  |  | 45.3 | 0.014913 | 1 | -0.146 | 44.607 | 20.077 | 0 | 0
|  |  | 45.35 | 0.014825 | 1 | -0.146 | 44.597 | 20.077 | 0 | 0
|  |  | 45.4 | 0.014736 | 1 | -0.146 | 44.587 | 20.077 | 0 | 0
|  |  | 45.45 | 0.014544 | 1 | -0.146 | 44.577 | 20.077 | 0 | 0
|  |  | 45.5 | 0.014352 | 1 | -0.146 | 44.567 | 20.077 | 0 | 0
|  |  | 45.55 | 0.01416 | 1 | -0.146 | 44.557 | 20.077 | 0 | 0
|  |  | 45.6 | 0.013968 | 1 | -0.146 | 44.547 | 20.077 | 0 | 0
|  |  | 45.65 | 0.01384 | 1 | -0.146 | 44.537 | 20.077 | 0 | 0
|  |  | 45.7 | 0.013713 | 1 | -0.146 | 44.527 | 20.077 | 0 | 0
|  |  | 45.75 | 0.013586 | 1 | -0.146 | 44.517 | 20.077 | 0 | 0
|  |  | 45.8 | 0.013459 | 1 | -0.146 | 44.507 | 20.077 | 0 | 0
|  |  | 45.85 | 0.013354 | 1 | -0.146 | 44.497 | 20.077 | 0 | 0
|  |  | 45.9 | 0.013249 | 1 | -0.146 | 44.487 | 20.077 | 0 | 0
|  |  | 45.95 | 0.013145 | 1 | -0.146 | 44.477 | 20.077 | 0 | 0
|  |  | 46 | 0.01304 | 1 | -0.146 | 44.467 | 20.077 | 0 | 0
|  |  | 46.05 | 0.012853 | 1 | -0.146 | 44.457 | 20.077 | 0 | 0
|  |  | 46.1 | 0.012666 | 1 | -0.146 | 44.447 | 20.077 | 0 | 0
|  |  | 46.15 | 0.012478 | 1 | -0.146 | 44.437 | 20.077 | 0 | 0
|  |  | 46.2 | 0.012291 | 1 | -0.146 | 44.427 | 20.077 | 0 | 0
|  |  | 46.25 | 0.012217 | 1 | -0.146 | 44.417 | 20.077 | 0 | 0
|  |  | 46.3 | 0.012142 | 1 | -0.146 | 44.407 | 20.077 | 0 | 0
|  |  | 46.35 | 0.012068 | 1 | -0.146 | 44.397 | 20.077 | 0 | 0
|  |  | 46.4 | 0.011994 | 1 | -0.146 | 44.387 | 20.077 | 0 | 0
|  |  | 46.45 | 0.01184 | 1 | -0.146 | 44.377 | 20.077 | 0 | 0
|  |  | 46.5 | 0.011687 | 1 | -0.146 | 44.367 | 20.077 | 0 | 0
|  |  | 46.55 | 0.011534 | 1 | -0.146 | 44.357 | 20.077 | 0 | 0
|  |  | 46.6 | 0.011381 | 1 | -0.146 | 44.347 | 20.077 | 0 | 0
|  |  | 46.65 | 0.011265 | 1 | -0.146 | 44.337 | 20.077 | 0 | 0
|  |  | 46.7 | 0.011149 | 1 | -0.146 | 44.327 | 20.077 | 0 | 0
|  |  | 46.75 | 0.011034 | 1 | -0.146 | 44.317 | 20.077 | 0 | 0
|  |  | 46.8 | 0.010918 | 1 | -0.146 | 44.307 | 20.077 | 0 | 0
|  |  | 46.85 | 0.010836 | 1 | -0.146 | 44.297 | 20.077 | 0 | 0
|  |  | 46.9 | 0.010754 | 1 | -0.146 | 44.287 | 20.077 | 0 | 0
|  |  | 46.95 | 0.010673 | 1 | -0.146 | 44.277 | 20.077 | 0 | 0
|  |  | 47 | 0.010591 | 1 | -0.146 | 44.267 | 20.077 | 0 | 0
|  |  | 47.05 | 0.010425 | 1 | -0.146 | 44.257 | 20.077 | 0 | 0
|  |  | 47.1 | 0.010259 | 1 | -0.146 | 44.247 | 20.077 | 0 | 0
|  |  | 47.15 | 0.010094 | 1 | -0.146 | 44.237 | 20.077 | 0 | 0
|  |  | 47.2 | 0.0099279 | 1 | -0.146 | 44.227 | 20.077 | 0 | 0
|  |  | 47.25 | 0.0098594 | 1 | -0.146 | 44.217 | 20.077 | 0 | 0
|  |  | 47.3 | 0.0097909 | 1 | -0.146 | 44.207 | 20.077 | 0 | 0
|  |  | 47.35 | 0.0097224 | 1 | -0.146 | 44.197 | 20.077 | 0 | 0
|  |  | 47.4 | 0.0096538 | 1 | -0.146 | 44.187 | 20.077 | 0 | 0
|  |  | 47.45 | 0.0095285 | 1 | -0.146 | 44.177 | 20.077 | 0 | 0
|  |  | 47.5 | 0.0094032 | 1 | -0.146 | 44.167 | 20.077 | 0 | 0
|  |  | 47.55 | 0.0092779 | 1 | -0.146 | 44.157 | 20.077 | 0 | 0
|  |  | 47.6 | 0.0091526 | 1 | -0.146 | 44.147 | 20.077 | 0 | 0
|  |  | 47.65 | 0.0090431 | 1 | -0.146 | 44.137 | 20.076 | 0 | 0
|  |  | 47.7 | 0.0089336 | 1 | -0.146 | 44.127 | 20.076 | 0 | 0
|  |  | 47.75 | 0.008824 | 1 | -0.146 | 44.117 | 20.076 | 0 | 0
|  |  | 47.8 | 0.0087145 | 1 | -0.146 | 44.107 | 20.076 | 0 | 0
|  |  | 47.85 | 0.0086467 | 1 | -0.146 | 44.097 | 20.076 | 0 | 0
|  |  | 47.9 | 0.0085789 | 1 | -0.146 | 44.087 | 20.076 | 0 | 0
|  |  | 47.95 | 0.008511 | 1 | -0.146 | 44.077 | 20.076 | 0 | 0
|  |  | 48 | 0.0084432 | 1 | -0.146 | 44.067 | 20.076 | 0 | 0
|  |  | 48.05 | 0.0083042 | 1 | -0.146 | 44.057 | 20.076 | 0 | 0
|  |  | 48.1 | 0.0081651 | 1 | -0.146 | 44.047 | 20.076 | 0 | 0
|  |  | 48.15 | 0.0080261 | 1 | -0.146 | 44.037 | 20.076 | 0 | 0
|  |  | 48.2 | 0.007887 | 1 | -0.146 | 44.027 | 20.076 | 0 | 0
|  |  | 48.25 | 0.0078258 | 1 | -0.146 | 44.017 | 20.076 | 0 | 0
|  |  | 48.3 | 0.0077646 | 1 | -0.146 | 44.007 | 20.076 | 0 | 0
|  |  | 48.35 | 0.0077034 | 1 | -0.146 | 43.997 | 20.076 | 0 | 0
|  |  | 48.4 | 0.0076421 | 1 | -0.146 | 43.987 | 20.076 | 0 | 0
|  |  | 48.45 | 0.0075458 | 1 | -0.146 | 43.977 | 20.076 | 0 | 0
|  |  | 48.5 | 0.0074494 | 1 | -0.146 | 43.967 | 20.076 | 0 | 0
|  |  | 48.55 | 0.007353 | 1 | -0.146 | 43.957 | 20.076 | 0 | 0
|  |  | 48.6 | 0.0072566 | 1 | -0.146 | 43.947 | 20.076 | 0 | 0
|  |  | 48.65 | 0.0071591 | 1 | -0.146 | 43.937 | 20.076 | 0 | 0
|  |  | 48.7 | 0.0070616 | 1 | -0.146 | 43.927 | 20.076 | 0 | 0
|  |  | 48.75 | 0.0069641 | 1 | -0.146 | 43.917 | 20.076 | 0 | 0
|  |  | 48.8 | 0.0068666 | 1 | -0.146 | 43.907 | 20.076 | 0 | 0
|  |  | 48.85 | 0.0068138 | 1 | -0.146 | 43.897 | 20.076 | 0 | 0
|  |  | 48.9 | 0.0067611 | 1 | -0.146 | 43.887 | 20.076 | 0 | 0
|  |  | 48.95 | 0.0067083 | 1 | -0.146 | 43.877 | 20.076 | 0 | 0
|  |  | 49 | 0.0066556 | 1 | -0.146 | 43.867 | 20.076 | 0 | 0
|  |  | 49.05 | 0.006542 | 1 | -0.146 | 43.857 | 20.076 | 0 | 0
|  |  | 49.1 | 0.0064285 | 1 | -0.146 | 43.847 | 20.076 | 0 | 0
|  |  | 49.15 | 0.0063149 | 1 | -0.146 | 43.837 | 20.076 | 0 | 0
|  |  | 49.2 | 0.0062014 | 1 | -0.146 | 43.827 | 20.076 | 0 | 0
|  |  | 49.25 | 0.0061479 | 1 | -0.146 | 43.817 | 20.076 | 0 | 0
|  |  | 49.3 | 0.0060944 | 1 | -0.146 | 43.807 | 20.076 | 0 | 0
|  |  | 49.35 | 0.0060409 | 1 | -0.146 | 43.797 | 20.076 | 0 | 0
|  |  | 49.4 | 0.0059875 | 1 | -0.146 | 43.787 | 20.076 | 0 | 0
|  |  | 49.45 | 0.0059145 | 1 | -0.146 | 43.777 | 20.076 | 0 | 0
|  |  | 49.5 | 0.0058415 | 1 | -0.146 | 43.767 | 20.076 | 0 | 0
|  |  | 49.55 | 0.0057686 | 1 | -0.146 | 43.757 | 20.076 | 0 | 0
|  |  | 49.6 | 0.0056956 | 1 | -0.146 | 43.747 | 20.076 | 0 | 0
|  |  | 49.65 | 0.0056106 | 1 | -0.146 | 43.737 | 20.076 | 0 | 0
|  |  | 49.7 | 0.0055255 | 1 | -0.146 | 43.727 | 20.076 | 0 | 0
|  |  | 49.75 | 0.0054405 | 1 | -0.146 | 43.717 | 20.076 | 0 | 0
|  |  | 49.8 | 0.0053554 | 1 | -0.146 | 43.707 | 20.076 | 0 | 0
|  |  | 49.85 | 0.0053139 | 1 | -0.146 | 43.697 | 20.076 | 0 | 0
|  |  | 49.9 | 0.0052724 | 1 | -0.146 | 43.687 | 20.076 | 0 | 0
|  |  | 49.95 | 0.0052308 | 1 | -0.146 | 43.677 | 20.076 | 0 | 0
|  |  | 50 | 0.0051893 | 1 | -0.146 | 43.667 | 20.076 | 0 | 0
|  |  | 50.05 | 0.0050992 | 1 | -0.146 | 43.657 | 20.076 | 0 | 0
|  |  | 50.1 | 0.0050092 | 1 | -0.146 | 43.647 | 20.076 | 0 | 0
|  |  | 50.15 | 0.0049191 | 1 | -0.146 | 43.637 | 20.076 | 0 | 0
|  |  | 50.2 | 0.0048291 | 1 | -0.146 | 43.627 | 20.076 | 0 | 0
|  |  | 50.25 | 0.0047824 | 1 | -0.146 | 43.617 | 20.076 | 0 | 0
|  |  | 50.3 | 0.0047357 | 1 | -0.146 | 43.607 | 20.076 | 0 | 0
|  |  | 50.35 | 0.0046891 | 1 | -0.146 | 43.597 | 20.076 | 0 | 0
|  |  | 50.4 | 0.0046424 | 1 | -0.146 | 43.587 | 20.076 | 0 | 0
|  |  | 50.45 | 0.0045884 | 1 | -0.146 | 43.577 | 20.076 | 0 | 0
|  |  | 50.5 | 0.0045344 | 1 | -0.146 | 43.567 | 20.076 | 0 | 0
|  |  | 50.55 | 0.0044804 | 1 | -0.146 | 43.557 | 20.076 | 0 | 0
|  |  | 50.6 | 0.0044264 | 1 | -0.146 | 43.547 | 20.076 | 0 | 0
|  |  | 50.65 | 0.0043534 | 1 | -0.146 | 43.537 | 20.076 | 0 | 0
|  |  | 50.7 | 0.0042803 | 1 | -0.146 | 43.527 | 20.076 | 0 | 0
|  |  | 50.75 | 0.0042073 | 1 | -0.146 | 43.517 | 20.076 | 0 | 0
|  |  | 50.8 | 0.0041342 | 1 | -0.146 | 43.507 | 20.076 | 0 | 0
|  |  | 50.85 | 0.0041017 | 1 | -0.146 | 43.497 | 20.076 | 0 | 0
|  |  | 50.9 | 0.0040693 | 1 | -0.146 | 43.487 | 20.076 | 0 | 0
|  |  | 50.95 | 0.0040368 | 1 | -0.146 | 43.477 | 20.076 | 0 | 0
|  |  | 51 | 0.0040043 | 1 | -0.146 | 43.467 | 20.076 | 0 | 0
|  |  | 51.05 | 0.0039339 | 1 | -0.146 | 43.457 | 20.076 | 0 | 0
|  |  | 51.1 | 0.0038634 | 1 | -0.146 | 43.447 | 20.075 | 0 | 0
|  |  | 51.15 | 0.003793 | 1 | -0.146 | 43.437 | 20.075 | 0 | 0
|  |  | 51.2 | 0.0037225 | 1 | -0.146 | 43.427 | 20.075 | 0 | 0
|  |  | 51.25 | 0.0036817 | 1 | -0.146 | 43.417 | 20.075 | 0 | 0
|  |  | 51.3 | 0.0036409 | 1 | -0.146 | 43.407 | 20.075 | 0 | 0
|  |  | 51.35 | 0.0036001 | 1 | -0.146 | 43.397 | 20.075 | 0 | 0
|  |  | 51.4 | 0.0035593 | 1 | -0.146 | 43.387 | 20.075 | 0 | 0
|  |  | 51.45 | 0.0035188 | 1 | -0.146 | 43.377 | 20.075 | 0 | 0
|  |  | 51.5 | 0.0034784 | 1 | -0.146 | 43.367 | 20.075 | 0 | 0
|  |  | 51.55 | 0.0034379 | 1 | -0.146 | 43.357 | 20.075 | 0 | 0
|  |  | 51.6 | 0.0033974 | 1 | -0.146 | 43.347 | 20.075 | 0 | 0
|  |  | 51.65 | 0.0033368 | 1 | -0.146 | 43.337 | 20.075 | 0 | 0
|  |  | 51.7 | 0.0032762 | 1 | -0.146 | 43.327 | 20.075 | 0 | 0
|  |  | 51.75 | 0.0032156 | 1 | -0.146 | 43.317 | 20.075 | 0 | 0
|  |  | 51.8 | 0.003155 | 1 | -0.146 | 43.307 | 20.075 | 0 | 0
|  |  | 51.85 | 0.0031293 | 1 | -0.146 | 43.297 | 20.075 | 0 | 0
|  |  | 51.9 | 0.0031036 | 1 | -0.146 | 43.287 | 20.075 | 0 | 0
|  |  | 51.95 | 0.003078 | 1 | -0.146 | 43.277 | 20.075 | 0 | 0
|  |  | 52 | 0.0030523 | 1 | -0.146 | 43.267 | 20.075 | 0 | 0
|  |  | 52.05 | 0.0029986 | 1 | -0.146 | 43.257 | 20.075 | 0 | 0
|  |  | 52.1 | 0.002945 | 1 | -0.146 | 43.247 | 20.075 | 0 | 0
|  |  | 52.15 | 0.0028913 | 1 | -0.146 | 43.237 | 20.075 | 0 | 0
|  |  | 52.2 | 0.0028376 | 1 | -0.146 | 43.227 | 20.075 | 0 | 0
|  |  | 52.25 | 0.002803 | 1 | -0.146 | 43.217 | 20.075 | 0 | 0
|  |  | 52.3 | 0.0027683 | 1 | -0.146 | 43.207 | 20.075 | 0 | 0
|  |  | 52.35 | 0.0027336 | 1 | -0.146 | 43.197 | 20.075 | 0 | 0
|  |  | 52.4 | 0.0026989 | 1 | -0.146 | 43.187 | 20.075 | 0 | 0
|  |  | 52.45 | 0.002669 | 1 | -0.146 | 43.177 | 20.075 | 0 | 0
|  |  | 52.5 | 0.0026392 | 1 | -0.146 | 43.167 | 20.075 | 0 | 0
|  |  | 52.55 | 0.0026093 | 1 | -0.146 | 43.157 | 20.075 | 0 | 0
|  |  | 52.6 | 0.0025794 | 1 | -0.146 | 43.147 | 20.075 | 0 | 0
|  |  | 52.65 | 0.0025304 | 1 | -0.146 | 43.137 | 20.075 | 0 | 0
|  |  | 52.7 | 0.0024814 | 1 | -0.146 | 43.127 | 20.075 | 0 | 0
|  |  | 52.75 | 0.0024324 | 1 | -0.146 | 43.117 | 20.075 | 0 | 0
|  |  | 52.8 | 0.0023834 | 1 | -0.146 | 43.107 | 20.075 | 0 | 0
|  |  | 52.85 | 0.0023625 | 1 | -0.146 | 43.097 | 20.075 | 0 | 0
|  |  | 52.9 | 0.0023416 | 1 | -0.146 | 43.087 | 20.075 | 0 | 0
|  |  | 52.95 | 0.0023207 | 1 | -0.146 | 43.077 | 20.075 | 0 | 0
|  |  | 53 | 0.0022999 | 1 | -0.146 | 43.067 | 20.075 | 0 | 0
|  |  | 53.05 | 0.0022599 | 1 | -0.146 | 43.057 | 20.075 | 0 | 0
|  |  | 53.1 | 0.0022198 | 1 | -0.146 | 43.047 | 20.075 | 0 | 0
|  |  | 53.15 | 0.0021798 | 1 | -0.146 | 43.037 | 20.075 | 0 | 0
|  |  | 53.2 | 0.0021398 | 1 | -0.146 | 43.027 | 20.075 | 0 | 0
|  |  | 53.25 | 0.0021101 | 1 | -0.146 | 43.017 | 20.075 | 0 | 0
|  |  | 53.3 | 0.0020804 | 1 | -0.146 | 43.007 | 20.075 | 0 | 0
|  |  | 53.35 | 0.0020506 | 1 | -0.146 | 42.997 | 20.075 | 0 | 0
|  |  | 53.4 | 0.0020209 | 1 | -0.146 | 42.987 | 20.075 | 0 | 0
|  |  | 53.45 | 0.0019992 | 1 | -0.146 | 42.978 | 20.075 | 0 | 0
|  |  | 53.5 | 0.0019776 | 1 | -0.146 | 42.968 | 20.075 | 0 | 0
|  |  | 53.55 | 0.0019559 | 1 | -0.146 | 42.958 | 20.075 | 0 | 0
|  |  | 53.6 | 0.0019342 | 1 | -0.146 | 42.948 | 20.075 | 0 | 0
|  |  | 53.65 | 0.0018955 | 1 | -0.146 | 42.938 | 20.075 | 0 | 0
|  |  | 53.7 | 0.0018569 | 1 | -0.146 | 42.928 | 20.075 | 0 | 0
|  |  | 53.75 | 0.0018183 | 1 | -0.146 | 42.918 | 20.075 | 0 | 0
|  |  | 53.8 | 0.0017796 | 1 | -0.146 | 42.908 | 20.075 | 0 | 0
|  |  | 53.85 | 0.0017628 | 1 | -0.146 | 42.898 | 20.075 | 0 | 0
|  |  | 53.9 | 0.0017461 | 1 | -0.146 | 42.888 | 20.075 | 0 | 0
|  |  | 53.95 | 0.0017293 | 1 | -0.146 | 42.878 | 20.075 | 0 | 0
|  |  | 54 | 0.0017125 | 1 | -0.146 | 42.868 | 20.075 | 0 | 0
|  |  | 54.05 | 0.0016837 | 1 | -0.146 | 42.858 | 20.075 | 0 | 0
|  |  | 54.1 | 0.001655 | 1 | -0.146 | 42.848 | 20.075 | 0 | 0
|  |  | 54.15 | 0.0016262 | 1 | -0.146 | 42.838 | 20.075 | 0 | 0
|  |  | 54.2 | 0.0015974 | 1 | -0.146 | 42.828 | 20.075 | 0 | 0
|  |  | 54.25 | 0.0015728 | 1 | -0.146 | 42.818 | 20.075 | 0 | 0
|  |  | 54.3 | 0.0015483 | 1 | -0.146 | 42.808 | 20.075 | 0 | 0
|  |  | 54.35 | 0.0015237 | 1 | -0.146 | 42.798 | 20.075 | 0 | 0
|  |  | 54.4 | 0.0014992 | 1 | -0.146 | 42.788 | 20.075 | 0 | 0
|  |  | 54.45 | 0.0014833 | 1 | -0.146 | 42.778 | 20.075 | 0 | 0
|  |  | 54.5 | 0.0014675 | 1 | -0.146 | 42.768 | 20.075 | 0 | 0
|  |  | 54.55 | 0.0014516 | 1 | -0.146 | 42.758 | 20.074 | 0 | 0
|  |  | 54.6 | 0.0014358 | 1 | -0.146 | 42.748 | 20.074 | 0 | 0
|  |  | 54.65 | 0.0014062 | 1 | -0.146 | 42.738 | 20.074 | 0 | 0
|  |  | 54.7 | 0.0013766 | 1 | -0.146 | 42.728 | 20.074 | 0 | 0
|  |  | 54.75 | 0.001347 | 1 | -0.146 | 42.718 | 20.074 | 0 | 0
|  |  | 54.8 | 0.0013174 | 1 | -0.146 | 42.708 | 20.074 | 0 | 0
|  |  | 54.85 | 0.0013082 | 1 | -0.146 | 42.698 | 20.074 | 0 | 0
|  |  | 54.9 | 0.001299 | 1 | -0.146 | 42.688 | 20.074 | 0 | 0
|  |  | 54.95 | 0.0012898 | 1 | -0.146 | 42.678 | 20.074 | 0 | 0
|  | END_DATA |  |  |  | -0.146 | 42.668 | 20.074 | 0 | 0
| END_SCAN 1 |  |  |  |  | -0.146 | 42.658 | 20.074 | 0 | 0
END_SCAN_DATA | [5 empty] | -0.146 | 42.648 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.638 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.628 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.618 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.608 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.598 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.588 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.578 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.568 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.558 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.548 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.538 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.528 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.518 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.508 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.498 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.488 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.478 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.468 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.458 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.448 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.438 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.428 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.418 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.408 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.398 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.388 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.378 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.368 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.358 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.348 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.338 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.328 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.318 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.308 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.298 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.288 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.278 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.268 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.258 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.248 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.238 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.228 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.218 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.208 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.198 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.188 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.178 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.168 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.158 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.148 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.138 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.128 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.118 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.108 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.098 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.088 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.078 | 20.074 | 0 | 0
[6 empty] | -0.146 | 42.068 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.058 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.048 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.038 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.028 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.018 | 20.073 | 0 | 0
[6 empty] | -0.146 | 42.008 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.998 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.988 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.978 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.968 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.958 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.948 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.938 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.928 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.918 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.908 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.898 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.888 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.878 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.868 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.858 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.848 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.838 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.828 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.818 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.808 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.798 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.788 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.778 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.768 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.758 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.748 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.738 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.728 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.718 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.708 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.698 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.688 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.678 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.668 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.658 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.648 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.638 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.628 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.618 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.608 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.598 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.588 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.578 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.568 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.558 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.548 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.538 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.528 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.518 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.508 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.498 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.488 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.478 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.468 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.458 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.448 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.438 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.428 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.418 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.408 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.398 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.388 | 20.073 | 0 | 0
[6 empty] | -0.146 | 41.378 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.368 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.358 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.348 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.338 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.328 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.318 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.308 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.298 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.288 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.278 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.268 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.258 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.248 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.238 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.228 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.218 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.208 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.198 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.188 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.178 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.168 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.158 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.148 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.138 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.128 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.118 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.108 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.098 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.088 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.078 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.068 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.058 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.048 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.038 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.028 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.018 | 20.072 | 0 | 0
[6 empty] | -0.146 | 41.008 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.998 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.988 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.978 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.968 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.958 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.948 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.938 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.928 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.918 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.908 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.898 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.888 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.878 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.868 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.858 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.848 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.838 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.828 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.818 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.808 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.798 | 20.072 | 0 | 0
[6 empty] | -0.146 | 40.788 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.778 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.768 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.758 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.748 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.738 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.728 | 20.072 | 0.001 | 0.1
[6 empty] | -0.146 | 40.718 | 20.072 | 0.002 | 0.2
[6 empty] | -0.146 | 40.708 | 20.072 | 0.002 | 0.2
[6 empty] | -0.146 | 40.698 | 20.072 | 0.002 | 0.2
[6 empty] | -0.146 | 40.688 | 20.071 | 0.002 | 0.2
[6 empty] | -0.146 | 40.678 | 20.071 | 0.002 | 0.2
[6 empty] | -0.146 | 40.668 | 20.071 | 0.002 | 0.2
[6 empty] | -0.146 | 40.658 | 20.071 | 0.002 | 0.2
[6 empty] | -0.146 | 40.648 | 20.071 | 0.002 | 0.2
[6 empty] | -0.146 | 40.638 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.628 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.618 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.608 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.598 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.588 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.578 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.568 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.558 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.548 | 20.071 | 0.003 | 0.3
[6 empty] | -0.146 | 40.538 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.528 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.518 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.508 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.498 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.488 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.478 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.468 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.458 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.448 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.438 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.428 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.418 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.408 | 20.071 | 0.004 | 0.4
[6 empty] | -0.146 | 40.398 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.388 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.378 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.368 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.358 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.348 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.338 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.328 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.318 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.308 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.298 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.288 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.278 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.268 | 20.071 | 0.005 | 0.5
[6 empty] | -0.146 | 40.258 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.248 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.238 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.228 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.218 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.208 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.198 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.188 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.178 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.168 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.158 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.148 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.138 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.128 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.118 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.108 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.098 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.088 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.078 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.068 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.058 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.048 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.038 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.028 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.018 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 40.008 | 20.071 | 0.006 | 0.6
[6 empty] | -0.146 | 39.998 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.988 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.978 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.969 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.959 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.949 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.939 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.929 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.919 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.909 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.899 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.889 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.879 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.869 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.859 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.849 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.839 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.829 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.819 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.809 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.799 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.789 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.779 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.769 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.759 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.749 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.739 | 20.07 | 0.006 | 0.6
[6 empty] | -0.146 | 39.729 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.719 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.709 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.699 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.689 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.679 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.669 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.659 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.649 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.639 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.629 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.619 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.609 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.599 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.589 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.579 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.569 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.559 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.549 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.539 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.529 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.519 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.509 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.499 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.489 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.479 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.469 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.459 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.449 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.439 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.429 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.419 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.409 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.399 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.389 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.379 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.369 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.359 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.349 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.339 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.329 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.319 | 20.07 | 0.007 | 0.7
[6 empty] | -0.146 | 39.309 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.299 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.289 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.279 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.269 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.259 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.249 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.239 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.229 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.219 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.209 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.199 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.189 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.179 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.169 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.159 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.149 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.139 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.129 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.119 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.109 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.099 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.089 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.079 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.069 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.059 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.049 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.039 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.029 | 20.069 | 0.007 | 0.7
[6 empty] | -0.146 | 39.019 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 39.009 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.999 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.989 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.979 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.969 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.959 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.949 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.939 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.929 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.919 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.909 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.899 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.889 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.879 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.869 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.859 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.849 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.839 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.829 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.819 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.809 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.799 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.789 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.779 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.769 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.759 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.749 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.739 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.729 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.719 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.709 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.699 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.689 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.679 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.669 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.659 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.649 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.639 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.629 | 20.069 | 0.008 | 0.8
[6 empty] | -0.146 | 38.619 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.609 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.599 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.589 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.579 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.569 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.559 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.549 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.539 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.529 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.519 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.509 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.499 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.489 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.479 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.469 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.459 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.449 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.439 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.429 | 20.068 | 0.008 | 0.8
[6 empty] | -0.146 | 38.419 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.409 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.399 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.389 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.379 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.369 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.359 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.349 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.339 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.329 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.319 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.309 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.299 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.289 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.279 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.269 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.259 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.249 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.239 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.229 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.219 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.209 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.199 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.189 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.179 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.169 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.159 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.149 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.139 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.129 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.119 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.109 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.099 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.089 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.079 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.069 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.059 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.049 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.039 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.029 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.019 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 38.009 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.999 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.989 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.979 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.969 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.959 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.949 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.939 | 20.068 | 0.009 | 0.9
[6 empty] | -0.146 | 37.929 | 20.067 | 0.009 | 0.9
[6 empty] | -0.146 | 37.919 | 20.067 | 0.009 | 0.9
[6 empty] | -0.146 | 37.909 | 20.067 | 0.009 | 0.9
[6 empty] | -0.146 | 37.899 | 20.067 | 0.009 | 0.9
[6 empty] | -0.146 | 37.889 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.879 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.869 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.859 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.849 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.839 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.829 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.819 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.809 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.799 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.789 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.779 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.769 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.759 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.749 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.739 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.729 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.719 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.709 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.699 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.689 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.679 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.669 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.659 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.649 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.639 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.629 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.619 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.609 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.599 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.589 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.579 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.569 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.559 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.549 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.539 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.529 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.519 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.509 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.499 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.489 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.479 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.469 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.459 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.449 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.439 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.429 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.419 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.409 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.399 | 20.067 | 0.01 | 1
[6 empty] | -0.146 | 37.389 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.379 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.369 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.359 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.349 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.339 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.329 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.319 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.309 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.299 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.289 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.279 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.269 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.259 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.249 | 20.067 | 0.011 | 1.1
[6 empty] | -0.146 | 37.239 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.229 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.219 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.209 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.199 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.189 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.179 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.169 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.159 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.149 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.139 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.129 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.119 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.109 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.099 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.089 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.079 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.069 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.059 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.049 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.039 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.029 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.019 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 37.009 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.999 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.989 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.979 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.97 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.96 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.95 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.94 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.93 | 20.066 | 0.011 | 1.1
[6 empty] | -0.146 | 36.92 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.91 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.9 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.89 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.88 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.87 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.86 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.85 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.84 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.83 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.82 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.81 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.8 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.79 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.78 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.77 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.76 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.75 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.74 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.73 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.72 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.71 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.7 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.69 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.68 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.67 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.66 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.65 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.64 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.63 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.62 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.61 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.6 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.59 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.58 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.57 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.56 | 20.066 | 0.012 | 1.2
[6 empty] | -0.146 | 36.55 | 20.065 | 0.012 | 1.2
[6 empty] | -0.146 | 36.54 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.53 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.52 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.51 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.5 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.49 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.48 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.47 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.46 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.45 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.44 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.43 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.42 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.41 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.4 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.39 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.38 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.37 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.36 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.35 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.34 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.33 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.32 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.31 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.3 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.29 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.28 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.27 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.26 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.25 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.24 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.23 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.22 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.21 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.2 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.19 | 20.065 | 0.013 | 1.3
[6 empty] | -0.146 | 36.18 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.17 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.16 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.15 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.14 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.13 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.12 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.11 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.1 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.09 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.08 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.07 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.06 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.05 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.04 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.03 | 20.065 | 0.014 | 1.4
[6 empty] | -0.146 | 36.02 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 36.01 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 36 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.99 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.98 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.97 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.96 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.95 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.94 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.93 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.92 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.91 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.9 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.89 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.88 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.87 | 20.065 | 0.015 | 1.5
[6 empty] | -0.146 | 35.86 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.85 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.84 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.83 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.82 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.81 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.8 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.79 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.78 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.77 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.76 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.75 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.74 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.73 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.72 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.71 | 20.064 | 0.016 | 1.6
[6 empty] | -0.146 | 35.7 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.69 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.68 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.67 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.66 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.65 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.64 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.63 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.62 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.61 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.6 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.59 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.58 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.57 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.56 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.55 | 20.064 | 0.017 | 1.7
[6 empty] | -0.146 | 35.54 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.53 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.52 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.51 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.5 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.49 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.48 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.47 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.46 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.45 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.44 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.43 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.42 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.41 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.4 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.39 | 20.064 | 0.018 | 1.8
[6 empty] | -0.146 | 35.38 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.37 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.36 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.35 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.34 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.33 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.32 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.31 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.3 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.29 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.28 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.27 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.26 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.25 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.24 | 20.064 | 0.019 | 1.9
[6 empty] | -0.146 | 35.23 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.22 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.21 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.2 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.19 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.18 | 20.064 | 0.02 | 2
[6 empty] | -0.146 | 35.17 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.16 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.15 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.14 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.13 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.12 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.11 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.1 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.09 | 20.063 | 0.02 | 2
[6 empty] | -0.146 | 35.08 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.07 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.06 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.05 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.04 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.03 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.02 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35.01 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 35 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.99 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.98 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.97 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.96 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.95 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.94 | 20.063 | 0.021 | 2.1
[6 empty] | -0.146 | 34.93 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.92 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.91 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.9 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.89 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.88 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.87 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.86 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.85 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.84 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.83 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.82 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.81 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.8 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.79 | 20.063 | 0.022 | 2.2
[6 empty] | -0.146 | 34.78 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.77 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.76 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.75 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.74 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.73 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.72 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.71 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.7 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.69 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.68 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.67 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.66 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.65 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.64 | 20.063 | 0.023 | 2.3
[6 empty] | -0.146 | 34.63 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.62 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.61 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.6 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.59 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.58 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.57 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.56 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.55 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.54 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.53 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.52 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.51 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.5 | 20.063 | 0.024 | 2.4
[6 empty] | -0.146 | 34.49 | 20.063 | 0.025 | 2.5
[6 empty] | -0.146 | 34.48 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.47 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.46 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.45 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.44 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.43 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.42 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.41 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.4 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.39 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.38 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.37 | 20.062 | 0.025 | 2.5
[6 empty] | -0.146 | 34.36 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.35 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.34 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.33 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.32 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.31 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.3 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.29 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.28 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.27 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.26 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.25 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.24 | 20.062 | 0.026 | 2.6
[6 empty] | -0.146 | 34.23 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.22 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.21 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.2 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.19 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.18 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.17 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.16 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.15 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.14 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.13 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.12 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.11 | 20.062 | 0.027 | 2.7
[6 empty] | -0.146 | 34.1 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.09 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.08 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.07 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.06 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.05 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.04 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.03 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.02 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34.01 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 34 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 33.99 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 33.98 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 33.971 | 20.062 | 0.028 | 2.8
[6 empty] | -0.146 | 33.961 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.951 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.941 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.931 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.921 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.911 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.901 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.891 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.881 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.871 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.861 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.851 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.841 | 20.062 | 0.029 | 2.9
[6 empty] | -0.146 | 33.831 | 20.062 | 0.03 | 3
[6 empty] | -0.146 | 33.821 | 20.062 | 0.03 | 3
[6 empty] | -0.146 | 33.811 | 20.062 | 0.03 | 3
[6 empty] | -0.146 | 33.801 | 20.062 | 0.03 | 3
[6 empty] | -0.146 | 33.791 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.781 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.771 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.761 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.751 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.741 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.731 | 20.061 | 0.03 | 3
[6 empty] | -0.146 | 33.721 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.711 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.701 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.691 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.681 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.671 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.661 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.651 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.641 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.631 | 20.061 | 0.031 | 3.1
[6 empty] | -0.146 | 33.621 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.611 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.601 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.591 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.581 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.571 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.561 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.551 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.541 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.531 | 20.061 | 0.032 | 3.2
[6 empty] | -0.146 | 33.521 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.511 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.501 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.491 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.481 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.471 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.461 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.451 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.441 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.431 | 20.061 | 0.033 | 3.3
[6 empty] | -0.146 | 33.421 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.411 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.401 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.391 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.381 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.371 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.361 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.351 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.341 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.331 | 20.061 | 0.034 | 3.4
[6 empty] | -0.146 | 33.321 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.311 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.301 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.291 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.281 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.271 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.261 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.251 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.241 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.231 | 20.061 | 0.035 | 3.5
[6 empty] | -0.146 | 33.221 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.211 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.201 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.191 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.181 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.171 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.161 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.151 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.141 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.131 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.121 | 20.061 | 0.036 | 3.6
[6 empty] | -0.146 | 33.111 | 20.061 | 0.037 | 3.7
[6 empty] | -0.146 | 33.101 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.091 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.081 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.071 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.061 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.051 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.041 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.031 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.021 | 20.06 | 0.037 | 3.7
[6 empty] | -0.146 | 33.011 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 33.001 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.991 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.981 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.971 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.961 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.951 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.941 | 20.06 | 0.038 | 3.8
[6 empty] | -0.146 | 32.931 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.921 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.911 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.901 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.891 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.881 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.871 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.861 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.851 | 20.06 | 0.039 | 3.9
[6 empty] | -0.146 | 32.841 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.831 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.821 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.811 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.801 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.791 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.781 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.771 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.761 | 20.06 | 0.04 | 4
[6 empty] | -0.146 | 32.751 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.741 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.731 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.721 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.711 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.701 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.691 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.681 | 20.06 | 0.041 | 4.1
[6 empty] | -0.146 | 32.671 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.661 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.651 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.641 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.631 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.621 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.611 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.601 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.591 | 20.06 | 0.042 | 4.2
[6 empty] | -0.146 | 32.581 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.571 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.561 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.551 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.541 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.531 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.521 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.511 | 20.06 | 0.043 | 4.3
[6 empty] | -0.146 | 32.501 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.491 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.481 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.471 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.461 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.451 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.441 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.431 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.421 | 20.06 | 0.044 | 4.4
[6 empty] | -0.146 | 32.411 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.401 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.391 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.381 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.371 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.361 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.351 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.341 | 20.059 | 0.045 | 4.5
[6 empty] | -0.146 | 32.331 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.321 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.311 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.301 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.291 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.281 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.271 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.261 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.251 | 20.059 | 0.046 | 4.6
[6 empty] | -0.146 | 32.241 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.231 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.221 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.211 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.201 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.191 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.181 | 20.059 | 0.047 | 4.7
[6 empty] | -0.146 | 32.171 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.161 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.151 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.141 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.131 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.121 | 20.059 | 0.048 | 4.8
[6 empty] | -0.146 | 32.111 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.101 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.091 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.081 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.071 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.061 | 20.059 | 0.049 | 4.9
[6 empty] | -0.146 | 32.051 | 20.059 | 0.05 | 5
[6 empty] | -0.146 | 32.041 | 20.059 | 0.05 | 5
[6 empty] | -0.146 | 32.031 | 20.059 | 0.05 | 5
[6 empty] | -0.146 | 32.021 | 20.059 | 0.05 | 5
[6 empty] | -0.146 | 32.011 | 20.059 | 0.05 | 5
[6 empty] | -0.146 | 32.001 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.991 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.981 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.971 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.961 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.951 | 20.059 | 0.051 | 5.1
[6 empty] | -0.146 | 31.941 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.931 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.921 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.911 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.901 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.891 | 20.059 | 0.052 | 5.2
[6 empty] | -0.146 | 31.881 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.871 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.861 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.851 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.841 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.831 | 20.059 | 0.053 | 5.3
[6 empty] | -0.146 | 31.821 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.811 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.801 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.791 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.781 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.771 | 20.059 | 0.054 | 5.4
[6 empty] | -0.146 | 31.761 | 20.059 | 0.055 | 5.5
[6 empty] | -0.146 | 31.751 | 20.059 | 0.055 | 5.5
[6 empty] | -0.146 | 31.741 | 20.059 | 0.055 | 5.5
[6 empty] | -0.146 | 31.731 | 20.059 | 0.055 | 5.5
[6 empty] | -0.146 | 31.721 | 20.058 | 0.055 | 5.5
[6 empty] | -0.146 | 31.711 | 20.058 | 0.055 | 5.5
[6 empty] | -0.146 | 31.701 | 20.058 | 0.056 | 5.6
[6 empty] | -0.146 | 31.691 | 20.058 | 0.056 | 5.6
[6 empty] | -0.146 | 31.681 | 20.058 | 0.056 | 5.6
[6 empty] | -0.146 | 31.671 | 20.058 | 0.056 | 5.6
[6 empty] | -0.146 | 31.661 | 20.058 | 0.056 | 5.6
[6 empty] | -0.146 | 31.651 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.641 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.631 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.621 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.611 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.601 | 20.058 | 0.057 | 5.7
[6 empty] | -0.146 | 31.591 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.581 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.571 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.561 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.551 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.541 | 20.058 | 0.058 | 5.8
[6 empty] | -0.146 | 31.531 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.521 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.511 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.501 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.491 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.481 | 20.058 | 0.059 | 5.9
[6 empty] | -0.146 | 31.471 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.461 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.451 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.441 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.431 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.421 | 20.058 | 0.06 | 6
[6 empty] | -0.146 | 31.411 | 20.058 | 0.061 | 6.1
[6 empty] | -0.146 | 31.401 | 20.058 | 0.061 | 6.1
[6 empty] | -0.146 | 31.391 | 20.058 | 0.061 | 6.1
[6 empty] | -0.146 | 31.381 | 20.058 | 0.062 | 6.2
[6 empty] | -0.146 | 31.371 | 20.058 | 0.062 | 6.2
[6 empty] | -0.146 | 31.361 | 20.058 | 0.062 | 6.2
[6 empty] | -0.146 | 31.351 | 20.058 | 0.063 | 6.3
[6 empty] | -0.146 | 31.341 | 20.058 | 0.063 | 6.3
[6 empty] | -0.146 | 31.331 | 20.058 | 0.064 | 6.4
[6 empty] | -0.146 | 31.321 | 20.058 | 0.064 | 6.4
[6 empty] | -0.146 | 31.311 | 20.058 | 0.064 | 6.4
[6 empty] | -0.146 | 31.301 | 20.058 | 0.065 | 6.5
[6 empty] | -0.146 | 31.291 | 20.058 | 0.065 | 6.5
[6 empty] | -0.146 | 31.281 | 20.058 | 0.065 | 6.5
[6 empty] | -0.146 | 31.271 | 20.058 | 0.066 | 6.6
[6 empty] | -0.146 | 31.261 | 20.058 | 0.066 | 6.6
[6 empty] | -0.146 | 31.251 | 20.058 | 0.067 | 6.7
[6 empty] | -0.146 | 31.241 | 20.058 | 0.067 | 6.7
[6 empty] | -0.146 | 31.231 | 20.058 | 0.067 | 6.7
[6 empty] | -0.146 | 31.221 | 20.058 | 0.068 | 6.8
[6 empty] | -0.146 | 31.211 | 20.058 | 0.068 | 6.8
[6 empty] | -0.146 | 31.201 | 20.058 | 0.069 | 6.9
[6 empty] | -0.146 | 31.191 | 20.058 | 0.069 | 6.9
[6 empty] | -0.146 | 31.181 | 20.058 | 0.069 | 6.9
[6 empty] | -0.146 | 31.171 | 20.058 | 0.07 | 7
[6 empty] | -0.146 | 31.161 | 20.058 | 0.07 | 7
[6 empty] | -0.146 | 31.151 | 20.058 | 0.07 | 7
[6 empty] | -0.146 | 31.141 | 20.058 | 0.071 | 7.1
[6 empty] | -0.146 | 31.131 | 20.058 | 0.071 | 7.1
[6 empty] | -0.146 | 31.121 | 20.058 | 0.072 | 7.2
[6 empty] | -0.146 | 31.111 | 20.058 | 0.072 | 7.2
[6 empty] | -0.146 | 31.101 | 20.058 | 0.072 | 7.2
[6 empty] | -0.146 | 31.091 | 20.058 | 0.073 | 7.3
[6 empty] | -0.146 | 31.081 | 20.058 | 0.073 | 7.3
[6 empty] | -0.146 | 31.071 | 20.058 | 0.074 | 7.4
[6 empty] | -0.146 | 31.061 | 20.058 | 0.074 | 7.4
[6 empty] | -0.146 | 31.051 | 20.058 | 0.074 | 7.4
[6 empty] | -0.146 | 31.041 | 20.058 | 0.075 | 7.5
[6 empty] | -0.146 | 31.031 | 20.057 | 0.075 | 7.5
[6 empty] | -0.146 | 31.021 | 20.057 | 0.075 | 7.5
[6 empty] | -0.146 | 31.011 | 20.057 | 0.076 | 7.6
[6 empty] | -0.146 | 31.001 | 20.057 | 0.076 | 7.6
[6 empty] | -0.146 | 30.991 | 20.057 | 0.077 | 7.7
[6 empty] | -0.146 | 30.981 | 20.057 | 0.078 | 7.8
[6 empty] | -0.146 | 30.972 | 20.057 | 0.079 | 7.9
[6 empty] | -0.146 | 30.962 | 20.057 | 0.08 | 8
[6 empty] | -0.146 | 30.952 | 20.057 | 0.081 | 8.1
[6 empty] | -0.146 | 30.942 | 20.057 | 0.082 | 8.2
[6 empty] | -0.146 | 30.932 | 20.057 | 0.083 | 8.3
[6 empty] | -0.146 | 30.922 | 20.057 | 0.084 | 8.4
[6 empty] | -0.146 | 30.912 | 20.057 | 0.084 | 8.4
[6 empty] | -0.146 | 30.902 | 20.057 | 0.085 | 8.5
[6 empty] | -0.146 | 30.892 | 20.057 | 0.086 | 8.6
[6 empty] | -0.146 | 30.882 | 20.057 | 0.087 | 8.7
[6 empty] | -0.146 | 30.872 | 20.057 | 0.088 | 8.8
[6 empty] | -0.146 | 30.862 | 20.057 | 0.089 | 8.9
[6 empty] | -0.146 | 30.852 | 20.057 | 0.09 | 9
[6 empty] | -0.146 | 30.842 | 20.057 | 0.091 | 9.1
[6 empty] | -0.146 | 30.832 | 20.057 | 0.092 | 9.2
[6 empty] | -0.146 | 30.822 | 20.057 | 0.092 | 9.2
[6 empty] | -0.146 | 30.812 | 20.057 | 0.093 | 9.3
[6 empty] | -0.146 | 30.802 | 20.057 | 0.095 | 9.5
[6 empty] | -0.146 | 30.792 | 20.057 | 0.098 | 9.8
[6 empty] | -0.146 | 30.782 | 20.057 | 0.101 | 10.1
[6 empty] | -0.146 | 30.772 | 20.057 | 0.104 | 10.4
[6 empty] | -0.146 | 30.762 | 20.057 | 0.107 | 10.7
[6 empty] | -0.146 | 30.752 | 20.057 | 0.11 | 11
[6 empty] | -0.146 | 30.742 | 20.057 | 0.113 | 11.3
[6 empty] | -0.146 | 30.732 | 20.057 | 0.116 | 11.6
[6 empty] | -0.146 | 30.722 | 20.057 | 0.119 | 11.9
[6 empty] | -0.146 | 30.712 | 20.057 | 0.122 | 12.2
[6 empty] | -0.146 | 30.702 | 20.057 | 0.125 | 12.5
[6 empty] | -0.146 | 30.692 | 20.057 | 0.128 | 12.8
[6 empty] | -0.146 | 30.682 | 20.057 | 0.131 | 13.1
[6 empty] | -0.146 | 30.672 | 20.057 | 0.133 | 13.3
[6 empty] | -0.146 | 30.662 | 20.057 | 0.136 | 13.6
[6 empty] | -0.146 | 30.652 | 20.057 | 0.139 | 13.9
[6 empty] | -0.146 | 30.642 | 20.057 | 0.142 | 14.2
[6 empty] | -0.146 | 30.632 | 20.057 | 0.145 | 14.5
[6 empty] | -0.146 | 30.622 | 20.057 | 0.148 | 14.8
[6 empty] | -0.146 | 30.612 | 20.057 | 0.151 | 15.1
[6 empty] | -0.146 | 30.602 | 20.057 | 0.156 | 15.6
[6 empty] | -0.146 | 30.592 | 20.057 | 0.165 | 16.5
[6 empty] | -0.146 | 30.582 | 20.057 | 0.175 | 17.5
[6 empty] | -0.146 | 30.572 | 20.057 | 0.184 | 18.4
[6 empty] | -0.146 | 30.562 | 20.057 | 0.194 | 19.4
[6 empty] | -0.146 | 30.552 | 20.057 | 0.203 | 20.3
[6 empty] | -0.146 | 30.542 | 20.057 | 0.213 | 21.3
[6 empty] | -0.146 | 30.532 | 20.057 | 0.222 | 22.2
[6 empty] | -0.146 | 30.522 | 20.057 | 0.232 | 23.2
[6 empty] | -0.146 | 30.512 | 20.057 | 0.241 | 24.1
[6 empty] | -0.146 | 30.502 | 20.057 | 0.251 | 25.1
[6 empty] | -0.146 | 30.492 | 20.057 | 0.26 | 26
[6 empty] | -0.146 | 30.482 | 20.057 | 0.27 | 27
[6 empty] | -0.146 | 30.472 | 20.057 | 0.279 | 27.9
[6 empty] | -0.146 | 30.462 | 20.057 | 0.288 | 28.8
[6 empty] | -0.146 | 30.452 | 20.057 | 0.298 | 29.8
[6 empty] | -0.146 | 30.442 | 20.057 | 0.307 | 30.7
[6 empty] | -0.146 | 30.432 | 20.057 | 0.317 | 31.7
[6 empty] | -0.146 | 30.422 | 20.057 | 0.326 | 32.6
[6 empty] | -0.146 | 30.412 | 20.057 | 0.336 | 33.6
[6 empty] | -0.146 | 30.402 | 20.057 | 0.344 | 34.4
[6 empty] | -0.146 | 30.392 | 20.057 | 0.349 | 34.9
[6 empty] | -0.146 | 30.382 | 20.057 | 0.355 | 35.5
[6 empty] | -0.146 | 30.372 | 20.057 | 0.36 | 36
[6 empty] | -0.146 | 30.362 | 20.057 | 0.365 | 36.5
[6 empty] | -0.146 | 30.352 | 20.057 | 0.37 | 37
[6 empty] | -0.146 | 30.342 | 20.056 | 0.375 | 37.5
[6 empty] | -0.146 | 30.332 | 20.056 | 0.381 | 38.1
[6 empty] | -0.146 | 30.322 | 20.056 | 0.386 | 38.6
[6 empty] | -0.146 | 30.312 | 20.056 | 0.391 | 39.1
[6 empty] | -0.146 | 30.302 | 20.056 | 0.396 | 39.6
[6 empty] | -0.146 | 30.292 | 20.056 | 0.401 | 40.1
[6 empty] | -0.146 | 30.282 | 20.056 | 0.407 | 40.7
[6 empty] | -0.146 | 30.272 | 20.056 | 0.412 | 41.2
[6 empty] | -0.146 | 30.262 | 20.056 | 0.417 | 41.7
[6 empty] | -0.146 | 30.252 | 20.056 | 0.422 | 42.2
[6 empty] | -0.146 | 30.242 | 20.056 | 0.427 | 42.7
[6 empty] | -0.146 | 30.232 | 20.056 | 0.433 | 43.3
[6 empty] | -0.146 | 30.222 | 20.056 | 0.438 | 43.8
[6 empty] | -0.146 | 30.212 | 20.056 | 0.443 | 44.3
[6 empty] | -0.146 | 30.202 | 20.056 | 0.447 | 44.7
[6 empty] | -0.146 | 30.192 | 20.056 | 0.449 | 44.9
[6 empty] | -0.146 | 30.182 | 20.056 | 0.45 | 45
[6 empty] | -0.146 | 30.172 | 20.056 | 0.452 | 45.2
[6 empty] | -0.146 | 30.162 | 20.056 | 0.453 | 45.3
[6 empty] | -0.146 | 30.152 | 20.056 | 0.455 | 45.5
[6 empty] | -0.146 | 30.142 | 20.056 | 0.456 | 45.6
[6 empty] | -0.146 | 30.132 | 20.056 | 0.458 | 45.8
[6 empty] | -0.146 | 30.122 | 20.056 | 0.459 | 45.9
[6 empty] | -0.146 | 30.112 | 20.056 | 0.461 | 46.1
[6 empty] | -0.146 | 30.102 | 20.056 | 0.463 | 46.3
[6 empty] | -0.146 | 30.092 | 20.056 | 0.464 | 46.4
[6 empty] | -0.146 | 30.082 | 20.056 | 0.466 | 46.6
[6 empty] | -0.146 | 30.072 | 20.056 | 0.467 | 46.7
[6 empty] | -0.146 | 30.062 | 20.056 | 0.469 | 46.9
[6 empty] | -0.146 | 30.052 | 20.056 | 0.47 | 47
[6 empty] | -0.146 | 30.042 | 20.056 | 0.472 | 47.2
[6 empty] | -0.146 | 30.032 | 20.056 | 0.473 | 47.3
[6 empty] | -0.146 | 30.022 | 20.056 | 0.475 | 47.5
[6 empty] | -0.146 | 30.012 | 20.056 | 0.476 | 47.6
[6 empty] | -0.146 | 30.002 | 20.056 | 0.478 | 47.8
[6 empty] | -0.146 | 29.992 | 20.056 | 0.478 | 47.8
[6 empty] | -0.146 | 29.982 | 20.056 | 0.479 | 47.9
[6 empty] | -0.146 | 29.972 | 20.056 | 0.479 | 47.9
[6 empty] | -0.146 | 29.962 | 20.056 | 0.48 | 48
[6 empty] | -0.146 | 29.952 | 20.056 | 0.481 | 48.1
[6 empty] | -0.146 | 29.942 | 20.056 | 0.481 | 48.1
[6 empty] | -0.146 | 29.932 | 20.056 | 0.482 | 48.2
[6 empty] | -0.146 | 29.922 | 20.056 | 0.483 | 48.3
[6 empty] | -0.146 | 29.912 | 20.056 | 0.483 | 48.3
[6 empty] | -0.146 | 29.902 | 20.056 | 0.484 | 48.4
[6 empty] | -0.146 | 29.892 | 20.056 | 0.485 | 48.5
[6 empty] | -0.146 | 29.882 | 20.056 | 0.485 | 48.5
[6 empty] | -0.146 | 29.872 | 20.056 | 0.486 | 48.6
[6 empty] | -0.146 | 29.862 | 20.056 | 0.486 | 48.6
[6 empty] | -0.146 | 29.852 | 20.056 | 0.487 | 48.7
[6 empty] | -0.146 | 29.842 | 20.056 | 0.488 | 48.8
[6 empty] | -0.146 | 29.832 | 20.056 | 0.488 | 48.8
[6 empty] | -0.146 | 29.822 | 20.056 | 0.489 | 48.9
[6 empty] | -0.146 | 29.812 | 20.056 | 0.49 | 49
[6 empty] | -0.146 | 29.802 | 20.056 | 0.49 | 49
[6 empty] | -0.146 | 29.792 | 20.056 | 0.491 | 49.1
[6 empty] | -0.146 | 29.782 | 20.056 | 0.491 | 49.1
[6 empty] | -0.146 | 29.772 | 20.056 | 0.491 | 49.1
[6 empty] | -0.146 | 29.762 | 20.056 | 0.492 | 49.2
[6 empty] | -0.146 | 29.752 | 20.056 | 0.492 | 49.2
[6 empty] | -0.146 | 29.742 | 20.056 | 0.492 | 49.2
[6 empty] | -0.146 | 29.732 | 20.056 | 0.493 | 49.3
[6 empty] | -0.146 | 29.722 | 20.056 | 0.493 | 49.3
[6 empty] | -0.146 | 29.712 | 20.056 | 0.493 | 49.3
[6 empty] | -0.146 | 29.702 | 20.056 | 0.494 | 49.4
[6 empty] | -0.146 | 29.692 | 20.056 | 0.494 | 49.4
[6 empty] | -0.146 | 29.682 | 20.056 | 0.494 | 49.4
[6 empty] | -0.146 | 29.672 | 20.056 | 0.495 | 49.5
[6 empty] | -0.146 | 29.662 | 20.056 | 0.495 | 49.5
[6 empty] | -0.146 | 29.652 | 20.055 | 0.495 | 49.5
[6 empty] | -0.146 | 29.642 | 20.055 | 0.496 | 49.6
[6 empty] | -0.146 | 29.632 | 20.055 | 0.496 | 49.6
[6 empty] | -0.146 | 29.622 | 20.055 | 0.496 | 49.6
[6 empty] | -0.146 | 29.612 | 20.055 | 0.497 | 49.7
[6 empty] | -0.146 | 29.602 | 20.055 | 0.497 | 49.7
[6 empty] | -0.146 | 29.592 | 20.055 | 0.497 | 49.7
[6 empty] | -0.146 | 29.582 | 20.055 | 0.498 | 49.8
[6 empty] | -0.146 | 29.572 | 20.055 | 0.498 | 49.8
[6 empty] | -0.146 | 29.562 | 20.055 | 0.498 | 49.8
[6 empty] | -0.146 | 29.552 | 20.055 | 0.498 | 49.8
[6 empty] | -0.146 | 29.542 | 20.055 | 0.499 | 49.9
[6 empty] | -0.146 | 29.532 | 20.055 | 0.499 | 49.9
[6 empty] | -0.146 | 29.522 | 20.055 | 0.499 | 49.9
[6 empty] | -0.146 | 29.512 | 20.055 | 0.499 | 49.9
[6 empty] | -0.146 | 29.502 | 20.055 | 0.5 | 50
[6 empty] | -0.146 | 29.492 | 20.055 | 0.5 | 50
[6 empty] | -0.146 | 29.482 | 20.055 | 0.5 | 50
[6 empty] | -0.146 | 29.472 | 20.055 | 0.5 | 50
[6 empty] | -0.146 | 29.462 | 20.055 | 0.501 | 50.1
[6 empty] | -0.146 | 29.452 | 20.055 | 0.501 | 50.1
[6 empty] | -0.146 | 29.442 | 20.055 | 0.501 | 50.1
[6 empty] | -0.146 | 29.432 | 20.055 | 0.501 | 50.1
[6 empty] | -0.146 | 29.422 | 20.055 | 0.502 | 50.2
[6 empty] | -0.146 | 29.412 | 20.055 | 0.502 | 50.2
[6 empty] | -0.146 | 29.402 | 20.055 | 0.502 | 50.2
[6 empty] | -0.146 | 29.392 | 20.055 | 0.502 | 50.2
[6 empty] | -0.146 | 29.382 | 20.055 | 0.503 | 50.3
[6 empty] | -0.146 | 29.372 | 20.055 | 0.503 | 50.3
[6 empty] | -0.146 | 29.362 | 20.055 | 0.503 | 50.3
[6 empty] | -0.146 | 29.352 | 20.055 | 0.503 | 50.3
[6 empty] | -0.146 | 29.342 | 20.055 | 0.503 | 50.3
[6 empty] | -0.146 | 29.332 | 20.055 | 0.504 | 50.4
[6 empty] | -0.146 | 29.322 | 20.055 | 0.504 | 50.4
[6 empty] | -0.146 | 29.312 | 20.055 | 0.504 | 50.4
[6 empty] | -0.146 | 29.302 | 20.055 | 0.504 | 50.4
[6 empty] | -0.146 | 29.292 | 20.055 | 0.504 | 50.4
[6 empty] | -0.146 | 29.282 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.272 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.262 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.252 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.242 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.232 | 20.055 | 0.505 | 50.5
[6 empty] | -0.146 | 29.222 | 20.055 | 0.506 | 50.6
[6 empty] | -0.146 | 29.212 | 20.055 | 0.506 | 50.6
[6 empty] | -0.146 | 29.202 | 20.055 | 0.506 | 50.6
[6 empty] | -0.146 | 29.192 | 20.055 | 0.506 | 50.6
[6 empty] | -0.146 | 29.182 | 20.055 | 0.506 | 50.6
[6 empty] | -0.146 | 29.172 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.162 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.152 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.142 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.132 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.122 | 20.055 | 0.507 | 50.7
[6 empty] | -0.146 | 29.112 | 20.055 | 0.508 | 50.8
[6 empty] | -0.146 | 29.102 | 20.055 | 0.508 | 50.8
[6 empty] | -0.146 | 29.092 | 20.055 | 0.508 | 50.8
[6 empty] | -0.146 | 29.082 | 20.055 | 0.508 | 50.8
[6 empty] | -0.146 | 29.072 | 20.055 | 0.508 | 50.8
[6 empty] | -0.146 | 29.062 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.052 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.042 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.032 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.022 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.012 | 20.055 | 0.509 | 50.9
[6 empty] | -0.146 | 29.002 | 20.055 | 0.51 | 51
[6 empty] | -0.146 | 28.992 | 20.055 | 0.51 | 51
[6 empty] | -0.146 | 28.982 | 20.055 | 0.51 | 51
[6 empty] | -0.146 | 28.972 | 20.055 | 0.51 | 51
[6 empty] | -0.146 | 28.962 | 20.054 | 0.51 | 51
[6 empty] | -0.146 | 28.952 | 20.054 | 0.51 | 51
[6 empty] | -0.146 | 28.942 | 20.054 | 0.51 | 51
[6 empty] | -0.146 | 28.932 | 20.054 | 0.51 | 51
[6 empty] | -0.146 | 28.922 | 20.054 | 0.51 | 51
[6 empty] | -0.146 | 28.912 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.902 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.892 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.882 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.872 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.862 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.852 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.842 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.832 | 20.054 | 0.511 | 51.1
[6 empty] | -0.146 | 28.822 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.812 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.802 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.792 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.782 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.772 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.762 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.752 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.742 | 20.054 | 0.512 | 51.2
[6 empty] | -0.146 | 28.732 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.722 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.712 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.702 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.692 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.682 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.672 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.662 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.652 | 20.054 | 0.513 | 51.3
[6 empty] | -0.146 | 28.642 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.632 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.622 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.612 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.602 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.592 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.582 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.572 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.562 | 20.054 | 0.514 | 51.4
[6 empty] | -0.146 | 28.552 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.542 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.532 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.522 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.512 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.502 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.492 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.482 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.472 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.462 | 20.054 | 0.515 | 51.5
[6 empty] | -0.146 | 28.452 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.442 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.432 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.422 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.412 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.402 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.392 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.382 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.372 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.362 | 20.054 | 0.516 | 51.6
[6 empty] | -0.146 | 28.352 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.342 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.332 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.322 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.312 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.302 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.292 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.282 | 20.054 | 0.517 | 51.7
[6 empty] | -0.146 | 28.272 | 20.053 | 0.517 | 51.7
[6 empty] | -0.146 | 28.262 | 20.053 | 0.517 | 51.7
[6 empty] | -0.146 | 28.252 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.242 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.232 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.222 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.212 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.202 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.192 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.182 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.172 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.162 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.152 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.142 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.132 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.122 | 20.053 | 0.518 | 51.8
[6 empty] | -0.146 | 28.112 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.102 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.092 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.082 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.072 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.062 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.052 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.042 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.032 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.022 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.012 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 28.002 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.992 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.982 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.973 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.963 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.953 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.943 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.933 | 20.053 | 0.519 | 51.9
[6 empty] | -0.146 | 27.923 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.913 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.903 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.893 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.883 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.873 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.863 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.853 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.843 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.833 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.823 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.813 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.803 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.793 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.783 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.773 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.763 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.753 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.743 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.733 | 20.053 | 0.52 | 52
[6 empty] | -0.146 | 27.723 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.713 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.703 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.693 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.683 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.673 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.663 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.653 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.643 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.633 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.623 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.613 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.603 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.593 | 20.053 | 0.521 | 52.1
[6 empty] | -0.146 | 27.583 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.573 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.563 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.553 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.543 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.533 | 20.052 | 0.521 | 52.1
[6 empty] | -0.146 | 27.523 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.513 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.503 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.493 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.483 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.473 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.463 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.453 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.443 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.433 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.423 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.413 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.403 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.393 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.383 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.373 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.363 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.353 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.343 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.333 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.323 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.313 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.303 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.293 | 20.052 | 0.522 | 52.2
[6 empty] | -0.146 | 27.283 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.273 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.263 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.253 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.243 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.233 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.223 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.213 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.203 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.193 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.183 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.173 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.163 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.153 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.143 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.133 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.123 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.113 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.103 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.093 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.083 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.073 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.063 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.053 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.043 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.033 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.023 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.013 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 27.003 | 20.052 | 0.523 | 52.3
[6 empty] | -0.146 | 26.993 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.983 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.973 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.963 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.953 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.943 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.933 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.923 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.913 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.903 | 20.052 | 0.524 | 52.4
[6 empty] | -0.146 | 26.893 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.883 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.873 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.863 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.853 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.843 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.833 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.823 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.813 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.803 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.793 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.783 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.773 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.763 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.753 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.743 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.733 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.723 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.713 | 20.051 | 0.524 | 52.4
[6 empty] | -0.146 | 26.703 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.693 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.683 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.673 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.663 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.653 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.643 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.633 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.623 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.613 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.603 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.593 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.583 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.573 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.563 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.553 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.543 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.533 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.523 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.513 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.503 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.493 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.483 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.473 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.463 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.453 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.443 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.433 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.423 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.413 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.403 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.393 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.383 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.373 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.363 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.353 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.343 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.333 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.323 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.313 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.303 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.293 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.283 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.273 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.263 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.253 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.243 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.233 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.223 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.213 | 20.051 | 0.525 | 52.5
[6 empty] | -0.146 | 26.203 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.193 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.183 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.173 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.163 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.153 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.143 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.133 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.123 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.113 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.103 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.093 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.083 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.073 | 20.05 | 0.525 | 52.5
[6 empty] | -0.146 | 26.063 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.053 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.043 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.033 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.023 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.013 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 26.003 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.993 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.983 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.973 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.963 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.953 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.943 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.933 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.923 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.913 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.903 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.893 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.883 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.873 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.863 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.853 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.843 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.833 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.823 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.813 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.803 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.793 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.783 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.773 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.763 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.753 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.743 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.733 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.723 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.713 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.703 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.693 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.683 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.673 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.663 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.653 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.643 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.633 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.623 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.613 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.603 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.593 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.583 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.573 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.563 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.553 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.543 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.533 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.523 | 20.05 | 0.526 | 52.6
[6 empty] | -0.146 | 25.513 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.503 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.493 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.483 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.473 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.463 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.453 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.443 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.433 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.423 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.413 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.403 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.393 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.383 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.373 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.363 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.353 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.343 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.333 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.323 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.313 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.303 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.293 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.283 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.273 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.263 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.253 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.243 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.233 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.223 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.213 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.203 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.193 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.183 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.173 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.163 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.153 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.143 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.133 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.123 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.113 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.103 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.093 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.083 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.073 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.063 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.053 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.043 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.033 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.023 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.013 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 25.003 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.993 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.983 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.974 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.964 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.954 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.944 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.934 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.924 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.914 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.904 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.894 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.884 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.874 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.864 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.854 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.844 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.834 | 20.049 | 0.526 | 52.6
[6 empty] | -0.146 | 24.824 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.814 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.804 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.794 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.784 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.774 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.764 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.754 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.744 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.734 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.724 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.714 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.704 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.694 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.684 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.674 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.664 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.654 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.644 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.634 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.624 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.614 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.604 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.594 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.584 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.574 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.564 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.554 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.544 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.534 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.524 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.514 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.504 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.494 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.484 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.474 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.464 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.454 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.444 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.434 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.424 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.414 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.404 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.394 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.384 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.374 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.364 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.354 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.344 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.334 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.324 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.314 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.304 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.294 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.284 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.274 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.264 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.254 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.244 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.234 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.224 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.214 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.204 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.194 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.184 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.174 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.164 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.154 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.144 | 20.048 | 0.526 | 52.6
[6 empty] | -0.146 | 24.134 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.124 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.114 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.104 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.094 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.084 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.074 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.064 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.054 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.044 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.034 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.024 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.014 | 20.047 | 0.526 | 52.6
[6 empty] | -0.146 | 24.004 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.994 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.984 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.974 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.964 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.954 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.944 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.934 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.924 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.914 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.904 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.894 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.884 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.874 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.864 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.854 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.844 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.834 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.824 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.814 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.804 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.794 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.784 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.774 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.764 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.754 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.744 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.734 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.724 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.714 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.704 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.694 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.684 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.674 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.664 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.654 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.644 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.634 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.624 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.614 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.604 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.594 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.584 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.574 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.564 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.554 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.544 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.534 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.524 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.514 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.504 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.494 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.484 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.474 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.464 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.454 | 20.047 | 0.525 | 52.5
[6 empty] | -0.146 | 23.444 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.434 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.424 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.414 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.404 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.394 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.384 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.374 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.364 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.354 | 20.046 | 0.525 | 52.5
[6 empty] | -0.146 | 23.344 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.334 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.324 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.314 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.304 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.294 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.284 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.274 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.264 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.254 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.244 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.234 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.224 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.214 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.204 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.194 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.184 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.174 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.164 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.154 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.144 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.134 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.124 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.114 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.104 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.094 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.084 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.074 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.064 | 20.046 | 0.524 | 52.4
[6 empty] | -0.146 | 23.054 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 23.044 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 23.034 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 23.024 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 23.014 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 23.004 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.994 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.984 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.974 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.964 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.954 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.944 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.934 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.924 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.914 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.904 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.894 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.884 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.874 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.864 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.854 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.844 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.834 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.824 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.814 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.804 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.794 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.784 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.774 | 20.046 | 0.523 | 52.3
[6 empty] | -0.146 | 22.764 | 20.046 | 0.522 | 52.2
[6 empty] | -0.146 | 22.754 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.744 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.734 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.724 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.714 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.704 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.694 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.684 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.674 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.664 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.654 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.644 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.634 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.624 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.614 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.604 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.594 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.584 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.574 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.564 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.554 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.544 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.534 | 20.045 | 0.522 | 52.2
[6 empty] | -0.146 | 22.524 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.514 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.504 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.494 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.484 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.474 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.464 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.454 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.444 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.434 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.424 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.414 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.404 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.394 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.384 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.374 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.364 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.354 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.344 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.334 | 20.045 | 0.521 | 52.1
[6 empty] | -0.146 | 22.324 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.314 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.304 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.294 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.284 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.274 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.264 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.254 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.244 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.234 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.224 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.214 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.204 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.194 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.184 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.174 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.164 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.154 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.144 | 20.045 | 0.52 | 52
[6 empty] | -0.146 | 22.134 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.124 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.114 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.104 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.094 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.084 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.074 | 20.045 | 0.519 | 51.9
[6 empty] | -0.146 | 22.064 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.054 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.044 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.034 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.024 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.014 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 22.004 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.994 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.984 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.975 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.965 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.955 | 20.044 | 0.519 | 51.9
[6 empty] | -0.146 | 21.945 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.935 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.925 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.915 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.905 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.895 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.885 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.875 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.865 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.855 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.845 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.835 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.825 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.815 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.805 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.795 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.785 | 20.044 | 0.518 | 51.8
[6 empty] | -0.146 | 21.775 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.765 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.755 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.745 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.735 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.725 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.715 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.705 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.695 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.685 | 20.044 | 0.517 | 51.7
[6 empty] | -0.146 | 21.675 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.665 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.655 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.645 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.635 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.625 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.615 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.605 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.595 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.585 | 20.044 | 0.516 | 51.6
[6 empty] | -0.146 | 21.575 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.565 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.555 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.545 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.535 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.525 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.515 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.505 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.495 | 20.044 | 0.515 | 51.5
[6 empty] | -0.146 | 21.485 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.475 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.465 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.455 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.445 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.435 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.425 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.415 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.405 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.395 | 20.044 | 0.514 | 51.4
[6 empty] | -0.146 | 21.385 | 20.044 | 0.513 | 51.3
[6 empty] | -0.146 | 21.375 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.365 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.355 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.345 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.335 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.325 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.315 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.305 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.295 | 20.043 | 0.513 | 51.3
[6 empty] | -0.146 | 21.285 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.275 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.265 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.255 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.245 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.235 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.225 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.215 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.205 | 20.043 | 0.512 | 51.2
[6 empty] | -0.146 | 21.195 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.185 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.175 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.165 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.155 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.145 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.135 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.125 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.115 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.105 | 20.043 | 0.511 | 51.1
[6 empty] | -0.146 | 21.095 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.085 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.075 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.065 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.055 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.045 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.035 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.025 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.015 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 21.005 | 20.043 | 0.51 | 51
[6 empty] | -0.146 | 20.995 | 20.043 | 0.509 | 50.9
[6 empty] | -0.146 | 20.985 | 20.043 | 0.509 | 50.9
[6 empty] | -0.146 | 20.975 | 20.043 | 0.509 | 50.9
[6 empty] | -0.146 | 20.965 | 20.043 | 0.509 | 50.9
[6 empty] | -0.146 | 20.955 | 20.043 | 0.509 | 50.9
[6 empty] | -0.146 | 20.945 | 20.043 | 0.508 | 50.8
[6 empty] | -0.146 | 20.935 | 20.043 | 0.508 | 50.8
[6 empty] | -0.146 | 20.925 | 20.043 | 0.508 | 50.8
[6 empty] | -0.146 | 20.915 | 20.043 | 0.508 | 50.8
[6 empty] | -0.146 | 20.905 | 20.043 | 0.508 | 50.8
[6 empty] | -0.146 | 20.895 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.885 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.875 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.865 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.855 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.845 | 20.043 | 0.507 | 50.7
[6 empty] | -0.146 | 20.835 | 20.043 | 0.506 | 50.6
[6 empty] | -0.146 | 20.825 | 20.043 | 0.506 | 50.6
[6 empty] | -0.146 | 20.815 | 20.043 | 0.506 | 50.6
[6 empty] | -0.146 | 20.805 | 20.043 | 0.506 | 50.6
[6 empty] | -0.146 | 20.795 | 20.043 | 0.506 | 50.6
[6 empty] | -0.146 | 20.785 | 20.043 | 0.505 | 50.5
[6 empty] | -0.146 | 20.775 | 20.043 | 0.505 | 50.5
[6 empty] | -0.146 | 20.765 | 20.043 | 0.505 | 50.5
[6 empty] | -0.146 | 20.755 | 20.043 | 0.505 | 50.5
[6 empty] | -0.146 | 20.745 | 20.043 | 0.505 | 50.5
[6 empty] | -0.146 | 20.735 | 20.043 | 0.504 | 50.4
[6 empty] | -0.146 | 20.725 | 20.043 | 0.504 | 50.4
[6 empty] | -0.146 | 20.715 | 20.043 | 0.504 | 50.4
[6 empty] | -0.146 | 20.705 | 20.043 | 0.504 | 50.4
[6 empty] | -0.146 | 20.695 | 20.043 | 0.504 | 50.4
[6 empty] | -0.146 | 20.685 | 20.042 | 0.503 | 50.3
[6 empty] | -0.146 | 20.675 | 20.042 | 0.503 | 50.3
[6 empty] | -0.146 | 20.665 | 20.042 | 0.503 | 50.3
[6 empty] | -0.146 | 20.655 | 20.042 | 0.503 | 50.3
[6 empty] | -0.146 | 20.645 | 20.042 | 0.503 | 50.3
[6 empty] | -0.146 | 20.635 | 20.042 | 0.502 | 50.2
[6 empty] | -0.146 | 20.625 | 20.042 | 0.502 | 50.2
[6 empty] | -0.146 | 20.615 | 20.042 | 0.502 | 50.2
[6 empty] | -0.146 | 20.605 | 20.042 | 0.502 | 50.2
[6 empty] | -0.146 | 20.595 | 20.042 | 0.502 | 50.2
[6 empty] | -0.146 | 20.585 | 20.042 | 0.501 | 50.1
[6 empty] | -0.146 | 20.575 | 20.042 | 0.501 | 50.1
[6 empty] | -0.146 | 20.565 | 20.042 | 0.501 | 50.1
[6 empty] | -0.146 | 20.555 | 20.042 | 0.501 | 50.1
[6 empty] | -0.146 | 20.545 | 20.042 | 0.5 | 50
[6 empty] | -0.146 | 20.535 | 20.042 | 0.5 | 50
[6 empty] | -0.146 | 20.525 | 20.042 | 0.5 | 50
[6 empty] | -0.146 | 20.515 | 20.042 | 0.499 | 49.9
[6 empty] | -0.146 | 20.505 | 20.042 | 0.499 | 49.9
[6 empty] | -0.146 | 20.495 | 20.042 | 0.499 | 49.9
[6 empty] | -0.146 | 20.485 | 20.042 | 0.499 | 49.9
[6 empty] | -0.146 | 20.475 | 20.042 | 0.498 | 49.8
[6 empty] | -0.146 | 20.465 | 20.042 | 0.498 | 49.8
[6 empty] | -0.146 | 20.455 | 20.042 | 0.498 | 49.8
[6 empty] | -0.146 | 20.445 | 20.042 | 0.498 | 49.8
[6 empty] | -0.146 | 20.435 | 20.042 | 0.497 | 49.7
[6 empty] | -0.146 | 20.425 | 20.042 | 0.497 | 49.7
[6 empty] | -0.146 | 20.415 | 20.042 | 0.497 | 49.7
[6 empty] | -0.146 | 20.405 | 20.042 | 0.497 | 49.7
[6 empty] | -0.146 | 20.395 | 20.042 | 0.496 | 49.6
[6 empty] | -0.146 | 20.385 | 20.042 | 0.496 | 49.6
[6 empty] | -0.146 | 20.375 | 20.042 | 0.496 | 49.6
[6 empty] | -0.146 | 20.365 | 20.042 | 0.495 | 49.5
[6 empty] | -0.146 | 20.355 | 20.042 | 0.495 | 49.5
[6 empty] | -0.146 | 20.345 | 20.042 | 0.495 | 49.5
[6 empty] | -0.146 | 20.335 | 20.042 | 0.494 | 49.4
[6 empty] | -0.146 | 20.325 | 20.042 | 0.494 | 49.4
[6 empty] | -0.146 | 20.315 | 20.042 | 0.494 | 49.4
[6 empty] | -0.146 | 20.305 | 20.042 | 0.493 | 49.3
[6 empty] | -0.146 | 20.295 | 20.042 | 0.493 | 49.3
[6 empty] | -0.146 | 20.285 | 20.042 | 0.492 | 49.2
[6 empty] | -0.146 | 20.275 | 20.042 | 0.492 | 49.2
[6 empty] | -0.146 | 20.265 | 20.042 | 0.492 | 49.2
[6 empty] | -0.146 | 20.255 | 20.042 | 0.491 | 49.1
[6 empty] | -0.146 | 20.245 | 20.042 | 0.491 | 49.1
[6 empty] | -0.146 | 20.235 | 20.042 | 0.491 | 49.1
[6 empty] | -0.146 | 20.225 | 20.042 | 0.49 | 49
[6 empty] | -0.146 | 20.215 | 20.042 | 0.49 | 49
[6 empty] | -0.146 | 20.205 | 20.042 | 0.49 | 49
[6 empty] | -0.146 | 20.195 | 20.042 | 0.489 | 48.9
[6 empty] | -0.146 | 20.185 | 20.042 | 0.488 | 48.8
[6 empty] | -0.146 | 20.175 | 20.042 | 0.488 | 48.8
[6 empty] | -0.146 | 20.165 | 20.042 | 0.487 | 48.7
[6 empty] | -0.146 | 20.155 | 20.042 | 0.487 | 48.7
[6 empty] | -0.146 | 20.145 | 20.042 | 0.486 | 48.6
[6 empty] | -0.146 | 20.135 | 20.042 | 0.485 | 48.5
[6 empty] | -0.146 | 20.125 | 20.042 | 0.485 | 48.5
[6 empty] | -0.146 | 20.115 | 20.042 | 0.484 | 48.4
[6 empty] | -0.146 | 20.105 | 20.042 | 0.483 | 48.3
[6 empty] | -0.146 | 20.095 | 20.042 | 0.483 | 48.3
[6 empty] | -0.146 | 20.085 | 20.042 | 0.482 | 48.2
[6 empty] | -0.146 | 20.075 | 20.042 | 0.482 | 48.2
[6 empty] | -0.146 | 20.065 | 20.042 | 0.481 | 48.1
[6 empty] | -0.146 | 20.055 | 20.042 | 0.48 | 48
[6 empty] | -0.146 | 20.045 | 20.042 | 0.48 | 48
[6 empty] | -0.146 | 20.035 | 20.042 | 0.479 | 47.9
[6 empty] | -0.146 | 20.025 | 20.042 | 0.478 | 47.8
[6 empty] | -0.146 | 20.015 | 20.042 | 0.478 | 47.8
[6 empty] | -0.146 | 20.005 | 20.042 | 0.477 | 47.7
[6 empty] | -0.146 | 19.995 | 20.041 | 0.476 | 47.6
[6 empty] | -0.146 | 19.985 | 20.041 | 0.474 | 47.4
[6 empty] | -0.146 | 19.975 | 20.041 | 0.472 | 47.2
[6 empty] | -0.146 | 19.965 | 20.041 | 0.471 | 47.1
[6 empty] | -0.146 | 19.955 | 20.041 | 0.469 | 46.9
[6 empty] | -0.146 | 19.945 | 20.041 | 0.468 | 46.8
[6 empty] | -0.146 | 19.935 | 20.041 | 0.466 | 46.6
[6 empty] | -0.146 | 19.925 | 20.041 | 0.465 | 46.5
[6 empty] | -0.146 | 19.915 | 20.041 | 0.463 | 46.3
[6 empty] | -0.146 | 19.905 | 20.041 | 0.462 | 46.2
[6 empty] | -0.146 | 19.895 | 20.041 | 0.46 | 46
[6 empty] | -0.146 | 19.885 | 20.041 | 0.459 | 45.9
[6 empty] | -0.146 | 19.875 | 20.041 | 0.457 | 45.7
[6 empty] | -0.146 | 19.865 | 20.041 | 0.456 | 45.6
[6 empty] | -0.146 | 19.855 | 20.041 | 0.454 | 45.4
[6 empty] | -0.146 | 19.845 | 20.041 | 0.453 | 45.3
[6 empty] | -0.146 | 19.835 | 20.041 | 0.451 | 45.1
[6 empty] | -0.146 | 19.825 | 20.041 | 0.45 | 45
[6 empty] | -0.146 | 19.815 | 20.041 | 0.448 | 44.8
[6 empty] | -0.146 | 19.805 | 20.041 | 0.446 | 44.6
[6 empty] | -0.146 | 19.795 | 20.041 | 0.441 | 44.1
[6 empty] | -0.146 | 19.785 | 20.041 | 0.436 | 43.6
[6 empty] | -0.146 | 19.775 | 20.041 | 0.431 | 43.1
[6 empty] | -0.146 | 19.765 | 20.041 | 0.426 | 42.6
[6 empty] | -0.146 | 19.755 | 20.041 | 0.421 | 42.1
[6 empty] | -0.146 | 19.745 | 20.041 | 0.415 | 41.5
[6 empty] | -0.146 | 19.735 | 20.041 | 0.41 | 41
[6 empty] | -0.146 | 19.725 | 20.041 | 0.405 | 40.5
[6 empty] | -0.146 | 19.715 | 20.041 | 0.4 | 40
[6 empty] | -0.146 | 19.705 | 20.041 | 0.395 | 39.5
[6 empty] | -0.146 | 19.695 | 20.041 | 0.39 | 39
[6 empty] | -0.146 | 19.685 | 20.041 | 0.384 | 38.4
[6 empty] | -0.146 | 19.675 | 20.041 | 0.379 | 37.9
[6 empty] | -0.146 | 19.665 | 20.041 | 0.374 | 37.4
[6 empty] | -0.146 | 19.655 | 20.041 | 0.369 | 36.9
[6 empty] | -0.146 | 19.645 | 20.041 | 0.364 | 36.4
[6 empty] | -0.146 | 19.635 | 20.041 | 0.359 | 35.9
[6 empty] | -0.146 | 19.625 | 20.041 | 0.353 | 35.3
[6 empty] | -0.146 | 19.615 | 20.041 | 0.348 | 34.8
[6 empty] | -0.146 | 19.605 | 20.041 | 0.343 | 34.3
[6 empty] | -0.146 | 19.595 | 20.041 | 0.333 | 33.3
[6 empty] | -0.146 | 19.585 | 20.041 | 0.324 | 32.4
[6 empty] | -0.146 | 19.575 | 20.041 | 0.314 | 31.4
[6 empty] | -0.146 | 19.565 | 20.041 | 0.305 | 30.5
[6 empty] | -0.146 | 19.555 | 20.041 | 0.296 | 29.6
[6 empty] | -0.146 | 19.545 | 20.041 | 0.286 | 28.6
[6 empty] | -0.146 | 19.535 | 20.041 | 0.277 | 27.7
[6 empty] | -0.146 | 19.525 | 20.041 | 0.267 | 26.7
[6 empty] | -0.146 | 19.515 | 20.041 | 0.258 | 25.8
[6 empty] | -0.146 | 19.505 | 20.041 | 0.248 | 24.8
[6 empty] | -0.146 | 19.495 | 20.041 | 0.239 | 23.9
[6 empty] | -0.146 | 19.485 | 20.041 | 0.229 | 22.9
[6 empty] | -0.146 | 19.475 | 20.041 | 0.22 | 22
[6 empty] | -0.146 | 19.465 | 20.041 | 0.21 | 21
[6 empty] | -0.146 | 19.455 | 20.041 | 0.201 | 20.1
[6 empty] | -0.146 | 19.445 | 20.041 | 0.191 | 19.1
[6 empty] | -0.146 | 19.435 | 20.041 | 0.182 | 18.2
[6 empty] | -0.146 | 19.425 | 20.041 | 0.172 | 17.2
[6 empty] | -0.146 | 19.415 | 20.041 | 0.163 | 16.3
[6 empty] | -0.146 | 19.405 | 20.041 | 0.154 | 15.4
[6 empty] | -0.146 | 19.395 | 20.041 | 0.151 | 15.1
[6 empty] | -0.146 | 19.385 | 20.041 | 0.148 | 14.8
[6 empty] | -0.146 | 19.375 | 20.041 | 0.145 | 14.5
[6 empty] | -0.146 | 19.365 | 20.041 | 0.142 | 14.2
[6 empty] | -0.146 | 19.355 | 20.041 | 0.139 | 13.9
[6 empty] | -0.146 | 19.345 | 20.041 | 0.136 | 13.6
[6 empty] | -0.146 | 19.335 | 20.041 | 0.133 | 13.3
[6 empty] | -0.146 | 19.325 | 20.041 | 0.13 | 13
[6 empty] | -0.146 | 19.315 | 20.041 | 0.127 | 12.7
[6 empty] | -0.146 | 19.305 | 20.04 | 0.124 | 12.4
[6 empty] | -0.146 | 19.295 | 20.04 | 0.121 | 12.1
[6 empty] | -0.146 | 19.285 | 20.04 | 0.118 | 11.8
[6 empty] | -0.146 | 19.275 | 20.04 | 0.115 | 11.5
[6 empty] | -0.146 | 19.265 | 20.04 | 0.112 | 11.2
[6 empty] | -0.146 | 19.255 | 20.04 | 0.109 | 10.9
[6 empty] | -0.146 | 19.245 | 20.04 | 0.106 | 10.6
[6 empty] | -0.146 | 19.235 | 20.04 | 0.103 | 10.3
[6 empty] | -0.146 | 19.225 | 20.04 | 0.1 | 10
[6 empty] | -0.146 | 19.215 | 20.04 | 0.097 | 9.7
[6 empty] | -0.146 | 19.205 | 20.04 | 0.094 | 9.4
[6 empty] | -0.146 | 19.195 | 20.04 | 0.093 | 9.3
[6 empty] | -0.146 | 19.185 | 20.04 | 0.092 | 9.2
[6 empty] | -0.146 | 19.175 | 20.04 | 0.091 | 9.1
[6 empty] | -0.146 | 19.165 | 20.04 | 0.09 | 9
[6 empty] | -0.146 | 19.155 | 20.04 | 0.09 | 9
[6 empty] | -0.146 | 19.145 | 20.04 | 0.089 | 8.9
[6 empty] | -0.146 | 19.135 | 20.04 | 0.088 | 8.8
[6 empty] | -0.146 | 19.125 | 20.04 | 0.087 | 8.7
[6 empty] | -0.146 | 19.115 | 20.04 | 0.086 | 8.6
[6 empty] | -0.146 | 19.105 | 20.04 | 0.085 | 8.5
[6 empty] | -0.146 | 19.095 | 20.04 | 0.084 | 8.4
[6 empty] | -0.146 | 19.085 | 20.04 | 0.083 | 8.3
[6 empty] | -0.146 | 19.075 | 20.04 | 0.082 | 8.2
[6 empty] | -0.146 | 19.065 | 20.04 | 0.081 | 8.1
[6 empty] | -0.146 | 19.055 | 20.04 | 0.081 | 8.1
[6 empty] | -0.146 | 19.045 | 20.04 | 0.08 | 8
[6 empty] | -0.146 | 19.035 | 20.04 | 0.079 | 7.9
[6 empty] | -0.146 | 19.025 | 20.04 | 0.078 | 7.8
[6 empty] | -0.146 | 19.015 | 20.04 | 0.077 | 7.7
[6 empty] | -0.146 | 19.005 | 20.04 | 0.076 | 7.6
[6 empty] | -0.146 | 18.995 | 20.04 | 0.076 | 7.6
[6 empty] | -0.146 | 18.985 | 20.04 | 0.075 | 7.5
[6 empty] | -0.146 | 18.976 | 20.04 | 0.075 | 7.5
[6 empty] | -0.146 | 18.966 | 20.04 | 0.075 | 7.5
[6 empty] | -0.146 | 18.956 | 20.04 | 0.074 | 7.4
[6 empty] | -0.146 | 18.946 | 20.04 | 0.074 | 7.4
[6 empty] | -0.146 | 18.936 | 20.04 | 0.073 | 7.3
[6 empty] | -0.146 | 18.926 | 20.04 | 0.073 | 7.3
[6 empty] | -0.146 | 18.916 | 20.04 | 0.073 | 7.3
[6 empty] | -0.146 | 18.906 | 20.04 | 0.072 | 7.2
[6 empty] | -0.146 | 18.896 | 20.04 | 0.072 | 7.2
[6 empty] | -0.146 | 18.886 | 20.04 | 0.072 | 7.2
[6 empty] | -0.146 | 18.876 | 20.04 | 0.071 | 7.1
[6 empty] | -0.146 | 18.866 | 20.04 | 0.071 | 7.1
[6 empty] | -0.146 | 18.856 | 20.04 | 0.07 | 7
[6 empty] | -0.146 | 18.846 | 20.04 | 0.07 | 7
[6 empty] | -0.146 | 18.836 | 20.04 | 0.07 | 7
[6 empty] | -0.146 | 18.826 | 20.04 | 0.069 | 6.9
[6 empty] | -0.146 | 18.816 | 20.04 | 0.069 | 6.9
[6 empty] | -0.146 | 18.806 | 20.04 | 0.068 | 6.8
[6 empty] | -0.146 | 18.796 | 20.04 | 0.068 | 6.8
[6 empty] | -0.146 | 18.786 | 20.04 | 0.068 | 6.8
[6 empty] | -0.146 | 18.776 | 20.04 | 0.067 | 6.7
[6 empty] | -0.146 | 18.766 | 20.04 | 0.067 | 6.7
[6 empty] | -0.146 | 18.756 | 20.04 | 0.067 | 6.7
[6 empty] | -0.146 | 18.746 | 20.04 | 0.066 | 6.6
[6 empty] | -0.146 | 18.736 | 20.04 | 0.066 | 6.6
[6 empty] | -0.146 | 18.726 | 20.04 | 0.065 | 6.5
[6 empty] | -0.146 | 18.716 | 20.04 | 0.065 | 6.5
[6 empty] | -0.146 | 18.706 | 20.04 | 0.065 | 6.5
[6 empty] | -0.146 | 18.696 | 20.04 | 0.064 | 6.4
[6 empty] | -0.146 | 18.686 | 20.04 | 0.064 | 6.4
[6 empty] | -0.146 | 18.676 | 20.04 | 0.063 | 6.3
[6 empty] | -0.146 | 18.666 | 20.04 | 0.063 | 6.3
[6 empty] | -0.146 | 18.656 | 20.04 | 0.063 | 6.3
[6 empty] | -0.146 | 18.646 | 20.04 | 0.062 | 6.2
[6 empty] | -0.146 | 18.636 | 20.04 | 0.062 | 6.2
[6 empty] | -0.146 | 18.626 | 20.04 | 0.062 | 6.2
[6 empty] | -0.146 | 18.616 | 20.039 | 0.061 | 6.1
[6 empty] | -0.146 | 18.606 | 20.039 | 0.061 | 6.1
[6 empty] | -0.146 | 18.596 | 20.039 | 0.061 | 6.1
[6 empty] | -0.146 | 18.586 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.576 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.566 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.556 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.546 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.536 | 20.039 | 0.06 | 6
[6 empty] | -0.146 | 18.526 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.516 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.506 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.496 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.486 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.476 | 20.039 | 0.059 | 5.9
[6 empty] | -0.146 | 18.466 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.456 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.446 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.436 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.426 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.416 | 20.039 | 0.058 | 5.8
[6 empty] | -0.146 | 18.406 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.396 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.386 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.376 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.366 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.356 | 20.039 | 0.057 | 5.7
[6 empty] | -0.146 | 18.346 | 20.039 | 0.056 | 5.6
[6 empty] | -0.146 | 18.336 | 20.039 | 0.056 | 5.6
[6 empty] | -0.146 | 18.326 | 20.039 | 0.056 | 5.6
[6 empty] | -0.146 | 18.316 | 20.039 | 0.056 | 5.6
[6 empty] | -0.146 | 18.306 | 20.039 | 0.056 | 5.6
[6 empty] | -0.146 | 18.296 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.286 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.276 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.266 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.256 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.246 | 20.039 | 0.055 | 5.5
[6 empty] | -0.146 | 18.236 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.226 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.216 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.206 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.196 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.186 | 20.039 | 0.054 | 5.4
[6 empty] | -0.146 | 18.176 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.166 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.156 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.146 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.136 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.126 | 20.039 | 0.053 | 5.3
[6 empty] | -0.146 | 18.116 | 20.039 | 0.052 | 5.2
[6 empty] | -0.146 | 18.106 | 20.039 | 0.052 | 5.2
[6 empty] | -0.146 | 18.096 | 20.039 | 0.052 | 5.2
[6 empty] | -0.146 | 18.086 | 20.039 | 0.052 | 5.2
[6 empty] | -0.146 | 18.076 | 20.039 | 0.052 | 5.2
[6 empty] | -0.146 | 18.066 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.056 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.046 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.036 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.026 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.016 | 20.039 | 0.051 | 5.1
[6 empty] | -0.146 | 18.006 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.996 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.986 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.976 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.966 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.956 | 20.039 | 0.05 | 5
[6 empty] | -0.146 | 17.946 | 20.039 | 0.049 | 4.9
[6 empty] | -0.146 | 17.936 | 20.039 | 0.049 | 4.9
[6 empty] | -0.146 | 17.926 | 20.038 | 0.049 | 4.9
[6 empty] | -0.146 | 17.916 | 20.038 | 0.049 | 4.9
[6 empty] | -0.146 | 17.906 | 20.038 | 0.049 | 4.9
[6 empty] | -0.146 | 17.896 | 20.038 | 0.049 | 4.9
[6 empty] | -0.146 | 17.886 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.876 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.866 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.856 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.846 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.836 | 20.038 | 0.048 | 4.8
[6 empty] | -0.146 | 17.826 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.816 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.806 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.796 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.786 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.776 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.766 | 20.038 | 0.047 | 4.7
[6 empty] | -0.146 | 17.756 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.746 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.736 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.726 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.716 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.706 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.696 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.686 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.676 | 20.038 | 0.046 | 4.6
[6 empty] | -0.146 | 17.666 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.656 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.646 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.636 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.626 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.616 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.606 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.596 | 20.038 | 0.045 | 4.5
[6 empty] | -0.146 | 17.586 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.576 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.566 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.556 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.546 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.536 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.526 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.516 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.506 | 20.038 | 0.044 | 4.4
[6 empty] | -0.146 | 17.496 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.486 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.476 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.466 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.456 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.446 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.436 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.426 | 20.038 | 0.043 | 4.3
[6 empty] | -0.146 | 17.416 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.406 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.396 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.386 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.376 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.366 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.356 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.346 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.336 | 20.038 | 0.042 | 4.2
[6 empty] | -0.146 | 17.326 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.316 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.306 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.296 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.286 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.276 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.266 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.256 | 20.038 | 0.041 | 4.1
[6 empty] | -0.146 | 17.246 | 20.038 | 0.04 | 4
[6 empty] | -0.146 | 17.236 | 20.038 | 0.04 | 4
[6 empty] | -0.146 | 17.226 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.216 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.206 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.196 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.186 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.176 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.166 | 20.037 | 0.04 | 4
[6 empty] | -0.146 | 17.156 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.146 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.136 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.126 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.116 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.106 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.096 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.086 | 20.037 | 0.039 | 3.9
[6 empty] | -0.146 | 17.076 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.066 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.056 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.046 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.036 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.026 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.016 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 17.006 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 16.996 | 20.037 | 0.038 | 3.8
[6 empty] | -0.146 | 16.986 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.976 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.966 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.956 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.946 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.936 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.926 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.916 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.906 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.896 | 20.037 | 0.037 | 3.7
[6 empty] | -0.146 | 16.886 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.876 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.866 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.856 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.846 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.836 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.826 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.816 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.806 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.796 | 20.037 | 0.036 | 3.6
[6 empty] | -0.146 | 16.786 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.776 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.766 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.756 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.746 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.736 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.726 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.716 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.706 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.696 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.686 | 20.037 | 0.035 | 3.5
[6 empty] | -0.146 | 16.676 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.666 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.656 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.646 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.636 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.626 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.616 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.606 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.596 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.586 | 20.037 | 0.034 | 3.4
[6 empty] | -0.146 | 16.576 | 20.037 | 0.033 | 3.3
[6 empty] | -0.146 | 16.566 | 20.037 | 0.033 | 3.3
[6 empty] | -0.146 | 16.556 | 20.037 | 0.033 | 3.3
[6 empty] | -0.146 | 16.546 | 20.037 | 0.033 | 3.3
[6 empty] | -0.146 | 16.536 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.526 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.516 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.506 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.496 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.486 | 20.036 | 0.033 | 3.3
[6 empty] | -0.146 | 16.476 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.466 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.456 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.446 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.436 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.426 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.416 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.406 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.396 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.386 | 20.036 | 0.032 | 3.2
[6 empty] | -0.146 | 16.376 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.366 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.356 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.346 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.336 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.326 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.316 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.306 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.296 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.286 | 20.036 | 0.031 | 3.1
[6 empty] | -0.146 | 16.276 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.266 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.256 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.246 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.236 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.226 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.216 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.206 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.196 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.186 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.176 | 20.036 | 0.03 | 3
[6 empty] | -0.146 | 16.166 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.156 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.146 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.136 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.126 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.116 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.106 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.096 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.086 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.076 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.066 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.056 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.046 | 20.036 | 0.029 | 2.9
[6 empty] | -0.146 | 16.036 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 16.026 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 16.016 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 16.006 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.996 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.986 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.977 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.967 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.957 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.947 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.937 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.927 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.917 | 20.036 | 0.028 | 2.8
[6 empty] | -0.146 | 15.907 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.897 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.887 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.877 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.867 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.857 | 20.036 | 0.027 | 2.7
[6 empty] | -0.146 | 15.847 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.837 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.827 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.817 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.807 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.797 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.787 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.777 | 20.035 | 0.027 | 2.7
[6 empty] | -0.146 | 15.767 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.757 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.747 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.737 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.727 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.717 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.707 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.697 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.687 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.677 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.667 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.657 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.647 | 20.035 | 0.026 | 2.6
[6 empty] | -0.146 | 15.637 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.627 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.617 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.607 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.597 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.587 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.577 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.567 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.557 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.547 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.537 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.527 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.517 | 20.035 | 0.025 | 2.5
[6 empty] | -0.146 | 15.507 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.497 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.487 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.477 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.467 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.457 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.447 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.437 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.427 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.417 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.407 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.397 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.387 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.377 | 20.035 | 0.024 | 2.4
[6 empty] | -0.146 | 15.367 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.357 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.347 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.337 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.327 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.317 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.307 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.297 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.287 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.277 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.267 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.257 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.247 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.237 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.227 | 20.035 | 0.023 | 2.3
[6 empty] | -0.146 | 15.217 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.207 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.197 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.187 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.177 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.167 | 20.035 | 0.022 | 2.2
[6 empty] | -0.146 | 15.157 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.147 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.137 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.127 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.117 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.107 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.097 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.087 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.077 | 20.034 | 0.022 | 2.2
[6 empty] | -0.146 | 15.067 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.057 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.047 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.037 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.027 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.017 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 15.007 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.997 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.987 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.977 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.967 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.957 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.947 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.937 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.927 | 20.034 | 0.021 | 2.1
[6 empty] | -0.146 | 14.917 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.907 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.897 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.887 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.877 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.867 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.857 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.847 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.837 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.827 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.817 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.807 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.797 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.787 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.777 | 20.034 | 0.02 | 2
[6 empty] | -0.146 | 14.767 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.757 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.747 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.737 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.727 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.717 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.707 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.697 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.687 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.677 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.667 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.657 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.647 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.637 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.627 | 20.034 | 0.019 | 1.9
[6 empty] | -0.146 | 14.617 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.607 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.597 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.587 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.577 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.567 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.557 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.547 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.537 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.527 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.517 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.507 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.497 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.487 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.477 | 20.034 | 0.018 | 1.8
[6 empty] | -0.146 | 14.467 | 20.033 | 0.018 | 1.8
[6 empty] | -0.146 | 14.457 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.447 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.437 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.427 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.417 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.407 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.397 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.387 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.377 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.367 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.357 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.347 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.337 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.327 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.317 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.307 | 20.033 | 0.017 | 1.7
[6 empty] | -0.146 | 14.297 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.287 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.277 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.267 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.257 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.247 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.237 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.227 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.217 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.207 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.197 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.187 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.177 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.167 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.157 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.147 | 20.033 | 0.016 | 1.6
[6 empty] | -0.146 | 14.137 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.127 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.117 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.107 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.097 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.087 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.077 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.067 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.057 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.047 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.037 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.027 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.017 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 14.007 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 13.997 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 13.987 | 20.033 | 0.015 | 1.5
[6 empty] | -0.146 | 13.977 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.967 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.957 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.947 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.937 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.927 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.917 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.907 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.897 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.887 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.877 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.867 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.857 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.847 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.837 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.827 | 20.033 | 0.014 | 1.4
[6 empty] | -0.146 | 13.817 | 20.033 | 0.013 | 1.3
[6 empty] | -0.146 | 13.807 | 20.033 | 0.013 | 1.3
[6 empty] | -0.146 | 13.797 | 20.033 | 0.013 | 1.3
[6 empty] | -0.146 | 13.787 | 20.033 | 0.013 | 1.3
[6 empty] | -0.146 | 13.777 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.767 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.757 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.747 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.737 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.727 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.717 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.707 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.697 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.687 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.677 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.667 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.657 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.647 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.637 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.627 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.617 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.607 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.597 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.587 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.577 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.567 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.557 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.547 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.537 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.527 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.517 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.507 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.497 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.487 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.477 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.467 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.457 | 20.032 | 0.013 | 1.3
[6 empty] | -0.146 | 13.447 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.437 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.427 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.417 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.407 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.397 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.387 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.377 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.367 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.357 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.347 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.337 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.327 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.317 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.307 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.297 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.287 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.277 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.267 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.257 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.247 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.237 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.227 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.217 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.207 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.197 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.187 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.177 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.167 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.157 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.147 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.137 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.127 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.117 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.107 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.097 | 20.032 | 0.012 | 1.2
[6 empty] | -0.146 | 13.087 | 20.031 | 0.012 | 1.2
[6 empty] | -0.146 | 13.077 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.067 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.057 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.047 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.037 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.027 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.017 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 13.007 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.997 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.987 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.978 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.968 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.958 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.948 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.938 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.928 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.918 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.908 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.898 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.888 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.878 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.868 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.858 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.848 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.838 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.828 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.818 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.808 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.798 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.788 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.778 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.768 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.758 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.748 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.738 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.728 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.718 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.708 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.698 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.688 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.678 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.668 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.658 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.648 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.638 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.628 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.618 | 20.031 | 0.011 | 1.1
[6 empty] | -0.146 | 12.608 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.598 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.588 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.578 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.568 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.558 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.548 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.538 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.528 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.518 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.508 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.498 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.488 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.478 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.468 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.458 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.448 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.438 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.428 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.418 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.408 | 20.031 | 0.01 | 1
[6 empty] | -0.146 | 12.398 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.388 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.378 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.368 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.358 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.348 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.338 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.328 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.318 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.308 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.298 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.288 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.278 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.268 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.258 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.248 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.238 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.228 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.218 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.208 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.198 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.188 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.178 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.168 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.158 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.148 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.138 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.128 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.118 | 20.03 | 0.01 | 1
[6 empty] | -0.146 | 12.108 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.098 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.088 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.078 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.068 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.058 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.048 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.038 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.028 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.018 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 12.008 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.998 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.988 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.978 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.968 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.958 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.948 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.938 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.928 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.918 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.908 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.898 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.888 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.878 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.868 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.858 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.848 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.838 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.828 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.818 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.808 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.798 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.788 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.778 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.768 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.758 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.748 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.738 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.728 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.718 | 20.03 | 0.009 | 0.9
[6 empty] | -0.146 | 11.708 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.698 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.688 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.678 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.668 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.658 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.648 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.638 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.628 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.618 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.608 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.598 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.588 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.578 | 20.029 | 0.009 | 0.9
[6 empty] | -0.146 | 11.568 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.558 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.548 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.538 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.528 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.518 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.508 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.498 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.488 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.478 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.468 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.458 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.448 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.438 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.428 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.418 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.408 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.398 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.388 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.378 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.368 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.358 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.348 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.338 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.328 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.318 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.308 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.298 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.288 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.278 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.268 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.258 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.248 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.238 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.228 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.218 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.208 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.198 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.188 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.178 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.168 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.158 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.148 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.138 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.128 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.118 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.108 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.098 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.088 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.078 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.068 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.058 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.048 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.038 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.028 | 20.029 | 0.008 | 0.8
[6 empty] | -0.146 | 11.018 | 20.028 | 0.008 | 0.8
[6 empty] | -0.146 | 11.008 | 20.028 | 0.008 | 0.8
[6 empty] | -0.146 | 10.998 | 20.028 | 0.008 | 0.8
[6 empty] | -0.146 | 10.988 | 20.028 | 0.008 | 0.8
[6 empty] | -0.146 | 10.978 | 20.028 | 0.008 | 0.8
[6 empty] | -0.146 | 10.968 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.958 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.948 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.938 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.928 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.918 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.908 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.898 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.888 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.878 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.868 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.858 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.848 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.838 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.828 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.818 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.808 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.798 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.788 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.778 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.768 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.758 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.748 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.738 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.728 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.718 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.708 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.698 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.688 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.678 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.668 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.658 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.648 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.638 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.628 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.618 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.608 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.598 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.588 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.578 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.568 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.558 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.548 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.538 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.528 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.518 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.508 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.498 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.488 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.478 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.468 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.458 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.448 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.438 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.428 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.418 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.408 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.398 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.388 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.378 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.368 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.358 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.348 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.338 | 20.028 | 0.007 | 0.7
[6 empty] | -0.146 | 10.328 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.318 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.308 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.298 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.288 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.278 | 20.027 | 0.007 | 0.7
[6 empty] | -0.146 | 10.268 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.258 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.248 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.238 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.228 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.218 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.208 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.198 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.188 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.178 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.168 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.158 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.148 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.138 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.128 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.118 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.108 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.098 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.088 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.078 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.068 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.058 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.048 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.038 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.028 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.018 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 10.008 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.998 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.988 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.978 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.969 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.959 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.949 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.939 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.929 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.919 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.909 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.899 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.889 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.879 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.869 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.859 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.849 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.839 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.829 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.819 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.809 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.799 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.789 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.779 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.769 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.759 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.749 | 20.027 | 0.006 | 0.6
[6 empty] | -0.146 | 9.739 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.729 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.719 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.709 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.699 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.689 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.679 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.669 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.659 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.649 | 20.027 | 0.005 | 0.5
[6 empty] | -0.146 | 9.639 | 20.026 | 0.005 | 0.5
[6 empty] | -0.146 | 9.629 | 20.026 | 0.005 | 0.5
[6 empty] | -0.146 | 9.619 | 20.026 | 0.005 | 0.5
[6 empty] | -0.146 | 9.609 | 20.026 | 0.005 | 0.5
[6 empty] | -0.146 | 9.599 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.589 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.579 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.569 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.559 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.549 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.539 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.529 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.519 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.509 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.499 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.489 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.479 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.469 | 20.026 | 0.004 | 0.4
[6 empty] | -0.146 | 9.459 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.449 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.439 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.429 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.419 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.409 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.399 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.389 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.379 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.369 | 20.026 | 0.003 | 0.3
[6 empty] | -0.146 | 9.359 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.349 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.339 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.329 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.319 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.309 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.299 | 20.026 | 0.002 | 0.2
[6 empty] | -0.146 | 9.289 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.279 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.269 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.259 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.249 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.239 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.229 | 20.026 | 0.001 | 0.1
[6 empty] | -0.146 | 9.219 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.209 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.199 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.189 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.179 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.169 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.159 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.149 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.139 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.129 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.119 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.109 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.099 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.089 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.079 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.069 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.059 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.049 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.039 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.029 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.019 | 20.026 | 0 | 0
[6 empty] | -0.146 | 9.009 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.999 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.989 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.979 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.969 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.959 | 20.026 | 0 | 0
[6 empty] | -0.146 | 8.949 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.939 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.929 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.919 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.909 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.899 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.889 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.879 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.869 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.859 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.849 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.839 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.829 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.819 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.809 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.799 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.789 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.779 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.769 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.759 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.749 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.739 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.729 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.719 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.709 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.699 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.689 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.679 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.669 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.659 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.649 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.639 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.629 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.619 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.609 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.599 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.589 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.579 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.569 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.559 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.549 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.539 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.529 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.519 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.509 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.499 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.489 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.479 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.469 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.459 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.449 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.439 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.429 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.419 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.409 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.399 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.389 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.379 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.369 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.359 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.349 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.339 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.329 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.319 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.309 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.299 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.289 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.279 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.269 | 20.025 | 0 | 0
[6 empty] | -0.146 | 8.259 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.249 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.239 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.229 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.219 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.209 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.199 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.189 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.179 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.169 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.159 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.149 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.139 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.129 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.119 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.109 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.099 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.089 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.079 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.069 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.059 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.049 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.039 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.029 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.019 | 20.024 | 0 | 0
[6 empty] | -0.146 | 8.009 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.999 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.989 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.979 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.969 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.959 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.949 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.939 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.929 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.919 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.909 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.899 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.889 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.879 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.869 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.859 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.849 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.839 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.829 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.819 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.809 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.799 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.789 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.779 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.769 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.759 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.749 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.739 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.729 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.719 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.709 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.699 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.689 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.679 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.669 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.659 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.649 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.639 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.629 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.619 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.609 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.599 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.589 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.579 | 20.024 | 0 | 0
[6 empty] | -0.146 | 7.569 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.559 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.549 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.539 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.529 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.519 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.509 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.499 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.489 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.479 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.469 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.459 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.449 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.439 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.429 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.419 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.409 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.399 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.389 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.379 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.369 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.359 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.349 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.339 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.329 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.319 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.309 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.299 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.289 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.279 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.269 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.259 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.249 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.239 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.229 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.219 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.209 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.199 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.189 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.179 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.169 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.159 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.149 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.139 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.129 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.119 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.109 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.099 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.089 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.079 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.069 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.059 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.049 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.039 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.029 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.019 | 20.023 | 0 | 0
[6 empty] | -0.146 | 7.009 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.999 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.989 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.979 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.97 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.96 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.95 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.94 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.93 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.92 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.91 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.9 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.89 | 20.023 | 0 | 0
[6 empty] | -0.146 | 6.88 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.87 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.86 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.85 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.84 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.83 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.82 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.81 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.8 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.79 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.78 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.77 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.76 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.75 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.74 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.73 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.72 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.71 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.7 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.69 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.68 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.67 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.66 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.65 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.64 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.63 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.62 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.61 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.6 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.59 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.58 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.57 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.56 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.55 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.54 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.53 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.52 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.51 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.5 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.49 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.48 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.47 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.46 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.45 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.44 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.43 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.42 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.41 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.4 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.39 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.38 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.37 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.36 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.35 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.34 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.33 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.32 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.31 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.3 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.29 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.28 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.27 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.26 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.25 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.24 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.23 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.22 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.21 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.2 | 20.022 | 0 | 0
[6 empty] | -0.146 | 6.19 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.18 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.17 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.16 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.15 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.14 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.13 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.12 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.11 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.1 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.09 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.08 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.07 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.06 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.05 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.04 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.03 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.02 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6.01 | 20.021 | 0 | 0
[6 empty] | -0.146 | 6 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.99 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.98 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.97 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.96 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.95 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.94 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.93 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.92 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.91 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.9 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.89 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.88 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.87 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.86 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.85 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.84 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.83 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.82 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.81 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.8 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.79 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.78 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.77 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.76 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.75 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.74 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.73 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.72 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.71 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.7 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.69 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.68 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.67 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.66 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.65 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.64 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.63 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.62 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.61 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.6 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.59 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.58 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.57 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.56 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.55 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.54 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.53 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.52 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.51 | 20.021 | 0 | 0
[6 empty] | -0.146 | 5.5 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.49 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.48 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.47 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.46 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.45 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.44 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.43 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.42 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.41 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.4 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.39 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.38 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.37 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.36 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.35 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.34 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.33 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.32 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.31 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.3 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.29 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.28 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.27 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.26 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.25 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.24 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.23 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.22 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.21 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.2 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.19 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.18 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.17 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.16 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.15 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.14 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.13 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.12 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.11 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.1 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.09 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.08 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.07 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.06 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.05 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.04 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.03 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.02 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5.01 | 20.02 | 0 | 0
[6 empty] | -0.146 | 5 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.99 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.98 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.97 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.96 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.95 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.94 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.93 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.92 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.91 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.9 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.89 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.88 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.87 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.86 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.85 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.84 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.83 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.82 | 20.02 | 0 | 0
[6 empty] | -0.146 | 4.81 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.8 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.79 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.78 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.77 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.76 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.75 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.74 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.73 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.72 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.71 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.7 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.69 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.68 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.67 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.66 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.65 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.64 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.63 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.62 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.61 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.6 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.59 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.58 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.57 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.56 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.55 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.54 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.53 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.52 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.51 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.5 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.49 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.48 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.47 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.46 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.45 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.44 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.43 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.42 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.41 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.4 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.39 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.38 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.37 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.36 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.35 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.34 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.33 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.32 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.31 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.3 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.29 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.28 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.27 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.26 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.25 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.24 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.23 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.22 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.21 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.2 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.19 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.18 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.17 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.16 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.15 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.14 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.13 | 20.019 | 0 | 0
[6 empty] | -0.146 | 4.12 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.11 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.1 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.09 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.08 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.07 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.06 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.05 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.04 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.03 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.02 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4.01 | 20.018 | 0 | 0
[6 empty] | -0.146 | 4 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.99 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.98 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.971 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.961 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.951 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.941 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.931 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.921 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.911 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.901 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.891 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.881 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.871 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.861 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.851 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.841 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.831 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.821 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.811 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.801 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.791 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.781 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.771 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.761 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.751 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.741 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.731 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.721 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.711 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.701 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.691 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.681 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.671 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.661 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.651 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.641 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.631 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.621 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.611 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.601 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.591 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.581 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.571 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.561 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.551 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.541 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.531 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.521 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.511 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.501 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.491 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.481 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.471 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.461 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.451 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.441 | 20.018 | 0 | 0
[6 empty] | -0.146 | 3.431 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.421 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.411 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.401 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.391 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.381 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.371 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.361 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.351 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.341 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.331 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.321 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.311 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.301 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.291 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.281 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.271 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.261 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.251 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.241 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.231 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.221 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.211 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.201 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.191 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.181 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.171 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.161 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.151 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.141 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.131 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.121 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.111 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.101 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.091 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.081 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.071 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.061 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.051 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.041 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.031 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.021 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.011 | 20.017 | 0 | 0
[6 empty] | -0.146 | 3.001 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.991 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.981 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.971 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.961 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.951 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.941 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.931 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.921 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.911 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.901 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.891 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.881 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.871 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.861 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.851 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.841 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.831 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.821 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.811 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.801 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.791 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.781 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.771 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.761 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.751 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.741 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.731 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.721 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.711 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.701 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.691 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.681 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.671 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.661 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.651 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.641 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.631 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.621 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.611 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.601 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.591 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.581 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.571 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.561 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.551 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.541 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.531 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.521 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.511 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.501 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.491 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.481 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.471 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.461 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.451 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.441 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.431 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.421 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.411 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.401 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.391 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.381 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.371 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.361 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.351 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.341 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.331 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.321 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.311 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.301 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.291 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.281 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.271 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.261 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.251 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.241 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.231 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.221 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.211 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.201 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.191 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.181 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.171 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.161 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.151 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.141 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.131 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.121 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.111 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.101 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.091 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.081 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.071 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.061 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.051 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.041 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.031 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.021 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.011 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.001 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.991 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.981 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.971 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.961 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.951 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.941 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.931 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.921 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.911 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.901 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.891 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.881 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.871 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.861 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.851 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.841 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.831 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.821 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.811 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.801 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.791 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.781 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.771 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.761 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.751 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.741 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.731 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.721 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.711 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.701 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.691 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.681 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.671 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.661 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.651 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.641 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.631 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.621 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.611 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.601 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.591 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.581 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.571 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.561 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.551 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.541 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.531 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.521 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.511 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.501 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.491 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.481 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.471 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.461 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.451 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.441 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.431 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.421 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.411 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.401 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.391 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.381 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.371 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.361 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.351 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.341 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.331 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.321 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.311 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.301 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.291 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.281 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.271 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.261 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.251 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.241 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.231 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.221 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.211 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.201 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.191 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.181 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.171 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.161 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.151 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.141 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.131 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.121 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.111 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.101 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.091 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.081 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.071 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.061 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.051 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.041 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.031 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.021 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.011 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.001 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.991 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.981 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.972 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.962 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.952 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.942 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.932 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.922 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.912 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.902 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.892 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.882 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.872 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.862 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.852 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.842 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.832 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.822 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.812 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.802 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.792 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.782 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.772 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.762 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.752 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.742 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.732 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.722 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.712 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.702 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.692 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.682 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.672 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.662 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.652 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.642 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.632 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.622 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.612 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.602 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.592 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.582 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.572 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.562 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.552 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.542 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.532 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.522 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.512 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.502 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.492 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.482 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.472 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.462 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.452 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.442 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.432 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.422 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.412 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.402 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.392 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.382 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.372 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.362 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.352 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.342 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.332 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.322 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.312 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.302 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.292 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.282 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.272 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.262 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.252 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.242 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.232 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.222 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.212 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.202 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.192 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.182 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.172 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.162 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.152 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.142 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.132 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.122 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.112 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.102 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.092 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.082 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.072 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.062 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.052 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.042 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.032 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.022 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.012 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.002 | 20.013 | 0 | 0
[6 empty] | -0.146 | -0.008 | 20.013 | 0 | 0
[6 empty] | -0.146 | -0.018 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.028 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.038 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.048 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.058 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.068 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.078 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.088 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.098 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.108 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.118 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.128 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.138 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.148 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.158 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.168 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.178 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.188 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.198 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.208 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.218 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.228 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.238 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.248 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.258 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.268 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.278 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.288 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.298 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.308 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.318 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.328 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.338 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.348 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.358 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.368 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.378 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.388 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.398 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.408 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.418 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.428 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.438 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.448 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.458 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.468 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.478 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.488 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.498 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.508 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.518 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.528 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.538 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.548 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.558 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.568 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.578 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.588 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.598 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.608 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.618 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.628 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.638 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.648 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.658 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.668 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.678 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.688 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.698 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.708 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.718 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.728 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.738 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.748 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.758 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.768 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.778 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.788 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.798 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.808 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.818 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.828 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.838 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.848 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.858 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.868 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.878 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.888 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.898 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.908 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.918 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.928 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.938 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.948 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.958 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.968 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.978 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.988 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.998 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.008 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.018 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.028 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.038 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.048 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.058 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.068 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.078 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.088 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.098 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.108 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.118 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.128 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.138 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.148 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.158 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.168 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.178 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.188 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.198 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.208 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.218 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.228 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.238 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.248 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.258 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.268 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.278 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.288 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.298 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.308 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.318 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.328 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.338 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.348 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.358 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.368 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.378 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.388 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.398 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.408 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.418 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.428 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.438 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.448 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.458 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.468 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.478 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.488 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.498 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.508 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.518 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.528 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.538 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.548 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.558 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.568 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.578 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.588 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.598 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.608 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.618 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.628 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.638 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.648 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.658 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.668 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.678 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.688 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.698 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.708 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.718 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.728 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.738 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.748 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.758 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.768 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.778 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.788 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.798 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.808 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.818 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.828 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.838 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.848 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.858 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.868 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.878 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.888 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.898 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.908 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.918 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.928 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.938 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.948 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.958 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.968 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.978 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.988 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.998 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.008 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.018 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.027 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.037 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.047 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.057 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.067 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.077 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.087 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.097 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.107 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.117 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.127 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.137 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.147 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.157 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.167 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.177 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.187 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.197 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.207 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.217 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.227 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.237 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.247 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.257 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.267 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.277 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.287 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.297 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.307 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.317 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.327 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.337 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.347 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.357 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.367 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.377 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.387 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.397 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.407 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.417 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.427 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.437 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.447 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.457 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.467 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.477 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.487 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.497 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.507 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.517 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.527 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.537 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.547 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.557 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.567 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.577 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.587 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.597 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.607 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.617 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.627 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.637 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.647 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.657 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.667 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.677 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.687 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.697 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.707 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.717 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.727 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.737 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.747 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.757 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.767 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.777 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.787 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.797 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.807 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.817 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.827 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.837 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.847 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.857 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.867 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.877 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.887 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.897 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.907 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.917 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.927 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.937 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.947 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.957 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.967 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.977 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.987 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.997 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.007 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.017 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.027 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.037 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.047 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.057 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.067 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.077 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.087 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.097 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.107 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.117 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.127 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.137 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.147 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.157 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.167 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.177 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.187 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.197 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.207 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.217 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.227 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.237 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.247 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.257 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.267 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.277 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.287 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.297 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.307 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.317 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.327 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.337 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.347 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.357 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.367 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.377 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.387 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.397 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.407 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.417 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.427 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.437 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.447 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.457 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.467 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.477 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.487 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.497 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.507 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.517 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.527 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.537 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.547 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.557 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.567 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.577 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.587 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.597 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.607 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.617 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.627 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.637 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.647 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.657 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.667 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.677 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.687 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.697 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.707 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.717 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.727 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.737 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.747 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.757 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.767 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.777 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.787 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.797 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.807 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.817 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.827 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.837 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.847 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.857 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.867 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.877 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.887 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.897 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.907 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.917 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.927 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.937 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.947 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.957 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.967 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.977 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.987 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.997 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.007 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.017 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.027 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.037 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.047 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.057 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.067 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.077 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.087 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.097 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.107 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.117 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.127 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.137 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.147 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.157 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.167 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.177 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.187 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.197 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.207 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.217 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.227 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.237 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.247 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.257 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.267 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.277 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.287 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.297 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.307 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.317 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.327 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.337 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.347 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.357 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.367 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.377 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.387 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.397 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.407 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.417 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.427 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.437 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.447 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.457 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.467 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.477 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.487 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.497 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.507 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.517 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.527 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.537 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.547 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.557 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.567 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.577 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.587 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.597 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.607 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.617 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.627 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.637 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.647 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.657 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.667 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.677 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.687 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.697 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.707 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.717 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.727 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.737 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.747 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.757 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.767 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.777 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.787 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.797 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.807 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.817 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.827 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.837 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.847 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.857 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.867 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.877 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.887 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.897 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.907 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.917 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.927 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.937 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.947 | 20.005 | 0 | 0
[6 empty] | ####################################################### |  |  |  | 0

## MLC
SCRIPT | [5 empty] | RAYSTATION
BEGIN_SCAN_DATA | [5 empty] | #######################################################
| FORMAT=CC-Export V1.9 |  |  |  |  | #RayStationVersion: | 13.1.0.144
| FILE_CREATION_DATE= |  |  |  |  | #PatientName: | RayTraining 40x40x40 Phantom
| LAST_MODIFIED=11-Jan-2023 11:56:56 |  |  |  |  | #PatientId: | Watertank cube
| BEGIN_SCAN 1 |  |  |  |  | #CoordinateSystem: | IEC patient coordinate system
|  | TASK_NAME=RayStation Phantom Profiles & PDDs |  |  |  | #######################################################
|  | PROGRAM=RayStation |  |  |  | #LineName: | 1
|  | COMMENT=RayStation Script Export |  |  |  | #DoseName: | Beam dose: MLC FIeld 1 (Validation, CT 1)
|  | MEAS_DATE=11-Jan-2023 11:56:56 |  |  |  | #DoseEngine: | Clinical: Collapsed Cone v5.7
|  | LINAC=RayStation |  |  |  | #TreatmentMachine: | RSL_TrueBeam Commission time: 11/20/2018 14:29:57
|  | MODALITY=X |  |  |  | #X [cm] | Y [cm] | Z [cm] | Dose [Gy] | cGy
|  | ISOCENTER=1000.00 |  |  |  | -0.146 | 55.043 | 20.092 | 0 | 0
|  | INPLANE_AXIS=Y |  |  |  | -0.146 | 55.033 | 20.092 | 0 | 0
|  | CROSSPLANE_AXIS=X |  |  |  | -0.146 | 55.023 | 20.092 | 0 | 0
|  | DEPTH_AXIS=Z |  |  |  | -0.146 | 55.013 | 20.092 | 0 | 0
|  | INPLANE_AXIS_DIR=GUN_TARGET |  |  |  | -0.146 | 55.003 | 20.092 | 0 | 0
|  | CROSSPLANE_AXIS_DIR=LEFT_RIGHT |  |  |  | -0.146 | 54.993 | 20.092 | 0 | 0
|  | DEPTH_AXIS_DIR=UP_DOWN |  |  |  | -0.146 | 54.983 | 20.092 | 0 | 0
|  | ENERGY=6.3 |  |  |  | -0.146 | 54.974 | 20.092 | 0 | 0
|  | NOMINAL_DMAX=20.00 |  |  |  | -0.146 | 54.964 | 20.092 | 0 | 0
|  | SSD=1000.0 |  |  |  | -0.146 | 54.954 | 20.092 | 0 | 0
|  | SCD=1000.00 |  |  |  | -0.146 | 54.944 | 20.092 | 0 | 0
|  | WEDGE_ANGLE=0.00 |  |  |  | -0.146 | 54.934 | 20.092 | 0 | 0
|  | FIELD_INPLANE=200.0 |  |  |  | -0.146 | 54.924 | 20.092 | 0 | 0
|  | FIELD_CROSSPLANE=200.0 |  |  |  | -0.146 | 54.914 | 20.092 | 0 | 0
|  | FIELD_TYPE=IRREGULAR |  |  |  | -0.146 | 54.904 | 20.092 | 0 | 0
|  | GANTRY=0.00 |  |  |  | -0.146 | 54.894 | 20.092 | 0 | 0
|  | GANTRY_UPRIGHT_POSITION=0 |  |  |  | -0.146 | 54.884 | 20.092 | 0 | 0
|  | GANTRY_ROTATION=CW |  |  |  | -0.146 | 54.874 | 20.092 | 0 | 0
|  | COLL_ANGLE=0.00 |  |  |  | -0.146 | 54.864 | 20.092 | 0 | 0
|  | REF_FIELD_DEPTH=50.00 |  |  |  | -0.146 | 54.854 | 20.092 | 0 | 0
|  | REF_FIELD_DEFINED=ISOCENTER |  |  |  | -0.146 | 54.844 | 20.092 | 0 | 0
|  | FIELD_INPLANE=200.0 |  |  |  | -0.146 | 54.834 | 20.092 | 0 | 0
|  | FIELD_CROSSPLANE=200.0 |  |  |  | -0.146 | 54.824 | 20.092 | 0 | 0
|  | SCAN_CURVETYPE=INPLANE_PROFILE |  |  |  | -0.146 | 54.814 | 20.092 | 0 | 0
|  | SCAN_DEPTH=100.0 |  |  |  | -0.146 | 54.804 | 20.092 | 0 | 0
|  | SCAN_OFFAXIS_INPLANE=0.00 |  |  |  | -0.146 | 54.794 | 20.092 | 0 | 0
|  | SCAN_OFFAXIS_CROSSPLANE=0.00 |  |  |  | -0.146 | 54.784 | 20.092 | 0 | 0
|  | SCAN_ANGLE=0.00 |  |  |  | -0.146 | 54.774 | 20.092 | 0 | 0
|  | SCAN_DIAGONAL=NOT_DIAGONAL |  |  |  | -0.146 | 54.764 | 20.092 | 0 | 0
|  | SCAN_DIRECTION=NEGATIVE |  |  |  | -0.146 | 54.754 | 20.092 | 0 | 0
|  | MEAS_MEDIUM=WATER |  |  |  | -0.146 | 54.744 | 20.092 | 0 | 0
|  | MEAS_UNIT=cGy |  |  |  | -0.146 | 54.734 | 20.092 | 0 | 0
|  | INCLINATION_ANGLE=0.00 |  |  |  | -0.146 | 54.724 | 20.092 | 0 | 0
|  | PRESSURE=1013.20 |  |  |  | -0.146 | 54.714 | 20.092 | 0 | 0
|  | TEMPERATURE=20.00 |  |  |  | -0.146 | 54.704 | 20.092 | 0 | 0
|  | NORM_TEMPERATURE=20.00 |  |  |  | -0.146 | 54.694 | 20.092 | 0 | 0
|  | CORRECTION_FACTOR=1.0000; |  |  |  | -0.146 | 54.684 | 20.092 | 0 | 0
|  | BEGIN_DATA |  |  |  | -0.146 | 54.674 | 20.092 | 0 | 0
|  |  | -5 | 0.0069873 | 1 | -0.146 | 54.664 | 20.092 | 0 | 0
|  |  | -4.95 | 0.0070314 | 1 | -0.146 | 54.654 | 20.092 | 0 | 0
|  |  | -4.9 | 0.0070755 | 1 | -0.146 | 54.644 | 20.092 | 0 | 0
|  |  | -4.85 | 0.0071196 | 1 | -0.146 | 54.634 | 20.092 | 0 | 0
|  |  | -4.8 | 0.0071637 | 1 | -0.146 | 54.624 | 20.092 | 0 | 0
|  |  | -4.75 | 0.0072828 | 1 | -0.146 | 54.614 | 20.092 | 0 | 0
|  |  | -4.7 | 0.007402 | 1 | -0.146 | 54.604 | 20.092 | 0 | 0
|  |  | -4.65 | 0.0075211 | 1 | -0.146 | 54.594 | 20.092 | 0 | 0
|  |  | -4.6 | 0.0076403 | 1 | -0.146 | 54.584 | 20.092 | 0 | 0
|  |  | -4.55 | 0.0077028 | 1 | -0.146 | 54.574 | 20.092 | 0 | 0
|  |  | -4.5 | 0.0077653 | 1 | -0.146 | 54.564 | 20.092 | 0 | 0
|  |  | -4.45 | 0.0078278 | 1 | -0.146 | 54.554 | 20.092 | 0 | 0
|  |  | -4.4 | 0.0078903 | 1 | -0.146 | 54.544 | 20.092 | 0 | 0
|  |  | -4.35 | 0.0079863 | 1 | -0.146 | 54.534 | 20.092 | 0 | 0
|  |  | -4.3 | 0.0080824 | 1 | -0.146 | 54.524 | 20.092 | 0 | 0
|  |  | -4.25 | 0.0081784 | 1 | -0.146 | 54.514 | 20.092 | 0 | 0
|  |  | -4.2 | 0.0082745 | 1 | -0.146 | 54.504 | 20.092 | 0 | 0
|  |  | -4.15 | 0.0083842 | 1 | -0.146 | 54.494 | 20.091 | 0 | 0
|  |  | -4.1 | 0.008494 | 1 | -0.146 | 54.484 | 20.091 | 0 | 0
|  |  | -4.05 | 0.0086037 | 1 | -0.146 | 54.474 | 20.091 | 0 | 0
|  |  | -4 | 0.0087135 | 1 | -0.146 | 54.464 | 20.091 | 0 | 0
|  |  | -3.95 | 0.0087773 | 1 | -0.146 | 54.454 | 20.091 | 0 | 0
|  |  | -3.9 | 0.0088412 | 1 | -0.146 | 54.444 | 20.091 | 0 | 0
|  |  | -3.85 | 0.0089051 | 1 | -0.146 | 54.434 | 20.091 | 0 | 0
|  |  | -3.8 | 0.008969 | 1 | -0.146 | 54.424 | 20.091 | 0 | 0
|  |  | -3.75 | 0.0091129 | 1 | -0.146 | 54.414 | 20.091 | 0 | 0
|  |  | -3.7 | 0.0092568 | 1 | -0.146 | 54.404 | 20.091 | 0 | 0
|  |  | -3.65 | 0.0094008 | 1 | -0.146 | 54.394 | 20.091 | 0 | 0
|  |  | -3.6 | 0.0095447 | 1 | -0.146 | 54.384 | 20.091 | 0 | 0
|  |  | -3.55 | 0.0096242 | 1 | -0.146 | 54.374 | 20.091 | 0 | 0
|  |  | -3.5 | 0.0097037 | 1 | -0.146 | 54.364 | 20.091 | 0 | 0
|  |  | -3.45 | 0.0097832 | 1 | -0.146 | 54.354 | 20.091 | 0 | 0
|  |  | -3.4 | 0.0098627 | 1 | -0.146 | 54.344 | 20.091 | 0 | 0
|  |  | -3.35 | 0.0099708 | 1 | -0.146 | 54.334 | 20.091 | 0 | 0
|  |  | -3.3 | 0.010079 | 1 | -0.146 | 54.324 | 20.091 | 0 | 0
|  |  | -3.25 | 0.010187 | 1 | -0.146 | 54.314 | 20.091 | 0 | 0
|  |  | -3.2 | 0.010295 | 1 | -0.146 | 54.304 | 20.091 | 0 | 0
|  |  | -3.15 | 0.010438 | 1 | -0.146 | 54.294 | 20.091 | 0 | 0
|  |  | -3.1 | 0.010581 | 1 | -0.146 | 54.284 | 20.091 | 0 | 0
|  |  | -3.05 | 0.010723 | 1 | -0.146 | 54.274 | 20.091 | 0 | 0
|  |  | -3 | 0.010866 | 1 | -0.146 | 54.264 | 20.091 | 0 | 0
|  |  | -2.95 | 0.01094 | 1 | -0.146 | 54.254 | 20.091 | 0 | 0
|  |  | -2.9 | 0.011014 | 1 | -0.146 | 54.244 | 20.091 | 0 | 0
|  |  | -2.85 | 0.011088 | 1 | -0.146 | 54.234 | 20.091 | 0 | 0
|  |  | -2.8 | 0.011163 | 1 | -0.146 | 54.224 | 20.091 | 0 | 0
|  |  | -2.75 | 0.011334 | 1 | -0.146 | 54.214 | 20.091 | 0 | 0
|  |  | -2.7 | 0.011505 | 1 | -0.146 | 54.204 | 20.091 | 0 | 0
|  |  | -2.65 | 0.011676 | 1 | -0.146 | 54.194 | 20.091 | 0 | 0
|  |  | -2.6 | 0.011847 | 1 | -0.146 | 54.184 | 20.091 | 0 | 0
|  |  | -2.55 | 0.011949 | 1 | -0.146 | 54.174 | 20.091 | 0 | 0
|  |  | -2.5 | 0.012051 | 1 | -0.146 | 54.164 | 20.091 | 0 | 0
|  |  | -2.45 | 0.012154 | 1 | -0.146 | 54.154 | 20.091 | 0 | 0
|  |  | -2.4 | 0.012256 | 1 | -0.146 | 54.144 | 20.091 | 0 | 0
|  |  | -2.35 | 0.012375 | 1 | -0.146 | 54.134 | 20.091 | 0 | 0
|  |  | -2.3 | 0.012494 | 1 | -0.146 | 54.124 | 20.091 | 0 | 0
|  |  | -2.25 | 0.012613 | 1 | -0.146 | 54.114 | 20.091 | 0 | 0
|  |  | -2.2 | 0.012733 | 1 | -0.146 | 54.104 | 20.091 | 0 | 0
|  |  | -2.15 | 0.012916 | 1 | -0.146 | 54.094 | 20.091 | 0 | 0
|  |  | -2.1 | 0.013099 | 1 | -0.146 | 54.084 | 20.091 | 0 | 0
|  |  | -2.05 | 0.013282 | 1 | -0.146 | 54.074 | 20.091 | 0 | 0
|  |  | -2 | 0.013465 | 1 | -0.146 | 54.064 | 20.091 | 0 | 0
|  |  | -1.95 | 0.013552 | 1 | -0.146 | 54.054 | 20.091 | 0 | 0
|  |  | -1.9 | 0.01364 | 1 | -0.146 | 54.044 | 20.091 | 0 | 0
|  |  | -1.85 | 0.013728 | 1 | -0.146 | 54.034 | 20.091 | 0 | 0
|  |  | -1.8 | 0.013816 | 1 | -0.146 | 54.024 | 20.091 | 0 | 0
|  |  | -1.75 | 0.014022 | 1 | -0.146 | 54.014 | 20.091 | 0 | 0
|  |  | -1.7 | 0.014229 | 1 | -0.146 | 54.004 | 20.091 | 0 | 0
|  |  | -1.65 | 0.014435 | 1 | -0.146 | 53.994 | 20.091 | 0 | 0
|  |  | -1.6 | 0.014642 | 1 | -0.146 | 53.984 | 20.091 | 0 | 0
|  |  | -1.55 | 0.014781 | 1 | -0.146 | 53.974 | 20.091 | 0 | 0
|  |  | -1.5 | 0.01492 | 1 | -0.146 | 53.964 | 20.091 | 0 | 0
|  |  | -1.45 | 0.015059 | 1 | -0.146 | 53.954 | 20.091 | 0 | 0
|  |  | -1.4 | 0.015198 | 1 | -0.146 | 53.944 | 20.091 | 0 | 0
|  |  | -1.35 | 0.015338 | 1 | -0.146 | 53.934 | 20.091 | 0 | 0
|  |  | -1.3 | 0.015478 | 1 | -0.146 | 53.924 | 20.091 | 0 | 0
|  |  | -1.25 | 0.015618 | 1 | -0.146 | 53.914 | 20.091 | 0 | 0
|  |  | -1.2 | 0.015757 | 1 | -0.146 | 53.904 | 20.091 | 0 | 0
|  |  | -1.15 | 0.016002 | 1 | -0.146 | 53.894 | 20.091 | 0 | 0
|  |  | -1.1 | 0.016247 | 1 | -0.146 | 53.884 | 20.091 | 0 | 0
|  |  | -1.05 | 0.016492 | 1 | -0.146 | 53.874 | 20.091 | 0 | 0
|  |  | -1 | 0.016736 | 1 | -0.146 | 53.864 | 20.091 | 0 | 0
|  |  | -0.95 | 0.01685 | 1 | -0.146 | 53.854 | 20.091 | 0 | 0
|  |  | -0.9 | 0.016964 | 1 | -0.146 | 53.844 | 20.091 | 0 | 0
|  |  | -0.85 | 0.017077 | 1 | -0.146 | 53.834 | 20.091 | 0 | 0
|  |  | -0.8 | 0.017191 | 1 | -0.146 | 53.824 | 20.091 | 0 | 0
|  |  | -0.75 | 0.017449 | 1 | -0.146 | 53.814 | 20.091 | 0 | 0
|  |  | -0.7 | 0.017707 | 1 | -0.146 | 53.804 | 20.09 | 0 | 0
|  |  | -0.65 | 0.017966 | 1 | -0.146 | 53.794 | 20.09 | 0 | 0
|  |  | -0.6 | 0.018224 | 1 | -0.146 | 53.784 | 20.09 | 0 | 0
|  |  | -0.55 | 0.018421 | 1 | -0.146 | 53.774 | 20.09 | 0 | 0
|  |  | -0.5 | 0.018618 | 1 | -0.146 | 53.764 | 20.09 | 0 | 0
|  |  | -0.45 | 0.018815 | 1 | -0.146 | 53.754 | 20.09 | 0 | 0
|  |  | -0.4 | 0.019012 | 1 | -0.146 | 53.744 | 20.09 | 0 | 0
|  |  | -0.35 | 0.019184 | 1 | -0.146 | 53.734 | 20.09 | 0 | 0
|  |  | -0.3 | 0.019355 | 1 | -0.146 | 53.724 | 20.09 | 0 | 0
|  |  | -0.25 | 0.019526 | 1 | -0.146 | 53.714 | 20.09 | 0 | 0
|  |  | -0.2 | 0.019698 | 1 | -0.146 | 53.704 | 20.09 | 0 | 0
|  |  | -0.15 | 0.020042 | 1 | -0.146 | 53.694 | 20.09 | 0 | 0
|  |  | -0.1 | 0.020386 | 1 | -0.146 | 53.684 | 20.09 | 0 | 0
|  |  | -0.05 | 0.02073 | 1 | -0.146 | 53.674 | 20.09 | 0 | 0
|  |  | 0 | 0.021074 | 1 | -0.146 | 53.664 | 20.09 | 0 | 0
|  |  | 0.05 | 0.021239 | 1 | -0.146 | 53.654 | 20.09 | 0 | 0
|  |  | 0.1 | 0.021404 | 1 | -0.146 | 53.644 | 20.09 | 0 | 0
|  |  | 0.15 | 0.021569 | 1 | -0.146 | 53.634 | 20.09 | 0 | 0
|  |  | 0.2 | 0.021734 | 1 | -0.146 | 53.624 | 20.09 | 0 | 0
|  |  | 0.25 | 0.022077 | 1 | -0.146 | 53.614 | 20.09 | 0 | 0
|  |  | 0.3 | 0.022421 | 1 | -0.146 | 53.604 | 20.09 | 0 | 0
|  |  | 0.35 | 0.022765 | 1 | -0.146 | 53.594 | 20.09 | 0 | 0
|  |  | 0.4 | 0.023109 | 1 | -0.146 | 53.584 | 20.09 | 0 | 0
|  |  | 0.45 | 0.023426 | 1 | -0.146 | 53.574 | 20.09 | 0 | 0
|  |  | 0.5 | 0.023744 | 1 | -0.146 | 53.564 | 20.09 | 0 | 0
|  |  | 0.55 | 0.024061 | 1 | -0.146 | 53.554 | 20.09 | 0 | 0
|  |  | 0.6 | 0.024378 | 1 | -0.146 | 53.544 | 20.09 | 0 | 0
|  |  | 0.65 | 0.024611 | 1 | -0.146 | 53.534 | 20.09 | 0 | 0
|  |  | 0.7 | 0.024844 | 1 | -0.146 | 53.524 | 20.09 | 0 | 0
|  |  | 0.75 | 0.025077 | 1 | -0.146 | 53.514 | 20.09 | 0 | 0
|  |  | 0.8 | 0.02531 | 1 | -0.146 | 53.504 | 20.09 | 0 | 0
|  |  | 0.85 | 0.025841 | 1 | -0.146 | 53.494 | 20.09 | 0 | 0
|  |  | 0.9 | 0.026373 | 1 | -0.146 | 53.484 | 20.09 | 0 | 0
|  |  | 0.95 | 0.026904 | 1 | -0.146 | 53.474 | 20.09 | 0 | 0
|  |  | 1 | 0.027436 | 1 | -0.146 | 53.464 | 20.09 | 0 | 0
|  |  | 1.05 | 0.027701 | 1 | -0.146 | 53.454 | 20.09 | 0 | 0
|  |  | 1.1 | 0.027966 | 1 | -0.146 | 53.444 | 20.09 | 0 | 0
|  |  | 1.15 | 0.028231 | 1 | -0.146 | 53.434 | 20.09 | 0 | 0
|  |  | 1.2 | 0.028496 | 1 | -0.146 | 53.424 | 20.09 | 0 | 0
|  |  | 1.25 | 0.029001 | 1 | -0.146 | 53.414 | 20.09 | 0 | 0
|  |  | 1.3 | 0.029505 | 1 | -0.146 | 53.404 | 20.09 | 0 | 0
|  |  | 1.35 | 0.03001 | 1 | -0.146 | 53.394 | 20.09 | 0 | 0
|  |  | 1.4 | 0.030514 | 1 | -0.146 | 53.384 | 20.09 | 0 | 0
|  |  | 1.45 | 0.031058 | 1 | -0.146 | 53.374 | 20.09 | 0 | 0
|  |  | 1.5 | 0.031603 | 1 | -0.146 | 53.364 | 20.09 | 0 | 0
|  |  | 1.55 | 0.032148 | 1 | -0.146 | 53.354 | 20.09 | 0 | 0
|  |  | 1.6 | 0.032692 | 1 | -0.146 | 53.344 | 20.09 | 0 | 0
|  |  | 1.65 | 0.033044 | 1 | -0.146 | 53.334 | 20.09 | 0 | 0
|  |  | 1.7 | 0.033396 | 1 | -0.146 | 53.324 | 20.09 | 0 | 0
|  |  | 1.75 | 0.033748 | 1 | -0.146 | 53.314 | 20.09 | 0 | 0
|  |  | 1.8 | 0.0341 | 1 | -0.146 | 53.304 | 20.09 | 0 | 0
|  |  | 1.85 | 0.034973 | 1 | -0.146 | 53.294 | 20.09 | 0 | 0
|  |  | 1.9 | 0.035846 | 1 | -0.146 | 53.284 | 20.09 | 0 | 0
|  |  | 1.95 | 0.036718 | 1 | -0.146 | 53.274 | 20.09 | 0 | 0
|  |  | 2 | 0.037591 | 1 | -0.146 | 53.264 | 20.09 | 0 | 0
|  |  | 2.05 | 0.038062 | 1 | -0.146 | 53.254 | 20.09 | 0 | 0
|  |  | 2.1 | 0.038534 | 1 | -0.146 | 53.244 | 20.09 | 0 | 0
|  |  | 2.15 | 0.039005 | 1 | -0.146 | 53.234 | 20.09 | 0 | 0
|  |  | 2.2 | 0.039476 | 1 | -0.146 | 53.224 | 20.09 | 0 | 0
|  |  | 2.25 | 0.04025 | 1 | -0.146 | 53.214 | 20.09 | 0 | 0
|  |  | 2.3 | 0.041023 | 1 | -0.146 | 53.204 | 20.09 | 0 | 0
|  |  | 2.35 | 0.041797 | 1 | -0.146 | 53.194 | 20.09 | 0 | 0
|  |  | 2.4 | 0.04257 | 1 | -0.146 | 53.184 | 20.09 | 0 | 0
|  |  | 2.45 | 0.043571 | 1 | -0.146 | 53.174 | 20.09 | 0 | 0
|  |  | 2.5 | 0.044572 | 1 | -0.146 | 53.164 | 20.09 | 0 | 0
|  |  | 2.55 | 0.045573 | 1 | -0.146 | 53.154 | 20.09 | 0 | 0
|  |  | 2.6 | 0.046573 | 1 | -0.146 | 53.144 | 20.09 | 0 | 0
|  |  | 2.65 | 0.047143 | 1 | -0.146 | 53.134 | 20.09 | 0 | 0
|  |  | 2.7 | 0.047712 | 1 | -0.146 | 53.124 | 20.09 | 0 | 0
|  |  | 2.75 | 0.048281 | 1 | -0.146 | 53.114 | 20.089 | 0 | 0
|  |  | 2.8 | 0.048851 | 1 | -0.146 | 53.104 | 20.089 | 0 | 0
|  |  | 2.85 | 0.050356 | 1 | -0.146 | 53.094 | 20.089 | 0 | 0
|  |  | 2.9 | 0.051862 | 1 | -0.146 | 53.084 | 20.089 | 0 | 0
|  |  | 2.95 | 0.053368 | 1 | -0.146 | 53.074 | 20.089 | 0 | 0
|  |  | 3 | 0.054873 | 1 | -0.146 | 53.064 | 20.089 | 0 | 0
|  |  | 3.05 | 0.055732 | 1 | -0.146 | 53.054 | 20.089 | 0 | 0
|  |  | 3.1 | 0.05659 | 1 | -0.146 | 53.044 | 20.089 | 0 | 0
|  |  | 3.15 | 0.057449 | 1 | -0.146 | 53.034 | 20.089 | 0 | 0
|  |  | 3.2 | 0.058307 | 1 | -0.146 | 53.024 | 20.089 | 0 | 0
|  |  | 3.25 | 0.059526 | 1 | -0.146 | 53.014 | 20.089 | 0 | 0
|  |  | 3.3 | 0.060746 | 1 | -0.146 | 53.004 | 20.089 | 0 | 0
|  |  | 3.35 | 0.061965 | 1 | -0.146 | 52.994 | 20.089 | 0 | 0
|  |  | 3.4 | 0.063184 | 1 | -0.146 | 52.984 | 20.089 | 0 | 0
|  |  | 3.45 | 0.064993 | 1 | -0.146 | 52.974 | 20.089 | 0 | 0
|  |  | 3.5 | 0.066802 | 1 | -0.146 | 52.964 | 20.089 | 0 | 0
|  |  | 3.55 | 0.068611 | 1 | -0.146 | 52.954 | 20.089 | 0 | 0
|  |  | 3.6 | 0.07042 | 1 | -0.146 | 52.944 | 20.089 | 0 | 0
|  |  | 3.65 | 0.071347 | 1 | -0.146 | 52.934 | 20.089 | 0 | 0
|  |  | 3.7 | 0.072275 | 1 | -0.146 | 52.924 | 20.089 | 0 | 0
|  |  | 3.75 | 0.073203 | 1 | -0.146 | 52.914 | 20.089 | 0 | 0
|  |  | 3.8 | 0.074131 | 1 | -0.146 | 52.904 | 20.089 | 0 | 0
|  |  | 3.85 | 0.12495 | 1 | -0.146 | 52.894 | 20.089 | 0 | 0
|  |  | 3.9 | 0.17577 | 1 | -0.146 | 52.884 | 20.089 | 0 | 0
|  |  | 3.95 | 0.22658 | 1 | -0.146 | 52.874 | 20.089 | 0 | 0
|  |  | 4 | 0.2774 | 1 | -0.146 | 52.864 | 20.089 | 0 | 0
|  |  | 4.05 | 0.32831 | 1 | -0.146 | 52.854 | 20.089 | 0 | 0
|  |  | 4.1 | 0.37922 | 1 | -0.146 | 52.844 | 20.089 | 0 | 0
|  |  | 4.15 | 0.43013 | 1 | -0.146 | 52.834 | 20.089 | 0 | 0
|  |  | 4.2 | 0.48104 | 1 | -0.146 | 52.824 | 20.089 | 0 | 0
|  |  | 4.25 | 0.48528 | 1 | -0.146 | 52.814 | 20.089 | 0 | 0
|  |  | 4.3 | 0.48952 | 1 | -0.146 | 52.804 | 20.089 | 0 | 0
|  |  | 4.35 | 0.49376 | 1 | -0.146 | 52.794 | 20.089 | 0 | 0
|  |  | 4.4 | 0.49801 | 1 | -0.146 | 52.784 | 20.089 | 0 | 0
|  |  | 4.45 | 0.50225 | 1 | -0.146 | 52.774 | 20.089 | 0 | 0
|  |  | 4.5 | 0.5065 | 1 | -0.146 | 52.764 | 20.089 | 0 | 0
|  |  | 4.55 | 0.51075 | 1 | -0.146 | 52.754 | 20.089 | 0 | 0
|  |  | 4.6 | 0.51499 | 1 | -0.146 | 52.744 | 20.089 | 0 | 0
|  |  | 4.65 | 0.51925 | 1 | -0.146 | 52.734 | 20.089 | 0 | 0
|  |  | 4.7 | 0.5235 | 1 | -0.146 | 52.724 | 20.089 | 0 | 0
|  |  | 4.75 | 0.52775 | 1 | -0.146 | 52.714 | 20.089 | 0 | 0
|  |  | 4.8 | 0.532 | 1 | -0.146 | 52.704 | 20.089 | 0 | 0
|  |  | 4.85 | 0.53625 | 1 | -0.146 | 52.694 | 20.089 | 0 | 0
|  |  | 4.9 | 0.54051 | 1 | -0.146 | 52.684 | 20.089 | 0 | 0
|  |  | 4.95 | 0.54477 | 1 | -0.146 | 52.674 | 20.089 | 0 | 0
|  |  | 5 | 0.54902 | 1 | -0.146 | 52.664 | 20.089 | 0 | 0
|  |  | 5.05 | 0.55435 | 1 | -0.146 | 52.654 | 20.089 | 0 | 0
|  |  | 5.1 | 0.55967 | 1 | -0.146 | 52.644 | 20.089 | 0 | 0
|  |  | 5.15 | 0.565 | 1 | -0.146 | 52.634 | 20.089 | 0 | 0
|  |  | 5.2 | 0.57032 | 1 | -0.146 | 52.624 | 20.089 | 0 | 0
|  |  | 5.25 | 0.57565 | 1 | -0.146 | 52.614 | 20.089 | 0 | 0
|  |  | 5.3 | 0.58098 | 1 | -0.146 | 52.604 | 20.089 | 0 | 0
|  |  | 5.35 | 0.58631 | 1 | -0.146 | 52.594 | 20.089 | 0 | 0
|  |  | 5.4 | 0.59164 | 1 | -0.146 | 52.584 | 20.089 | 0 | 0
|  |  | 5.45 | 0.59698 | 1 | -0.146 | 52.574 | 20.089 | 0 | 0
|  |  | 5.5 | 0.60231 | 1 | -0.146 | 52.564 | 20.089 | 0 | 0
|  |  | 5.55 | 0.60765 | 1 | -0.146 | 52.554 | 20.089 | 0 | 0
|  |  | 5.6 | 0.61299 | 1 | -0.146 | 52.544 | 20.089 | 0 | 0
|  |  | 5.65 | 0.61833 | 1 | -0.146 | 52.534 | 20.089 | 0 | 0
|  |  | 5.7 | 0.62367 | 1 | -0.146 | 52.524 | 20.089 | 0 | 0
|  |  | 5.75 | 0.62901 | 1 | -0.146 | 52.514 | 20.089 | 0 | 0
|  |  | 5.8 | 0.63435 | 1 | -0.146 | 52.504 | 20.089 | 0 | 0
|  |  | 5.85 | 0.6408 | 1 | -0.146 | 52.494 | 20.089 | 0 | 0
|  |  | 5.9 | 0.64724 | 1 | -0.146 | 52.484 | 20.089 | 0 | 0
|  |  | 5.95 | 0.65369 | 1 | -0.146 | 52.474 | 20.089 | 0 | 0
|  |  | 6 | 0.66013 | 1 | -0.146 | 52.464 | 20.089 | 0 | 0
|  |  | 6.05 | 0.66658 | 1 | -0.146 | 52.454 | 20.089 | 0 | 0
|  |  | 6.1 | 0.67303 | 1 | -0.146 | 52.444 | 20.089 | 0 | 0
|  |  | 6.15 | 0.67949 | 1 | -0.146 | 52.434 | 20.089 | 0 | 0
|  |  | 6.2 | 0.68594 | 1 | -0.146 | 52.424 | 20.088 | 0 | 0
|  |  | 6.25 | 0.6924 | 1 | -0.146 | 52.414 | 20.088 | 0 | 0
|  |  | 6.3 | 0.69885 | 1 | -0.146 | 52.404 | 20.088 | 0 | 0
|  |  | 6.35 | 0.70531 | 1 | -0.146 | 52.394 | 20.088 | 0 | 0
|  |  | 6.4 | 0.71177 | 1 | -0.146 | 52.384 | 20.088 | 0 | 0
|  |  | 6.45 | 0.71824 | 1 | -0.146 | 52.374 | 20.088 | 0 | 0
|  |  | 6.5 | 0.7247 | 1 | -0.146 | 52.364 | 20.088 | 0 | 0
|  |  | 6.55 | 0.73117 | 1 | -0.146 | 52.354 | 20.088 | 0 | 0
|  |  | 6.6 | 0.73763 | 1 | -0.146 | 52.344 | 20.088 | 0 | 0
|  |  | 6.65 | 0.74543 | 1 | -0.146 | 52.334 | 20.088 | 0 | 0
|  |  | 6.7 | 0.75323 | 1 | -0.146 | 52.324 | 20.088 | 0 | 0
|  |  | 6.75 | 0.76102 | 1 | -0.146 | 52.314 | 20.088 | 0 | 0
|  |  | 6.8 | 0.76882 | 1 | -0.146 | 52.304 | 20.088 | 0 | 0
|  |  | 6.85 | 0.77663 | 1 | -0.146 | 52.294 | 20.088 | 0 | 0
|  |  | 6.9 | 0.78443 | 1 | -0.146 | 52.284 | 20.088 | 0 | 0
|  |  | 6.95 | 0.79224 | 1 | -0.146 | 52.274 | 20.088 | 0 | 0
|  |  | 7 | 0.80004 | 1 | -0.146 | 52.264 | 20.088 | 0 | 0
|  |  | 7.05 | 0.80786 | 1 | -0.146 | 52.254 | 20.088 | 0 | 0
|  |  | 7.1 | 0.81567 | 1 | -0.146 | 52.244 | 20.088 | 0 | 0
|  |  | 7.15 | 0.82349 | 1 | -0.146 | 52.234 | 20.088 | 0 | 0
|  |  | 7.2 | 0.8313 | 1 | -0.146 | 52.224 | 20.088 | 0 | 0
|  |  | 7.25 | 0.83912 | 1 | -0.146 | 52.214 | 20.088 | 0 | 0
|  |  | 7.3 | 0.84694 | 1 | -0.146 | 52.204 | 20.088 | 0 | 0
|  |  | 7.35 | 0.85476 | 1 | -0.146 | 52.194 | 20.088 | 0 | 0
|  |  | 7.4 | 0.86258 | 1 | -0.146 | 52.184 | 20.088 | 0 | 0
|  |  | 7.45 | 0.87234 | 1 | -0.146 | 52.174 | 20.088 | 0 | 0
|  |  | 7.5 | 0.88211 | 1 | -0.146 | 52.164 | 20.088 | 0 | 0
|  |  | 7.55 | 0.89188 | 1 | -0.146 | 52.154 | 20.088 | 0 | 0
|  |  | 7.6 | 0.90164 | 1 | -0.146 | 52.144 | 20.088 | 0 | 0
|  |  | 7.65 | 0.91142 | 1 | -0.146 | 52.134 | 20.088 | 0 | 0
|  |  | 7.7 | 0.92119 | 1 | -0.146 | 52.124 | 20.088 | 0 | 0
|  |  | 7.75 | 0.93096 | 1 | -0.146 | 52.114 | 20.088 | 0 | 0
|  |  | 7.8 | 0.94074 | 1 | -0.146 | 52.104 | 20.088 | 0 | 0
|  |  | 7.85 | 0.95052 | 1 | -0.146 | 52.094 | 20.088 | 0 | 0
|  |  | 7.9 | 0.96031 | 1 | -0.146 | 52.084 | 20.088 | 0 | 0
|  |  | 7.95 | 0.97009 | 1 | -0.146 | 52.074 | 20.088 | 0 | 0
|  |  | 8 | 0.97988 | 1 | -0.146 | 52.064 | 20.088 | 0 | 0
|  |  | 8.05 | 0.98967 | 1 | -0.146 | 52.054 | 20.088 | 0 | 0
|  |  | 8.1 | 0.99946 | 1 | -0.146 | 52.044 | 20.088 | 0 | 0
|  |  | 8.15 | 1.0093 | 1 | -0.146 | 52.034 | 20.088 | 0 | 0
|  |  | 8.2 | 1.0191 | 1 | -0.146 | 52.024 | 20.088 | 0 | 0
|  |  | 8.25 | 1.0404 | 1 | -0.146 | 52.014 | 20.088 | 0 | 0
|  |  | 8.3 | 1.0617 | 1 | -0.146 | 52.004 | 20.088 | 0 | 0
|  |  | 8.35 | 1.083 | 1 | -0.146 | 51.994 | 20.088 | 0 | 0
|  |  | 8.4 | 1.1043 | 1 | -0.146 | 51.984 | 20.088 | 0 | 0
|  |  | 8.45 | 1.1257 | 1 | -0.146 | 51.975 | 20.088 | 0 | 0
|  |  | 8.5 | 1.147 | 1 | -0.146 | 51.965 | 20.088 | 0 | 0
|  |  | 8.55 | 1.1683 | 1 | -0.146 | 51.955 | 20.088 | 0 | 0
|  |  | 8.6 | 1.1897 | 1 | -0.146 | 51.945 | 20.088 | 0 | 0
|  |  | 8.65 | 1.2111 | 1 | -0.146 | 51.935 | 20.088 | 0 | 0
|  |  | 8.7 | 1.2324 | 1 | -0.146 | 51.925 | 20.088 | 0 | 0
|  |  | 8.75 | 1.2538 | 1 | -0.146 | 51.915 | 20.088 | 0 | 0
|  |  | 8.8 | 1.2751 | 1 | -0.146 | 51.905 | 20.088 | 0 | 0
|  |  | 8.85 | 1.2965 | 1 | -0.146 | 51.895 | 20.088 | 0 | 0
|  |  | 8.9 | 1.3179 | 1 | -0.146 | 51.885 | 20.088 | 0 | 0
|  |  | 8.95 | 1.3393 | 1 | -0.146 | 51.875 | 20.088 | 0 | 0
|  |  | 9 | 1.3607 | 1 | -0.146 | 51.865 | 20.088 | 0 | 0
|  |  | 9.05 | 1.3849 | 1 | -0.146 | 51.855 | 20.088 | 0 | 0
|  |  | 9.1 | 1.4092 | 1 | -0.146 | 51.845 | 20.088 | 0 | 0
|  |  | 9.15 | 1.4335 | 1 | -0.146 | 51.835 | 20.088 | 0 | 0
|  |  | 9.2 | 1.4578 | 1 | -0.146 | 51.825 | 20.088 | 0 | 0
|  |  | 9.25 | 1.4821 | 1 | -0.146 | 51.815 | 20.088 | 0 | 0
|  |  | 9.3 | 1.5064 | 1 | -0.146 | 51.805 | 20.088 | 0 | 0
|  |  | 9.35 | 1.5307 | 1 | -0.146 | 51.795 | 20.088 | 0 | 0
|  |  | 9.4 | 1.555 | 1 | -0.146 | 51.785 | 20.088 | 0 | 0
|  |  | 9.45 | 1.5794 | 1 | -0.146 | 51.775 | 20.088 | 0 | 0
|  |  | 9.5 | 1.6037 | 1 | -0.146 | 51.765 | 20.088 | 0 | 0
|  |  | 9.55 | 1.628 | 1 | -0.146 | 51.755 | 20.088 | 0 | 0
|  |  | 9.6 | 1.6523 | 1 | -0.146 | 51.745 | 20.088 | 0 | 0
|  |  | 9.65 | 1.6767 | 1 | -0.146 | 51.735 | 20.087 | 0 | 0
|  |  | 9.7 | 1.701 | 1 | -0.146 | 51.725 | 20.087 | 0 | 0
|  |  | 9.75 | 1.7254 | 1 | -0.146 | 51.715 | 20.087 | 0 | 0
|  |  | 9.8 | 1.7497 | 1 | -0.146 | 51.705 | 20.087 | 0 | 0
|  |  | 9.85 | 1.7719 | 1 | -0.146 | 51.695 | 20.087 | 0 | 0
|  |  | 9.9 | 1.794 | 1 | -0.146 | 51.685 | 20.087 | 0 | 0
|  |  | 9.95 | 1.8162 | 1 | -0.146 | 51.675 | 20.087 | 0 | 0
|  |  | 10 | 1.8383 | 1 | -0.146 | 51.665 | 20.087 | 0 | 0
|  |  | 10.05 | 1.8605 | 1 | -0.146 | 51.655 | 20.087 | 0 | 0
|  |  | 10.1 | 1.8826 | 1 | -0.146 | 51.645 | 20.087 | 0 | 0
|  |  | 10.15 | 1.9048 | 1 | -0.146 | 51.635 | 20.087 | 0 | 0
|  |  | 10.2 | 1.927 | 1 | -0.146 | 51.625 | 20.087 | 0 | 0
|  |  | 10.25 | 1.9491 | 1 | -0.146 | 51.615 | 20.087 | 0 | 0
|  |  | 10.3 | 1.9713 | 1 | -0.146 | 51.605 | 20.087 | 0 | 0
|  |  | 10.35 | 1.9935 | 1 | -0.146 | 51.595 | 20.087 | 0 | 0
|  |  | 10.4 | 2.0157 | 1 | -0.146 | 51.585 | 20.087 | 0 | 0
|  |  | 10.45 | 2.0378 | 1 | -0.146 | 51.575 | 20.087 | 0 | 0
|  |  | 10.5 | 2.06 | 1 | -0.146 | 51.565 | 20.087 | 0 | 0
|  |  | 10.55 | 2.0822 | 1 | -0.146 | 51.555 | 20.087 | 0 | 0
|  |  | 10.6 | 2.1044 | 1 | -0.146 | 51.545 | 20.087 | 0 | 0
|  |  | 10.65 | 2.1304 | 1 | -0.146 | 51.535 | 20.087 | 0 | 0
|  |  | 10.7 | 2.1563 | 1 | -0.146 | 51.525 | 20.087 | 0 | 0
|  |  | 10.75 | 2.1822 | 1 | -0.146 | 51.515 | 20.087 | 0 | 0
|  |  | 10.8 | 2.2082 | 1 | -0.146 | 51.505 | 20.087 | 0 | 0
|  |  | 10.85 | 2.2341 | 1 | -0.146 | 51.495 | 20.087 | 0 | 0
|  |  | 10.9 | 2.2601 | 1 | -0.146 | 51.485 | 20.087 | 0 | 0
|  |  | 10.95 | 2.2861 | 1 | -0.146 | 51.475 | 20.087 | 0 | 0
|  |  | 11 | 2.312 | 1 | -0.146 | 51.465 | 20.087 | 0 | 0
|  |  | 11.05 | 2.338 | 1 | -0.146 | 51.455 | 20.087 | 0 | 0
|  |  | 11.1 | 2.364 | 1 | -0.146 | 51.445 | 20.087 | 0 | 0
|  |  | 11.15 | 2.3899 | 1 | -0.146 | 51.435 | 20.087 | 0 | 0
|  |  | 11.2 | 2.4159 | 1 | -0.146 | 51.425 | 20.087 | 0 | 0
|  |  | 11.25 | 2.4419 | 1 | -0.146 | 51.415 | 20.087 | 0 | 0
|  |  | 11.3 | 2.4679 | 1 | -0.146 | 51.405 | 20.087 | 0 | 0
|  |  | 11.35 | 2.4939 | 1 | -0.146 | 51.395 | 20.087 | 0 | 0
|  |  | 11.4 | 2.5199 | 1 | -0.146 | 51.385 | 20.087 | 0 | 0
|  |  | 11.45 | 2.555 | 1 | -0.146 | 51.375 | 20.087 | 0 | 0
|  |  | 11.5 | 2.5901 | 1 | -0.146 | 51.365 | 20.087 | 0 | 0
|  |  | 11.55 | 2.6252 | 1 | -0.146 | 51.355 | 20.087 | 0 | 0
|  |  | 11.6 | 2.6603 | 1 | -0.146 | 51.345 | 20.087 | 0 | 0
|  |  | 11.65 | 2.6955 | 1 | -0.146 | 51.335 | 20.087 | 0 | 0
|  |  | 11.7 | 2.7306 | 1 | -0.146 | 51.325 | 20.087 | 0 | 0
|  |  | 11.75 | 2.7657 | 1 | -0.146 | 51.315 | 20.087 | 0 | 0
|  |  | 11.8 | 2.8009 | 1 | -0.146 | 51.305 | 20.087 | 0 | 0
|  |  | 11.85 | 2.836 | 1 | -0.146 | 51.295 | 20.087 | 0 | 0
|  |  | 11.9 | 2.8712 | 1 | -0.146 | 51.285 | 20.087 | 0 | 0
|  |  | 11.95 | 2.9064 | 1 | -0.146 | 51.275 | 20.087 | 0 | 0
|  |  | 12 | 2.9415 | 1 | -0.146 | 51.265 | 20.087 | 0 | 0
|  |  | 12.05 | 2.9767 | 1 | -0.146 | 51.255 | 20.087 | 0 | 0
|  |  | 12.1 | 3.0119 | 1 | -0.146 | 51.245 | 20.087 | 0 | 0
|  |  | 12.15 | 3.0471 | 1 | -0.146 | 51.235 | 20.087 | 0 | 0
|  |  | 12.2 | 3.0823 | 1 | -0.146 | 51.225 | 20.087 | 0 | 0
|  |  | 12.25 | 3.1277 | 1 | -0.146 | 51.215 | 20.087 | 0 | 0
|  |  | 12.3 | 3.1731 | 1 | -0.146 | 51.205 | 20.087 | 0 | 0
|  |  | 12.35 | 3.2184 | 1 | -0.146 | 51.195 | 20.087 | 0 | 0
|  |  | 12.4 | 3.2638 | 1 | -0.146 | 51.185 | 20.087 | 0 | 0
|  |  | 12.45 | 3.3092 | 1 | -0.146 | 51.175 | 20.087 | 0 | 0
|  |  | 12.5 | 3.3547 | 1 | -0.146 | 51.165 | 20.087 | 0 | 0
|  |  | 12.55 | 3.4001 | 1 | -0.146 | 51.155 | 20.087 | 0 | 0
|  |  | 12.6 | 3.4455 | 1 | -0.146 | 51.145 | 20.087 | 0 | 0
|  |  | 12.65 | 3.491 | 1 | -0.146 | 51.135 | 20.087 | 0 | 0
|  |  | 12.7 | 3.5364 | 1 | -0.146 | 51.125 | 20.087 | 0 | 0
|  |  | 12.75 | 3.5819 | 1 | -0.146 | 51.115 | 20.087 | 0 | 0
|  |  | 12.8 | 3.6274 | 1 | -0.146 | 51.105 | 20.087 | 0 | 0
|  |  | 12.85 | 3.6728 | 1 | -0.146 | 51.095 | 20.087 | 0 | 0
|  |  | 12.9 | 3.7183 | 1 | -0.146 | 51.085 | 20.087 | 0 | 0
|  |  | 12.95 | 3.7638 | 1 | -0.146 | 51.075 | 20.087 | 0 | 0
|  |  | 13 | 3.8093 | 1 | -0.146 | 51.065 | 20.087 | 0 | 0
|  |  | 13.05 | 3.8981 | 1 | -0.146 | 51.055 | 20.087 | 0 | 0
|  |  | 13.1 | 3.9868 | 1 | -0.146 | 51.045 | 20.086 | 0 | 0
|  |  | 13.15 | 4.0756 | 1 | -0.146 | 51.035 | 20.086 | 0 | 0
|  |  | 13.2 | 4.1644 | 1 | -0.146 | 51.025 | 20.086 | 0 | 0
|  |  | 13.25 | 4.2535 | 1 | -0.146 | 51.015 | 20.086 | 0 | 0
|  |  | 13.3 | 4.3427 | 1 | -0.146 | 51.005 | 20.086 | 0 | 0
|  |  | 13.35 | 4.4318 | 1 | -0.146 | 50.995 | 20.086 | 0 | 0
|  |  | 13.4 | 4.521 | 1 | -0.146 | 50.985 | 20.086 | 0 | 0
|  |  | 13.45 | 4.7105 | 1 | -0.146 | 50.975 | 20.086 | 0 | 0
|  |  | 13.5 | 4.8999 | 1 | -0.146 | 50.965 | 20.086 | 0 | 0
|  |  | 13.55 | 5.0894 | 1 | -0.146 | 50.955 | 20.086 | 0 | 0
|  |  | 13.6 | 5.2788 | 1 | -0.146 | 50.945 | 20.086 | 0 | 0
|  |  | 13.65 | 5.8337 | 1 | -0.146 | 50.935 | 20.086 | 0 | 0
|  |  | 13.7 | 6.3885 | 1 | -0.146 | 50.925 | 20.086 | 0 | 0
|  |  | 13.75 | 6.9434 | 1 | -0.146 | 50.915 | 20.086 | 0 | 0
|  |  | 13.8 | 7.4982 | 1 | -0.146 | 50.905 | 20.086 | 0 | 0
|  |  | 13.85 | 10.393 | 1 | -0.146 | 50.895 | 20.086 | 0 | 0
|  |  | 13.9 | 13.288 | 1 | -0.146 | 50.885 | 20.086 | 0 | 0
|  |  | 13.95 | 16.183 | 1 | -0.146 | 50.875 | 20.086 | 0 | 0
|  |  | 14 | 19.078 | 1 | -0.146 | 50.865 | 20.086 | 0 | 0
|  |  | 14.05 | 23.24 | 1 | -0.146 | 50.855 | 20.086 | 0 | 0
|  |  | 14.1 | 27.403 | 1 | -0.146 | 50.845 | 20.086 | 0 | 0
|  |  | 14.15 | 31.566 | 1 | -0.146 | 50.835 | 20.086 | 0 | 0
|  |  | 14.2 | 35.729 | 1 | -0.146 | 50.825 | 20.086 | 0 | 0
|  |  | 14.25 | 35.944 | 1 | -0.146 | 50.815 | 20.086 | 0 | 0
|  |  | 14.3 | 36.158 | 1 | -0.146 | 50.805 | 20.086 | 0 | 0
|  |  | 14.35 | 36.373 | 1 | -0.146 | 50.795 | 20.086 | 0 | 0
|  |  | 14.4 | 36.588 | 1 | -0.146 | 50.785 | 20.086 | 0 | 0
|  |  | 14.45 | 32.4 | 1 | -0.146 | 50.775 | 20.086 | 0 | 0
|  |  | 14.5 | 28.213 | 1 | -0.146 | 50.765 | 20.086 | 0 | 0
|  |  | 14.55 | 24.025 | 1 | -0.146 | 50.755 | 20.086 | 0 | 0
|  |  | 14.6 | 19.837 | 1 | -0.146 | 50.745 | 20.086 | 0 | 0
|  |  | 14.65 | 17.708 | 1 | -0.146 | 50.735 | 20.086 | 0 | 0
|  |  | 14.7 | 15.58 | 1 | -0.146 | 50.725 | 20.086 | 0 | 0
|  |  | 14.75 | 13.451 | 1 | -0.146 | 50.715 | 20.086 | 0 | 0
|  |  | 14.8 | 11.322 | 1 | -0.146 | 50.705 | 20.086 | 0 | 0
|  |  | 14.85 | 12.022 | 1 | -0.146 | 50.695 | 20.086 | 0 | 0
|  |  | 14.9 | 12.723 | 1 | -0.146 | 50.685 | 20.086 | 0 | 0
|  |  | 14.95 | 13.424 | 1 | -0.146 | 50.675 | 20.086 | 0 | 0
|  |  | 15 | 14.124 | 1 | -0.146 | 50.665 | 20.086 | 0 | 0
|  |  | 15.05 | 18.918 | 1 | -0.146 | 50.655 | 20.086 | 0 | 0
|  |  | 15.1 | 23.712 | 1 | -0.146 | 50.645 | 20.086 | 0 | 0
|  |  | 15.15 | 28.505 | 1 | -0.146 | 50.635 | 20.086 | 0 | 0
|  |  | 15.2 | 33.299 | 1 | -0.146 | 50.625 | 20.086 | 0 | 0
|  |  | 15.25 | 34.945 | 1 | -0.146 | 50.615 | 20.086 | 0 | 0
|  |  | 15.3 | 36.592 | 1 | -0.146 | 50.605 | 20.086 | 0 | 0
|  |  | 15.35 | 38.238 | 1 | -0.146 | 50.595 | 20.086 | 0 | 0
|  |  | 15.4 | 39.885 | 1 | -0.146 | 50.585 | 20.086 | 0 | 0
|  |  | 15.45 | 37.639 | 1 | -0.146 | 50.575 | 20.086 | 0 | 0
|  |  | 15.5 | 35.392 | 1 | -0.146 | 50.565 | 20.086 | 0 | 0
|  |  | 15.55 | 33.146 | 1 | -0.146 | 50.555 | 20.086 | 0 | 0
|  |  | 15.6 | 30.9 | 1 | -0.146 | 50.545 | 20.086 | 0 | 0
|  |  | 15.65 | 26.77 | 1 | -0.146 | 50.535 | 20.086 | 0 | 0
|  |  | 15.7 | 22.64 | 1 | -0.146 | 50.525 | 20.086 | 0 | 0
|  |  | 15.75 | 18.51 | 1 | -0.146 | 50.515 | 20.086 | 0 | 0
|  |  | 15.8 | 14.38 | 1 | -0.146 | 50.505 | 20.086 | 0 | 0
|  |  | 15.85 | 13.935 | 1 | -0.146 | 50.495 | 20.086 | 0 | 0
|  |  | 15.9 | 13.489 | 1 | -0.146 | 50.485 | 20.086 | 0 | 0
|  |  | 15.95 | 13.044 | 1 | -0.146 | 50.475 | 20.086 | 0 | 0
|  |  | 16 | 12.598 | 1 | -0.146 | 50.465 | 20.086 | 0 | 0
|  |  | 16.05 | 15.416 | 1 | -0.146 | 50.455 | 20.086 | 0 | 0
|  |  | 16.1 | 18.234 | 1 | -0.146 | 50.445 | 20.086 | 0 | 0
|  |  | 16.15 | 21.052 | 1 | -0.146 | 50.435 | 20.086 | 0 | 0
|  |  | 16.2 | 23.871 | 1 | -0.146 | 50.425 | 20.086 | 0 | 0
|  |  | 16.25 | 27.71 | 1 | -0.146 | 50.415 | 20.086 | 0 | 0
|  |  | 16.3 | 31.549 | 1 | -0.146 | 50.405 | 20.086 | 0 | 0
|  |  | 16.35 | 35.388 | 1 | -0.146 | 50.395 | 20.086 | 0 | 0
|  |  | 16.4 | 39.227 | 1 | -0.146 | 50.385 | 20.086 | 0 | 0
|  |  | 16.45 | 39.017 | 1 | -0.146 | 50.375 | 20.086 | 0 | 0
|  |  | 16.5 | 38.807 | 1 | -0.146 | 50.365 | 20.086 | 0 | 0
|  |  | 16.55 | 38.597 | 1 | -0.146 | 50.355 | 20.085 | 0 | 0
|  |  | 16.6 | 38.387 | 1 | -0.146 | 50.345 | 20.085 | 0 | 0
|  |  | 16.65 | 34.134 | 1 | -0.146 | 50.335 | 20.085 | 0 | 0
|  |  | 16.7 | 29.881 | 1 | -0.146 | 50.325 | 20.085 | 0 | 0
|  |  | 16.75 | 25.627 | 1 | -0.146 | 50.315 | 20.085 | 0 | 0
|  |  | 16.8 | 21.374 | 1 | -0.146 | 50.305 | 20.085 | 0 | 0
|  |  | 16.85 | 19.276 | 1 | -0.146 | 50.295 | 20.085 | 0 | 0
|  |  | 16.9 | 17.178 | 1 | -0.146 | 50.285 | 20.085 | 0 | 0
|  |  | 16.95 | 15.08 | 1 | -0.146 | 50.275 | 20.085 | 0 | 0
|  |  | 17 | 12.981 | 1 | -0.146 | 50.265 | 20.085 | 0 | 0
|  |  | 17.05 | 13.686 | 1 | -0.146 | 50.255 | 20.085 | 0 | 0
|  |  | 17.1 | 14.391 | 1 | -0.146 | 50.245 | 20.085 | 0 | 0
|  |  | 17.15 | 15.096 | 1 | -0.146 | 50.235 | 20.085 | 0 | 0
|  |  | 17.2 | 15.801 | 1 | -0.146 | 50.225 | 20.085 | 0 | 0
|  |  | 17.25 | 20.527 | 1 | -0.146 | 50.215 | 20.085 | 0 | 0
|  |  | 17.3 | 25.253 | 1 | -0.146 | 50.205 | 20.085 | 0 | 0
|  |  | 17.35 | 29.978 | 1 | -0.146 | 50.195 | 20.085 | 0 | 0
|  |  | 17.4 | 34.704 | 1 | -0.146 | 50.185 | 20.085 | 0 | 0
|  |  | 17.45 | 36.231 | 1 | -0.146 | 50.175 | 20.085 | 0 | 0
|  |  | 17.5 | 37.758 | 1 | -0.146 | 50.165 | 20.085 | 0 | 0
|  |  | 17.55 | 39.285 | 1 | -0.146 | 50.155 | 20.085 | 0 | 0
|  |  | 17.6 | 40.812 | 1 | -0.146 | 50.145 | 20.085 | 0 | 0
|  |  | 17.65 | 38.495 | 1 | -0.146 | 50.135 | 20.085 | 0 | 0
|  |  | 17.7 | 36.178 | 1 | -0.146 | 50.125 | 20.085 | 0 | 0
|  |  | 17.75 | 33.861 | 1 | -0.146 | 50.115 | 20.085 | 0 | 0
|  |  | 17.8 | 31.544 | 1 | -0.146 | 50.105 | 20.085 | 0 | 0
|  |  | 17.85 | 27.481 | 1 | -0.146 | 50.095 | 20.085 | 0 | 0
|  |  | 17.9 | 23.418 | 1 | -0.146 | 50.085 | 20.085 | 0 | 0
|  |  | 17.95 | 19.355 | 1 | -0.146 | 50.075 | 20.085 | 0 | 0
|  |  | 18 | 15.292 | 1 | -0.146 | 50.065 | 20.085 | 0 | 0
|  |  | 18.05 | 14.885 | 1 | -0.146 | 50.055 | 20.085 | 0 | 0
|  |  | 18.1 | 14.478 | 1 | -0.146 | 50.045 | 20.085 | 0 | 0
|  |  | 18.15 | 14.071 | 1 | -0.146 | 50.035 | 20.085 | 0 | 0
|  |  | 18.2 | 13.665 | 1 | -0.146 | 50.025 | 20.085 | 0 | 0
|  |  | 18.25 | 16.479 | 1 | -0.146 | 50.015 | 20.085 | 0 | 0
|  |  | 18.3 | 19.293 | 1 | -0.146 | 50.005 | 20.085 | 0 | 0
|  |  | 18.35 | 22.106 | 1 | -0.146 | 49.995 | 20.085 | 0 | 0
|  |  | 18.4 | 24.92 | 1 | -0.146 | 49.985 | 20.085 | 0 | 0
|  |  | 18.45 | 28.655 | 1 | -0.146 | 49.975 | 20.085 | 0 | 0
|  |  | 18.5 | 32.389 | 1 | -0.146 | 49.965 | 20.085 | 0 | 0
|  |  | 18.55 | 36.123 | 1 | -0.146 | 49.955 | 20.085 | 0 | 0
|  |  | 18.6 | 39.857 | 1 | -0.146 | 49.945 | 20.085 | 0 | 0
|  |  | 18.65 | 39.578 | 1 | -0.146 | 49.935 | 20.085 | 0 | 0
|  |  | 18.7 | 39.298 | 1 | -0.146 | 49.925 | 20.085 | 0 | 0
|  |  | 18.75 | 39.019 | 1 | -0.146 | 49.915 | 20.085 | 0 | 0
|  |  | 18.8 | 38.739 | 1 | -0.146 | 49.905 | 20.085 | 0 | 0
|  |  | 18.85 | 34.493 | 1 | -0.146 | 49.895 | 20.085 | 0 | 0
|  |  | 18.9 | 30.247 | 1 | -0.146 | 49.885 | 20.085 | 0 | 0
|  |  | 18.95 | 26.001 | 1 | -0.146 | 49.875 | 20.085 | 0 | 0
|  |  | 19 | 21.755 | 1 | -0.146 | 49.865 | 20.085 | 0 | 0
|  |  | 19.05 | 19.727 | 1 | -0.146 | 49.855 | 20.085 | 0 | 0
|  |  | 19.1 | 17.698 | 1 | -0.146 | 49.845 | 20.085 | 0 | 0
|  |  | 19.15 | 15.67 | 1 | -0.146 | 49.835 | 20.085 | 0 | 0
|  |  | 19.2 | 13.642 | 1 | -0.146 | 49.825 | 20.085 | 0 | 0
|  |  | 19.25 | 14.379 | 1 | -0.146 | 49.815 | 20.085 | 0 | 0
|  |  | 19.3 | 15.115 | 1 | -0.146 | 49.805 | 20.085 | 0 | 0
|  |  | 19.35 | 15.852 | 1 | -0.146 | 49.795 | 20.085 | 0 | 0
|  |  | 19.4 | 16.589 | 1 | -0.146 | 49.785 | 20.085 | 0 | 0
|  |  | 19.45 | 21.29 | 1 | -0.146 | 49.775 | 20.085 | 0 | 0
|  |  | 19.5 | 25.99 | 1 | -0.146 | 49.765 | 20.085 | 0 | 0
|  |  | 19.55 | 30.69 | 1 | -0.146 | 49.755 | 20.085 | 0 | 0
|  |  | 19.6 | 35.39 | 1 | -0.146 | 49.745 | 20.085 | 0 | 0
|  |  | 19.65 | 36.82 | 1 | -0.146 | 49.735 | 20.085 | 0 | 0
|  |  | 19.7 | 38.251 | 1 | -0.146 | 49.725 | 20.085 | 0 | 0
|  |  | 19.75 | 39.681 | 1 | -0.146 | 49.715 | 20.085 | 0 | 0
|  |  | 19.8 | 41.111 | 1 | -0.146 | 49.705 | 20.085 | 0 | 0
|  |  | 19.85 | 38.746 | 1 | -0.146 | 49.695 | 20.085 | 0 | 0
|  |  | 19.9 | 36.381 | 1 | -0.146 | 49.685 | 20.085 | 0 | 0
|  |  | 19.95 | 34.015 | 1 | -0.146 | 49.675 | 20.085 | 0 | 0
|  |  | 20 | 31.65 | 1 | -0.146 | 49.665 | 20.084 | 0 | 0
|  |  | 20.05 | 27.674 | 1 | -0.146 | 49.655 | 20.084 | 0 | 0
|  |  | 20.1 | 23.699 | 1 | -0.146 | 49.645 | 20.084 | 0 | 0
|  |  | 20.15 | 19.723 | 1 | -0.146 | 49.635 | 20.084 | 0 | 0
|  |  | 20.2 | 15.748 | 1 | -0.146 | 49.625 | 20.084 | 0 | 0
|  |  | 20.25 | 15.367 | 1 | -0.146 | 49.615 | 20.084 | 0 | 0
|  |  | 20.3 | 14.986 | 1 | -0.146 | 49.605 | 20.084 | 0 | 0
|  |  | 20.35 | 14.605 | 1 | -0.146 | 49.595 | 20.084 | 0 | 0
|  |  | 20.4 | 14.224 | 1 | -0.146 | 49.585 | 20.084 | 0 | 0
|  |  | 20.45 | 17.036 | 1 | -0.146 | 49.575 | 20.084 | 0 | 0
|  |  | 20.5 | 19.848 | 1 | -0.146 | 49.565 | 20.084 | 0 | 0
|  |  | 20.55 | 22.66 | 1 | -0.146 | 49.555 | 20.084 | 0 | 0
|  |  | 20.6 | 25.472 | 1 | -0.146 | 49.545 | 20.084 | 0 | 0
|  |  | 20.65 | 29.131 | 1 | -0.146 | 49.535 | 20.084 | 0 | 0
|  |  | 20.7 | 32.79 | 1 | -0.146 | 49.525 | 20.084 | 0 | 0
|  |  | 20.75 | 36.449 | 1 | -0.146 | 49.515 | 20.084 | 0 | 0
|  |  | 20.8 | 40.108 | 1 | -0.146 | 49.505 | 20.084 | 0 | 0
|  |  | 20.85 | 39.757 | 1 | -0.146 | 49.495 | 20.084 | 0 | 0
|  |  | 20.9 | 39.406 | 1 | -0.146 | 49.485 | 20.084 | 0 | 0
|  |  | 20.95 | 39.055 | 1 | -0.146 | 49.475 | 20.084 | 0 | 0
|  |  | 21 | 38.705 | 1 | -0.146 | 49.465 | 20.084 | 0 | 0
|  |  | 21.05 | 34.495 | 1 | -0.146 | 49.455 | 20.084 | 0 | 0
|  |  | 21.1 | 30.286 | 1 | -0.146 | 49.445 | 20.084 | 0 | 0
|  |  | 21.15 | 26.077 | 1 | -0.146 | 49.435 | 20.084 | 0 | 0
|  |  | 21.2 | 21.868 | 1 | -0.146 | 49.425 | 20.084 | 0 | 0
|  |  | 21.25 | 19.899 | 1 | -0.146 | 49.415 | 20.084 | 0 | 0
|  |  | 21.3 | 17.931 | 1 | -0.146 | 49.405 | 20.084 | 0 | 0
|  |  | 21.35 | 15.962 | 1 | -0.146 | 49.395 | 20.084 | 0 | 0
|  |  | 21.4 | 13.994 | 1 | -0.146 | 49.385 | 20.084 | 0 | 0
|  |  | 21.45 | 14.753 | 1 | -0.146 | 49.375 | 20.084 | 0 | 0
|  |  | 21.5 | 15.513 | 1 | -0.146 | 49.365 | 20.084 | 0 | 0
|  |  | 21.55 | 16.272 | 1 | -0.146 | 49.355 | 20.084 | 0 | 0
|  |  | 21.6 | 17.032 | 1 | -0.146 | 49.345 | 20.084 | 0 | 0
|  |  | 21.65 | 21.698 | 1 | -0.146 | 49.335 | 20.084 | 0 | 0
|  |  | 21.7 | 26.364 | 1 | -0.146 | 49.325 | 20.084 | 0 | 0
|  |  | 21.75 | 31.031 | 1 | -0.146 | 49.315 | 20.084 | 0 | 0
|  |  | 21.8 | 35.697 | 1 | -0.146 | 49.305 | 20.084 | 0 | 0
|  |  | 21.85 | 37.039 | 1 | -0.146 | 49.295 | 20.084 | 0 | 0
|  |  | 21.9 | 38.381 | 1 | -0.146 | 49.285 | 20.084 | 0 | 0
|  |  | 21.95 | 39.723 | 1 | -0.146 | 49.275 | 20.084 | 0 | 0
|  |  | 22 | 41.065 | 1 | -0.146 | 49.265 | 20.084 | 0 | 0
|  |  | 22.05 | 38.666 | 1 | -0.146 | 49.255 | 20.084 | 0 | 0
|  |  | 22.1 | 36.267 | 1 | -0.146 | 49.245 | 20.084 | 0 | 0
|  |  | 22.15 | 33.868 | 1 | -0.146 | 49.235 | 20.084 | 0 | 0
|  |  | 22.2 | 31.469 | 1 | -0.146 | 49.225 | 20.084 | 0 | 0
|  |  | 22.25 | 27.571 | 1 | -0.146 | 49.215 | 20.084 | 0 | 0
|  |  | 22.3 | 23.674 | 1 | -0.146 | 49.205 | 20.084 | 0 | 0
|  |  | 22.35 | 19.776 | 1 | -0.146 | 49.195 | 20.084 | 0 | 0
|  |  | 22.4 | 15.879 | 1 | -0.146 | 49.185 | 20.084 | 0 | 0
|  |  | 22.45 | 15.539 | 1 | -0.146 | 49.175 | 20.084 | 0 | 0
|  |  | 22.5 | 15.2 | 1 | -0.146 | 49.165 | 20.084 | 0 | 0
|  |  | 22.55 | 14.86 | 1 | -0.146 | 49.155 | 20.084 | 0 | 0
|  |  | 22.6 | 14.52 | 1 | -0.146 | 49.145 | 20.084 | 0 | 0
|  |  | 22.65 | 17.326 | 1 | -0.146 | 49.135 | 20.084 | 0 | 0
|  |  | 22.7 | 20.132 | 1 | -0.146 | 49.125 | 20.084 | 0 | 0
|  |  | 22.75 | 22.937 | 1 | -0.146 | 49.115 | 20.084 | 0 | 0
|  |  | 22.8 | 25.743 | 1 | -0.146 | 49.105 | 20.084 | 0 | 0
|  |  | 22.85 | 29.317 | 1 | -0.146 | 49.095 | 20.084 | 0 | 0
|  |  | 22.9 | 32.891 | 1 | -0.146 | 49.085 | 20.084 | 0 | 0
|  |  | 22.95 | 36.465 | 1 | -0.146 | 49.075 | 20.084 | 0 | 0
|  |  | 23 | 40.039 | 1 | -0.146 | 49.065 | 20.084 | 0 | 0
|  |  | 23.05 | 39.638 | 1 | -0.146 | 49.055 | 20.084 | 0 | 0
|  |  | 23.1 | 39.236 | 1 | -0.146 | 49.045 | 20.084 | 0 | 0
|  |  | 23.15 | 38.835 | 1 | -0.146 | 49.035 | 20.084 | 0 | 0
|  |  | 23.2 | 38.433 | 1 | -0.146 | 49.025 | 20.084 | 0 | 0
|  |  | 23.25 | 34.251 | 1 | -0.146 | 49.015 | 20.084 | 0 | 0
|  |  | 23.3 | 30.068 | 1 | -0.146 | 49.005 | 20.084 | 0 | 0
|  |  | 23.35 | 25.885 | 1 | -0.146 | 48.995 | 20.084 | 0 | 0
|  |  | 23.4 | 21.702 | 1 | -0.146 | 48.985 | 20.084 | 0 | 0
|  |  | 23.45 | 19.795 | 1 | -0.146 | 48.976 | 20.083 | 0 | 0
|  |  | 23.5 | 17.888 | 1 | -0.146 | 48.966 | 20.083 | 0 | 0
|  |  | 23.55 | 15.98 | 1 | -0.146 | 48.956 | 20.083 | 0 | 0
|  |  | 23.6 | 14.073 | 1 | -0.146 | 48.946 | 20.083 | 0 | 0
|  |  | 23.65 | 14.874 | 1 | -0.146 | 48.936 | 20.083 | 0 | 0
|  |  | 23.7 | 15.674 | 1 | -0.146 | 48.926 | 20.083 | 0 | 0
|  |  | 23.75 | 16.475 | 1 | -0.146 | 48.916 | 20.083 | 0 | 0
|  |  | 23.8 | 17.275 | 1 | -0.146 | 48.906 | 20.083 | 0 | 0
|  |  | 23.85 | 21.892 | 1 | -0.146 | 48.896 | 20.083 | 0 | 0
|  |  | 23.9 | 26.508 | 1 | -0.146 | 48.886 | 20.083 | 0 | 0
|  |  | 23.95 | 31.124 | 1 | -0.146 | 48.876 | 20.083 | 0 | 0
|  |  | 24 | 35.741 | 1 | -0.146 | 48.866 | 20.083 | 0 | 0
|  |  | 24.05 | 37.015 | 1 | -0.146 | 48.856 | 20.083 | 0 | 0
|  |  | 24.1 | 38.289 | 1 | -0.146 | 48.846 | 20.083 | 0 | 0
|  |  | 24.15 | 39.563 | 1 | -0.146 | 48.836 | 20.083 | 0 | 0
|  |  | 24.2 | 40.837 | 1 | -0.146 | 48.826 | 20.083 | 0 | 0
|  |  | 24.25 | 38.403 | 1 | -0.146 | 48.816 | 20.083 | 0 | 0
|  |  | 24.3 | 35.969 | 1 | -0.146 | 48.806 | 20.083 | 0 | 0
|  |  | 24.35 | 33.536 | 1 | -0.146 | 48.796 | 20.083 | 0 | 0
|  |  | 24.4 | 31.102 | 1 | -0.146 | 48.786 | 20.083 | 0 | 0
|  |  | 24.45 | 27.298 | 1 | -0.146 | 48.776 | 20.083 | 0 | 0
|  |  | 24.5 | 23.495 | 1 | -0.146 | 48.766 | 20.083 | 0 | 0
|  |  | 24.55 | 19.692 | 1 | -0.146 | 48.756 | 20.083 | 0 | 0
|  |  | 24.6 | 15.889 | 1 | -0.146 | 48.746 | 20.083 | 0 | 0
|  |  | 24.65 | 15.589 | 1 | -0.146 | 48.736 | 20.083 | 0 | 0
|  |  | 24.7 | 15.289 | 1 | -0.146 | 48.726 | 20.083 | 0 | 0
|  |  | 24.75 | 14.989 | 1 | -0.146 | 48.716 | 20.083 | 0 | 0
|  |  | 24.8 | 14.689 | 1 | -0.146 | 48.706 | 20.083 | 0 | 0
|  |  | 24.85 | 17.507 | 1 | -0.146 | 48.696 | 20.083 | 0 | 0
|  |  | 24.9 | 20.325 | 1 | -0.146 | 48.686 | 20.083 | 0 | 0
|  |  | 24.95 | 23.143 | 1 | -0.146 | 48.676 | 20.083 | 0 | 0
|  |  | 25 | 25.962 | 1 | -0.146 | 48.666 | 20.083 | 0 | 0
|  |  | 25.05 | 29.45 | 1 | -0.146 | 48.656 | 20.083 | 0 | 0
|  |  | 25.1 | 32.939 | 1 | -0.146 | 48.646 | 20.083 | 0 | 0
|  |  | 25.15 | 36.427 | 1 | -0.146 | 48.636 | 20.083 | 0 | 0
|  |  | 25.2 | 39.916 | 1 | -0.146 | 48.626 | 20.083 | 0 | 0
|  |  | 25.25 | 39.479 | 1 | -0.146 | 48.616 | 20.083 | 0 | 0
|  |  | 25.3 | 39.042 | 1 | -0.146 | 48.606 | 20.083 | 0 | 0
|  |  | 25.35 | 38.605 | 1 | -0.146 | 48.596 | 20.083 | 0 | 0
|  |  | 25.4 | 38.168 | 1 | -0.146 | 48.586 | 20.083 | 0 | 0
|  |  | 25.45 | 34.009 | 1 | -0.146 | 48.576 | 20.083 | 0 | 0
|  |  | 25.5 | 29.85 | 1 | -0.146 | 48.566 | 20.083 | 0 | 0
|  |  | 25.55 | 25.691 | 1 | -0.146 | 48.556 | 20.083 | 0 | 0
|  |  | 25.6 | 21.532 | 1 | -0.146 | 48.546 | 20.083 | 0 | 0
|  |  | 25.65 | 19.677 | 1 | -0.146 | 48.536 | 20.083 | 0 | 0
|  |  | 25.7 | 17.823 | 1 | -0.146 | 48.526 | 20.083 | 0 | 0
|  |  | 25.75 | 15.969 | 1 | -0.146 | 48.516 | 20.083 | 0 | 0
|  |  | 25.8 | 14.114 | 1 | -0.146 | 48.506 | 20.083 | 0 | 0
|  |  | 25.85 | 14.985 | 1 | -0.146 | 48.496 | 20.083 | 0 | 0
|  |  | 25.9 | 15.855 | 1 | -0.146 | 48.486 | 20.083 | 0 | 0
|  |  | 25.95 | 16.725 | 1 | -0.146 | 48.476 | 20.083 | 0 | 0
|  |  | 26 | 17.596 | 1 | -0.146 | 48.466 | 20.083 | 0 | 0
|  |  | 26.05 | 22.216 | 1 | -0.146 | 48.456 | 20.083 | 0 | 0
|  |  | 26.1 | 26.836 | 1 | -0.146 | 48.446 | 20.083 | 0 | 0
|  |  | 26.15 | 31.456 | 1 | -0.146 | 48.436 | 20.083 | 0 | 0
|  |  | 26.2 | 36.076 | 1 | -0.146 | 48.426 | 20.083 | 0 | 0
|  |  | 26.25 | 37.288 | 1 | -0.146 | 48.416 | 20.083 | 0 | 0
|  |  | 26.3 | 38.501 | 1 | -0.146 | 48.406 | 20.083 | 0 | 0
|  |  | 26.35 | 39.713 | 1 | -0.146 | 48.396 | 20.083 | 0 | 0
|  |  | 26.4 | 40.926 | 1 | -0.146 | 48.386 | 20.083 | 0 | 0
|  |  | 26.45 | 38.451 | 1 | -0.146 | 48.376 | 20.083 | 0 | 0
|  |  | 26.5 | 35.976 | 1 | -0.146 | 48.366 | 20.083 | 0 | 0
|  |  | 26.55 | 33.501 | 1 | -0.146 | 48.356 | 20.083 | 0 | 0
|  |  | 26.6 | 31.026 | 1 | -0.146 | 48.346 | 20.083 | 0 | 0
|  |  | 26.65 | 27.212 | 1 | -0.146 | 48.336 | 20.083 | 0 | 0
|  |  | 26.7 | 23.398 | 1 | -0.146 | 48.326 | 20.083 | 0 | 0
|  |  | 26.75 | 19.584 | 1 | -0.146 | 48.316 | 20.083 | 0 | 0
|  |  | 26.8 | 15.77 | 1 | -0.146 | 48.306 | 20.083 | 0 | 0
|  |  | 26.85 | 15.53 | 1 | -0.146 | 48.296 | 20.083 | 0 | 0
|  |  | 26.9 | 15.291 | 1 | -0.146 | 48.286 | 20.082 | 0 | 0
|  |  | 26.95 | 15.052 | 1 | -0.146 | 48.276 | 20.082 | 0 | 0
|  |  | 27 | 14.812 | 1 | -0.146 | 48.266 | 20.082 | 0 | 0
|  |  | 27.05 | 17.708 | 1 | -0.146 | 48.256 | 20.082 | 0 | 0
|  |  | 27.1 | 20.605 | 1 | -0.146 | 48.246 | 20.082 | 0 | 0
|  |  | 27.15 | 23.501 | 1 | -0.146 | 48.236 | 20.082 | 0 | 0
|  |  | 27.2 | 26.398 | 1 | -0.146 | 48.226 | 20.082 | 0 | 0
|  |  | 27.25 | 29.893 | 1 | -0.146 | 48.216 | 20.082 | 0 | 0
|  |  | 27.3 | 33.388 | 1 | -0.146 | 48.206 | 20.082 | 0 | 0
|  |  | 27.35 | 36.883 | 1 | -0.146 | 48.196 | 20.082 | 0 | 0
|  |  | 27.4 | 40.378 | 1 | -0.146 | 48.186 | 20.082 | 0 | 0
|  |  | 27.45 | 39.877 | 1 | -0.146 | 48.176 | 20.082 | 0 | 0
|  |  | 27.5 | 39.375 | 1 | -0.146 | 48.166 | 20.082 | 0 | 0
|  |  | 27.55 | 38.873 | 1 | -0.146 | 48.156 | 20.082 | 0 | 0
|  |  | 27.6 | 38.372 | 1 | -0.146 | 48.146 | 20.082 | 0 | 0
|  |  | 27.65 | 34.115 | 1 | -0.146 | 48.136 | 20.082 | 0 | 0
|  |  | 27.7 | 29.858 | 1 | -0.146 | 48.126 | 20.082 | 0 | 0
|  |  | 27.75 | 25.601 | 1 | -0.146 | 48.116 | 20.082 | 0 | 0
|  |  | 27.8 | 21.344 | 1 | -0.146 | 48.106 | 20.082 | 0 | 0
|  |  | 27.85 | 19.52 | 1 | -0.146 | 48.096 | 20.082 | 0 | 0
|  |  | 27.9 | 17.696 | 1 | -0.146 | 48.086 | 20.082 | 0 | 0
|  |  | 27.95 | 15.872 | 1 | -0.146 | 48.076 | 20.082 | 0 | 0
|  |  | 28 | 14.047 | 1 | -0.146 | 48.066 | 20.082 | 0 | 0
|  |  | 28.05 | 14.969 | 1 | -0.146 | 48.056 | 20.082 | 0 | 0
|  |  | 28.1 | 15.89 | 1 | -0.146 | 48.046 | 20.082 | 0 | 0
|  |  | 28.15 | 16.811 | 1 | -0.146 | 48.036 | 20.082 | 0 | 0
|  |  | 28.2 | 17.732 | 1 | -0.146 | 48.026 | 20.082 | 0 | 0
|  |  | 28.25 | 22.421 | 1 | -0.146 | 48.016 | 20.082 | 0 | 0
|  |  | 28.3 | 27.11 | 1 | -0.146 | 48.006 | 20.082 | 0 | 0
|  |  | 28.35 | 31.799 | 1 | -0.146 | 47.996 | 20.082 | 0 | 0
|  |  | 28.4 | 36.488 | 1 | -0.146 | 47.986 | 20.082 | 0 | 0
|  |  | 28.45 | 37.648 | 1 | -0.146 | 47.976 | 20.082 | 0 | 0
|  |  | 28.5 | 38.808 | 1 | -0.146 | 47.966 | 20.082 | 0 | 0
|  |  | 28.55 | 39.969 | 1 | -0.146 | 47.956 | 20.082 | 0 | 0
|  |  | 28.6 | 41.129 | 1 | -0.146 | 47.946 | 20.082 | 0 | 0
|  |  | 28.65 | 38.55 | 1 | -0.146 | 47.936 | 20.082 | 0 | 0
|  |  | 28.7 | 35.97 | 1 | -0.146 | 47.926 | 20.082 | 0 | 0
|  |  | 28.75 | 33.391 | 1 | -0.146 | 47.916 | 20.082 | 0 | 0
|  |  | 28.8 | 30.812 | 1 | -0.146 | 47.906 | 20.082 | 0 | 0
|  |  | 28.85 | 26.982 | 1 | -0.146 | 47.896 | 20.082 | 0 | 0
|  |  | 28.9 | 23.152 | 1 | -0.146 | 47.886 | 20.082 | 0 | 0
|  |  | 28.95 | 19.322 | 1 | -0.146 | 47.876 | 20.082 | 0 | 0
|  |  | 29 | 15.492 | 1 | -0.146 | 47.866 | 20.082 | 0 | 0
|  |  | 29.05 | 15.284 | 1 | -0.146 | 47.856 | 20.082 | 0 | 0
|  |  | 29.1 | 15.076 | 1 | -0.146 | 47.846 | 20.082 | 0 | 0
|  |  | 29.15 | 14.868 | 1 | -0.146 | 47.836 | 20.082 | 0 | 0
|  |  | 29.2 | 14.66 | 1 | -0.146 | 47.826 | 20.082 | 0 | 0
|  |  | 29.25 | 17.628 | 1 | -0.146 | 47.816 | 20.082 | 0 | 0
|  |  | 29.3 | 20.596 | 1 | -0.146 | 47.806 | 20.082 | 0 | 0
|  |  | 29.35 | 23.564 | 1 | -0.146 | 47.796 | 20.082 | 0 | 0
|  |  | 29.4 | 26.532 | 1 | -0.146 | 47.786 | 20.082 | 0 | 0
|  |  | 29.45 | 30.031 | 1 | -0.146 | 47.776 | 20.082 | 0 | 0
|  |  | 29.5 | 33.53 | 1 | -0.146 | 47.766 | 20.082 | 0 | 0
|  |  | 29.55 | 37.028 | 1 | -0.146 | 47.756 | 20.082 | 0 | 0
|  |  | 29.6 | 40.527 | 1 | -0.146 | 47.746 | 20.082 | 0 | 0
|  |  | 29.65 | 39.958 | 1 | -0.146 | 47.736 | 20.082 | 0 | 0
|  |  | 29.7 | 39.39 | 1 | -0.146 | 47.726 | 20.082 | 0 | 0
|  |  | 29.75 | 38.821 | 1 | -0.146 | 47.716 | 20.082 | 0 | 0
|  |  | 29.8 | 38.253 | 1 | -0.146 | 47.706 | 20.082 | 0 | 0
|  |  | 29.85 | 33.906 | 1 | -0.146 | 47.696 | 20.082 | 0 | 0
|  |  | 29.9 | 29.56 | 1 | -0.146 | 47.686 | 20.082 | 0 | 0
|  |  | 29.95 | 25.213 | 1 | -0.146 | 47.676 | 20.082 | 0 | 0
|  |  | 30 | 20.867 | 1 | -0.146 | 47.666 | 20.082 | 0 | 0
|  |  | 30.05 | 19.061 | 1 | -0.146 | 47.656 | 20.082 | 0 | 0
|  |  | 30.1 | 17.256 | 1 | -0.146 | 47.646 | 20.082 | 0 | 0
|  |  | 30.15 | 15.451 | 1 | -0.146 | 47.636 | 20.082 | 0 | 0
|  |  | 30.2 | 13.646 | 1 | -0.146 | 47.626 | 20.082 | 0 | 0
|  |  | 30.25 | 14.635 | 1 | -0.146 | 47.616 | 20.082 | 0 | 0
|  |  | 30.3 | 15.624 | 1 | -0.146 | 47.606 | 20.082 | 0 | 0
|  |  | 30.35 | 16.614 | 1 | -0.146 | 47.596 | 20.082 | 0 | 0
|  |  | 30.4 | 17.603 | 1 | -0.146 | 47.586 | 20.081 | 0 | 0
|  |  | 30.45 | 22.318 | 1 | -0.146 | 47.576 | 20.081 | 0 | 0
|  |  | 30.5 | 27.033 | 1 | -0.146 | 47.566 | 20.081 | 0 | 0
|  |  | 30.55 | 31.748 | 1 | -0.146 | 47.556 | 20.081 | 0 | 0
|  |  | 30.6 | 36.462 | 1 | -0.146 | 47.546 | 20.081 | 0 | 0
|  |  | 30.65 | 37.586 | 1 | -0.146 | 47.536 | 20.081 | 0 | 0
|  |  | 30.7 | 38.71 | 1 | -0.146 | 47.526 | 20.081 | 0 | 0
|  |  | 30.75 | 39.833 | 1 | -0.146 | 47.516 | 20.081 | 0 | 0
|  |  | 30.8 | 40.957 | 1 | -0.146 | 47.506 | 20.081 | 0 | 0
|  |  | 30.85 | 38.291 | 1 | -0.146 | 47.496 | 20.081 | 0 | 0
|  |  | 30.9 | 35.625 | 1 | -0.146 | 47.486 | 20.081 | 0 | 0
|  |  | 30.95 | 32.958 | 1 | -0.146 | 47.476 | 20.081 | 0 | 0
|  |  | 31 | 30.292 | 1 | -0.146 | 47.466 | 20.081 | 0 | 0
|  |  | 31.05 | 26.437 | 1 | -0.146 | 47.456 | 20.081 | 0 | 0
|  |  | 31.1 | 22.581 | 1 | -0.146 | 47.446 | 20.081 | 0 | 0
|  |  | 31.15 | 18.725 | 1 | -0.146 | 47.436 | 20.081 | 0 | 0
|  |  | 31.2 | 14.869 | 1 | -0.146 | 47.426 | 20.081 | 0 | 0
|  |  | 31.25 | 14.702 | 1 | -0.146 | 47.416 | 20.081 | 0 | 0
|  |  | 31.3 | 14.535 | 1 | -0.146 | 47.406 | 20.081 | 0 | 0
|  |  | 31.35 | 14.368 | 1 | -0.146 | 47.396 | 20.081 | 0 | 0
|  |  | 31.4 | 14.201 | 1 | -0.146 | 47.386 | 20.081 | 0 | 0
|  |  | 31.45 | 17.242 | 1 | -0.146 | 47.376 | 20.081 | 0 | 0
|  |  | 31.5 | 20.282 | 1 | -0.146 | 47.366 | 20.081 | 0 | 0
|  |  | 31.55 | 23.323 | 1 | -0.146 | 47.356 | 20.081 | 0 | 0
|  |  | 31.6 | 26.364 | 1 | -0.146 | 47.346 | 20.081 | 0 | 0
|  |  | 31.65 | 29.849 | 1 | -0.146 | 47.336 | 20.081 | 0 | 0
|  |  | 31.7 | 33.334 | 1 | -0.146 | 47.326 | 20.081 | 0 | 0
|  |  | 31.75 | 36.82 | 1 | -0.146 | 47.316 | 20.081 | 0 | 0
|  |  | 31.8 | 40.305 | 1 | -0.146 | 47.306 | 20.081 | 0 | 0
|  |  | 31.85 | 39.68 | 1 | -0.146 | 47.296 | 20.081 | 0 | 0
|  |  | 31.9 | 39.054 | 1 | -0.146 | 47.286 | 20.081 | 0 | 0
|  |  | 31.95 | 38.429 | 1 | -0.146 | 47.276 | 20.081 | 0 | 0
|  |  | 32 | 37.803 | 1 | -0.146 | 47.266 | 20.081 | 0 | 0
|  |  | 32.05 | 33.356 | 1 | -0.146 | 47.256 | 20.081 | 0 | 0
|  |  | 32.1 | 28.908 | 1 | -0.146 | 47.246 | 20.081 | 0 | 0
|  |  | 32.15 | 24.461 | 1 | -0.146 | 47.236 | 20.081 | 0 | 0
|  |  | 32.2 | 20.013 | 1 | -0.146 | 47.226 | 20.081 | 0 | 0
|  |  | 32.25 | 18.25 | 1 | -0.146 | 47.216 | 20.081 | 0 | 0
|  |  | 32.3 | 16.487 | 1 | -0.146 | 47.206 | 20.081 | 0 | 0
|  |  | 32.35 | 14.723 | 1 | -0.146 | 47.196 | 20.081 | 0 | 0
|  |  | 32.4 | 12.96 | 1 | -0.146 | 47.186 | 20.081 | 0 | 0
|  |  | 32.45 | 14.01 | 1 | -0.146 | 47.176 | 20.081 | 0 | 0
|  |  | 32.5 | 15.06 | 1 | -0.146 | 47.166 | 20.081 | 0 | 0
|  |  | 32.55 | 16.111 | 1 | -0.146 | 47.156 | 20.081 | 0 | 0
|  |  | 32.6 | 17.161 | 1 | -0.146 | 47.146 | 20.081 | 0 | 0
|  |  | 32.65 | 21.905 | 1 | -0.146 | 47.136 | 20.081 | 0 | 0
|  |  | 32.7 | 26.649 | 1 | -0.146 | 47.126 | 20.081 | 0 | 0
|  |  | 32.75 | 31.394 | 1 | -0.146 | 47.116 | 20.081 | 0 | 0
|  |  | 32.8 | 36.138 | 1 | -0.146 | 47.106 | 20.081 | 0 | 0
|  |  | 32.85 | 37.203 | 1 | -0.146 | 47.096 | 20.081 | 0 | 0
|  |  | 32.9 | 38.267 | 1 | -0.146 | 47.086 | 20.081 | 0 | 0
|  |  | 32.95 | 39.332 | 1 | -0.146 | 47.076 | 20.081 | 0 | 0
|  |  | 33 | 40.397 | 1 | -0.146 | 47.066 | 20.081 | 0 | 0
|  |  | 33.05 | 37.626 | 1 | -0.146 | 47.056 | 20.081 | 0 | 0
|  |  | 33.1 | 34.855 | 1 | -0.146 | 47.046 | 20.081 | 0 | 0
|  |  | 33.15 | 32.083 | 1 | -0.146 | 47.036 | 20.081 | 0 | 0
|  |  | 33.2 | 29.312 | 1 | -0.146 | 47.026 | 20.081 | 0 | 0
|  |  | 33.25 | 25.436 | 1 | -0.146 | 47.016 | 20.081 | 0 | 0
|  |  | 33.3 | 21.559 | 1 | -0.146 | 47.006 | 20.081 | 0 | 0
|  |  | 33.35 | 17.682 | 1 | -0.146 | 46.996 | 20.081 | 0 | 0
|  |  | 33.4 | 13.805 | 1 | -0.146 | 46.986 | 20.081 | 0 | 0
|  |  | 33.45 | 13.673 | 1 | -0.146 | 46.976 | 20.081 | 0 | 0
|  |  | 33.5 | 13.541 | 1 | -0.146 | 46.966 | 20.081 | 0 | 0
|  |  | 33.55 | 13.409 | 1 | -0.146 | 46.956 | 20.081 | 0 | 0
|  |  | 33.6 | 13.277 | 1 | -0.146 | 46.946 | 20.081 | 0 | 0
|  |  | 33.65 | 16.385 | 1 | -0.146 | 46.936 | 20.081 | 0 | 0
|  |  | 33.7 | 19.493 | 1 | -0.146 | 46.926 | 20.081 | 0 | 0
|  |  | 33.75 | 22.601 | 1 | -0.146 | 46.916 | 20.081 | 0 | 0
|  |  | 33.8 | 25.709 | 1 | -0.146 | 46.906 | 20.081 | 0 | 0
|  |  | 33.85 | 29.157 | 1 | -0.146 | 46.896 | 20.08 | 0 | 0
|  |  | 33.9 | 32.606 | 1 | -0.146 | 46.886 | 20.08 | 0 | 0
|  |  | 33.95 | 36.054 | 1 | -0.146 | 46.876 | 20.08 | 0 | 0
|  |  | 34 | 39.503 | 1 | -0.146 | 46.866 | 20.08 | 0 | 0
|  |  | 34.05 | 38.782 | 1 | -0.146 | 46.856 | 20.08 | 0 | 0
|  |  | 34.1 | 38.06 | 1 | -0.146 | 46.846 | 20.08 | 0 | 0
|  |  | 34.15 | 37.339 | 1 | -0.146 | 46.836 | 20.08 | 0 | 0
|  |  | 34.2 | 36.617 | 1 | -0.146 | 46.826 | 20.08 | 0 | 0
|  |  | 34.25 | 32.06 | 1 | -0.146 | 46.816 | 20.08 | 0 | 0
|  |  | 34.3 | 27.502 | 1 | -0.146 | 46.806 | 20.08 | 0 | 0
|  |  | 34.35 | 22.945 | 1 | -0.146 | 46.796 | 20.08 | 0 | 0
|  |  | 34.4 | 18.387 | 1 | -0.146 | 46.786 | 20.08 | 0 | 0
|  |  | 34.45 | 16.638 | 1 | -0.146 | 46.776 | 20.08 | 0 | 0
|  |  | 34.5 | 14.889 | 1 | -0.146 | 46.766 | 20.08 | 0 | 0
|  |  | 34.55 | 13.141 | 1 | -0.146 | 46.756 | 20.08 | 0 | 0
|  |  | 34.6 | 11.392 | 1 | -0.146 | 46.746 | 20.08 | 0 | 0
|  |  | 34.65 | 12.476 | 1 | -0.146 | 46.736 | 20.08 | 0 | 0
|  |  | 34.7 | 13.561 | 1 | -0.146 | 46.726 | 20.08 | 0 | 0
|  |  | 34.75 | 14.645 | 1 | -0.146 | 46.716 | 20.08 | 0 | 0
|  |  | 34.8 | 15.729 | 1 | -0.146 | 46.706 | 20.08 | 0 | 0
|  |  | 34.85 | 20.477 | 1 | -0.146 | 46.696 | 20.08 | 0 | 0
|  |  | 34.9 | 25.224 | 1 | -0.146 | 46.686 | 20.08 | 0 | 0
|  |  | 34.95 | 29.972 | 1 | -0.146 | 46.676 | 20.08 | 0.001 | 0.1
|  |  | 35 | 34.719 | 1 | -0.146 | 46.666 | 20.08 | 0.001 | 0.1
|  |  | 35.05 | 35.675 | 1 | -0.146 | 46.656 | 20.08 | 0.001 | 0.1
|  |  | 35.1 | 36.63 | 1 | -0.146 | 46.646 | 20.08 | 0.001 | 0.1
|  |  | 35.15 | 37.586 | 1 | -0.146 | 46.636 | 20.08 | 0.001 | 0.1
|  |  | 35.2 | 38.542 | 1 | -0.146 | 46.626 | 20.08 | 0.001 | 0.1
|  |  | 35.25 | 35.541 | 1 | -0.146 | 46.616 | 20.08 | 0.001 | 0.1
|  |  | 35.3 | 32.541 | 1 | -0.146 | 46.606 | 20.08 | 0.001 | 0.1
|  |  | 35.35 | 29.541 | 1 | -0.146 | 46.596 | 20.08 | 0.001 | 0.1
|  |  | 35.4 | 26.541 | 1 | -0.146 | 46.586 | 20.08 | 0.001 | 0.1
|  |  | 35.45 | 22.374 | 1 | -0.146 | 46.576 | 20.08 | 0.001 | 0.1
|  |  | 35.5 | 18.207 | 1 | -0.146 | 46.566 | 20.08 | 0.001 | 0.1
|  |  | 35.55 | 14.04 | 1 | -0.146 | 46.556 | 20.08 | 0.001 | 0.1
|  |  | 35.6 | 9.8725 | 1 | -0.146 | 46.546 | 20.08 | 0.001 | 0.1
|  |  | 35.65 | 8.97 | 1 | -0.146 | 46.536 | 20.08 | 0.001 | 0.1
|  |  | 35.7 | 8.0674 | 1 | -0.146 | 46.526 | 20.08 | 0.001 | 0.1
|  |  | 35.75 | 7.1649 | 1 | -0.146 | 46.516 | 20.08 | 0.001 | 0.1
|  |  | 35.8 | 6.2623 | 1 | -0.146 | 46.506 | 20.08 | 0.001 | 0.1
|  |  | 35.85 | 6.0211 | 1 | -0.146 | 46.496 | 20.08 | 0.001 | 0.1
|  |  | 35.9 | 5.7798 | 1 | -0.146 | 46.486 | 20.08 | 0.001 | 0.1
|  |  | 35.95 | 5.5386 | 1 | -0.146 | 46.476 | 20.08 | 0.001 | 0.1
|  |  | 36 | 5.2973 | 1 | -0.146 | 46.466 | 20.08 | 0.001 | 0.1
|  |  | 36.05 | 5.0562 | 1 | -0.146 | 46.456 | 20.08 | 0.001 | 0.1
|  |  | 36.1 | 4.8151 | 1 | -0.146 | 46.446 | 20.08 | 0.001 | 0.1
|  |  | 36.15 | 4.574 | 1 | -0.146 | 46.436 | 20.08 | 0.001 | 0.1
|  |  | 36.2 | 4.3329 | 1 | -0.146 | 46.426 | 20.08 | 0.001 | 0.1
|  |  | 36.25 | 4.2673 | 1 | -0.146 | 46.416 | 20.08 | 0.001 | 0.1
|  |  | 36.3 | 4.2016 | 1 | -0.146 | 46.406 | 20.08 | 0.001 | 0.1
|  |  | 36.35 | 4.1359 | 1 | -0.146 | 46.396 | 20.08 | 0.001 | 0.1
|  |  | 36.4 | 4.0702 | 1 | -0.146 | 46.386 | 20.08 | 0.001 | 0.1
|  |  | 36.45 | 4.0046 | 1 | -0.146 | 46.376 | 20.08 | 0.001 | 0.1
|  |  | 36.5 | 3.9389 | 1 | -0.146 | 46.366 | 20.08 | 0.001 | 0.1
|  |  | 36.55 | 3.8733 | 1 | -0.146 | 46.356 | 20.08 | 0.001 | 0.1
|  |  | 36.6 | 3.8077 | 1 | -0.146 | 46.346 | 20.08 | 0.001 | 0.1
|  |  | 36.65 | 3.7421 | 1 | -0.146 | 46.336 | 20.08 | 0.001 | 0.1
|  |  | 36.7 | 3.6765 | 1 | -0.146 | 46.326 | 20.08 | 0.001 | 0.1
|  |  | 36.75 | 3.6109 | 1 | -0.146 | 46.316 | 20.08 | 0.001 | 0.1
|  |  | 36.8 | 3.5453 | 1 | -0.146 | 46.306 | 20.08 | 0.001 | 0.1
|  |  | 36.85 | 3.4797 | 1 | -0.146 | 46.296 | 20.08 | 0.001 | 0.1
|  |  | 36.9 | 3.4142 | 1 | -0.146 | 46.286 | 20.08 | 0.001 | 0.1
|  |  | 36.95 | 3.3486 | 1 | -0.146 | 46.276 | 20.08 | 0.001 | 0.1
|  |  | 37 | 3.2831 | 1 | -0.146 | 46.266 | 20.08 | 0.001 | 0.1
|  |  | 37.05 | 3.2478 | 1 | -0.146 | 46.256 | 20.08 | 0.001 | 0.1
|  |  | 37.1 | 3.2126 | 1 | -0.146 | 46.246 | 20.08 | 0.001 | 0.1
|  |  | 37.15 | 3.1773 | 1 | -0.146 | 46.236 | 20.08 | 0.001 | 0.1
|  |  | 37.2 | 3.142 | 1 | -0.146 | 46.226 | 20.08 | 0.001 | 0.1
|  |  | 37.25 | 3.1068 | 1 | -0.146 | 46.216 | 20.08 | 0.001 | 0.1
|  |  | 37.3 | 3.0716 | 1 | -0.146 | 46.206 | 20.079 | 0.001 | 0.1
|  |  | 37.35 | 3.0363 | 1 | -0.146 | 46.196 | 20.079 | 0.001 | 0.1
|  |  | 37.4 | 3.0011 | 1 | -0.146 | 46.186 | 20.079 | 0.001 | 0.1
|  |  | 37.45 | 2.9659 | 1 | -0.146 | 46.176 | 20.079 | 0.001 | 0.1
|  |  | 37.5 | 2.9307 | 1 | -0.146 | 46.166 | 20.079 | 0.001 | 0.1
|  |  | 37.55 | 2.8954 | 1 | -0.146 | 46.156 | 20.079 | 0.001 | 0.1
|  |  | 37.6 | 2.8602 | 1 | -0.146 | 46.146 | 20.079 | 0.001 | 0.1
|  |  | 37.65 | 2.825 | 1 | -0.146 | 46.136 | 20.079 | 0.001 | 0.1
|  |  | 37.7 | 2.7898 | 1 | -0.146 | 46.126 | 20.079 | 0.001 | 0.1
|  |  | 37.75 | 2.7547 | 1 | -0.146 | 46.116 | 20.079 | 0.001 | 0.1
|  |  | 37.8 | 2.7195 | 1 | -0.146 | 46.106 | 20.079 | 0.001 | 0.1
|  |  | 37.85 | 2.6881 | 1 | -0.146 | 46.096 | 20.079 | 0.001 | 0.1
|  |  | 37.9 | 2.6567 | 1 | -0.146 | 46.086 | 20.079 | 0.001 | 0.1
|  |  | 37.95 | 2.6254 | 1 | -0.146 | 46.076 | 20.079 | 0.001 | 0.1
|  |  | 38 | 2.594 | 1 | -0.146 | 46.066 | 20.079 | 0.001 | 0.1
|  |  | 38.05 | 2.5626 | 1 | -0.146 | 46.056 | 20.079 | 0.001 | 0.1
|  |  | 38.1 | 2.5313 | 1 | -0.146 | 46.046 | 20.079 | 0.001 | 0.1
|  |  | 38.15 | 2.4999 | 1 | -0.146 | 46.036 | 20.079 | 0.001 | 0.1
|  |  | 38.2 | 2.4686 | 1 | -0.146 | 46.026 | 20.079 | 0.001 | 0.1
|  |  | 38.25 | 2.4373 | 1 | -0.146 | 46.016 | 20.079 | 0.001 | 0.1
|  |  | 38.3 | 2.406 | 1 | -0.146 | 46.006 | 20.079 | 0.001 | 0.1
|  |  | 38.35 | 2.3746 | 1 | -0.146 | 45.996 | 20.079 | 0.001 | 0.1
|  |  | 38.4 | 2.3433 | 1 | -0.146 | 45.986 | 20.079 | 0.001 | 0.1
|  |  | 38.45 | 2.312 | 1 | -0.146 | 45.977 | 20.079 | 0.001 | 0.1
|  |  | 38.5 | 2.2807 | 1 | -0.146 | 45.967 | 20.079 | 0.001 | 0.1
|  |  | 38.55 | 2.2494 | 1 | -0.146 | 45.957 | 20.079 | 0.001 | 0.1
|  |  | 38.6 | 2.2181 | 1 | -0.146 | 45.947 | 20.079 | 0.001 | 0.1
|  |  | 38.65 | 2.1959 | 1 | -0.146 | 45.937 | 20.079 | 0.001 | 0.1
|  |  | 38.7 | 2.1737 | 1 | -0.146 | 45.927 | 20.079 | 0.001 | 0.1
|  |  | 38.75 | 2.1515 | 1 | -0.146 | 45.917 | 20.079 | 0.001 | 0.1
|  |  | 38.8 | 2.1293 | 1 | -0.146 | 45.907 | 20.079 | 0.001 | 0.1
|  |  | 38.85 | 2.1071 | 1 | -0.146 | 45.897 | 20.079 | 0.001 | 0.1
|  |  | 38.9 | 2.0849 | 1 | -0.146 | 45.887 | 20.079 | 0.001 | 0.1
|  |  | 38.95 | 2.0627 | 1 | -0.146 | 45.877 | 20.079 | 0.001 | 0.1
|  |  | 39 | 2.0405 | 1 | -0.146 | 45.867 | 20.079 | 0.001 | 0.1
|  |  | 39.05 | 2.0183 | 1 | -0.146 | 45.857 | 20.079 | 0.001 | 0.1
|  |  | 39.1 | 1.9961 | 1 | -0.146 | 45.847 | 20.079 | 0.001 | 0.1
|  |  | 39.15 | 1.9739 | 1 | -0.146 | 45.837 | 20.079 | 0.001 | 0.1
|  |  | 39.2 | 1.9517 | 1 | -0.146 | 45.827 | 20.079 | 0.001 | 0.1
|  |  | 39.25 | 1.9296 | 1 | -0.146 | 45.817 | 20.079 | 0.001 | 0.1
|  |  | 39.3 | 1.9074 | 1 | -0.146 | 45.807 | 20.079 | 0.001 | 0.1
|  |  | 39.35 | 1.8853 | 1 | -0.146 | 45.797 | 20.079 | 0.001 | 0.1
|  |  | 39.4 | 1.8631 | 1 | -0.146 | 45.787 | 20.079 | 0.001 | 0.1
|  |  | 39.45 | 1.8401 | 1 | -0.146 | 45.777 | 20.079 | 0.001 | 0.1
|  |  | 39.5 | 1.817 | 1 | -0.146 | 45.767 | 20.079 | 0.001 | 0.1
|  |  | 39.55 | 1.794 | 1 | -0.146 | 45.757 | 20.079 | 0.001 | 0.1
|  |  | 39.6 | 1.771 | 1 | -0.146 | 45.747 | 20.079 | 0.001 | 0.1
|  |  | 39.65 | 1.7479 | 1 | -0.146 | 45.737 | 20.079 | 0.001 | 0.1
|  |  | 39.7 | 1.7249 | 1 | -0.146 | 45.727 | 20.079 | 0.001 | 0.1
|  |  | 39.75 | 1.7019 | 1 | -0.146 | 45.717 | 20.079 | 0.001 | 0.1
|  |  | 39.8 | 1.6789 | 1 | -0.146 | 45.707 | 20.079 | 0.001 | 0.1
|  |  | 39.85 | 1.6559 | 1 | -0.146 | 45.697 | 20.079 | 0.001 | 0.1
|  |  | 39.9 | 1.6329 | 1 | -0.146 | 45.687 | 20.079 | 0.001 | 0.1
|  |  | 39.95 | 1.6099 | 1 | -0.146 | 45.677 | 20.079 | 0.001 | 0.1
|  |  | 40 | 1.5869 | 1 | -0.146 | 45.667 | 20.079 | 0.001 | 0.1
|  |  | 40.05 | 1.5639 | 1 | -0.146 | 45.657 | 20.079 | 0.001 | 0.1
|  |  | 40.1 | 1.541 | 1 | -0.146 | 45.647 | 20.079 | 0.001 | 0.1
|  |  | 40.15 | 1.518 | 1 | -0.146 | 45.637 | 20.079 | 0.001 | 0.1
|  |  | 40.2 | 1.495 | 1 | -0.146 | 45.627 | 20.079 | 0.001 | 0.1
|  |  | 40.25 | 1.4691 | 1 | -0.146 | 45.617 | 20.079 | 0.001 | 0.1
|  |  | 40.3 | 1.4433 | 1 | -0.146 | 45.607 | 20.079 | 0.001 | 0.1
|  |  | 40.35 | 1.4174 | 1 | -0.146 | 45.597 | 20.079 | 0.001 | 0.1
|  |  | 40.4 | 1.3915 | 1 | -0.146 | 45.587 | 20.079 | 0.001 | 0.1
|  |  | 40.45 | 1.3656 | 1 | -0.146 | 45.577 | 20.079 | 0.001 | 0.1
|  |  | 40.5 | 1.3398 | 1 | -0.146 | 45.567 | 20.079 | 0.001 | 0.1
|  |  | 40.55 | 1.3139 | 1 | -0.146 | 45.557 | 20.079 | 0.001 | 0.1
|  |  | 40.6 | 1.2881 | 1 | -0.146 | 45.547 | 20.079 | 0.001 | 0.1
|  |  | 40.65 | 1.2622 | 1 | -0.146 | 45.537 | 20.079 | 0.001 | 0.1
|  |  | 40.7 | 1.2364 | 1 | -0.146 | 45.527 | 20.079 | 0.002 | 0.2
|  |  | 40.75 | 1.2106 | 1 | -0.146 | 45.517 | 20.078 | 0.002 | 0.2
|  |  | 40.8 | 1.1848 | 1 | -0.146 | 45.507 | 20.078 | 0.002 | 0.2
|  |  | 40.85 | 1.1589 | 1 | -0.146 | 45.497 | 20.078 | 0.002 | 0.2
|  |  | 40.9 | 1.1331 | 1 | -0.146 | 45.487 | 20.078 | 0.002 | 0.2
|  |  | 40.95 | 1.1073 | 1 | -0.146 | 45.477 | 20.078 | 0.002 | 0.2
|  |  | 41 | 1.0815 | 1 | -0.146 | 45.467 | 20.078 | 0.002 | 0.2
|  |  | 41.05 | 1.0711 | 1 | -0.146 | 45.457 | 20.078 | 0.002 | 0.2
|  |  | 41.1 | 1.0607 | 1 | -0.146 | 45.447 | 20.078 | 0.002 | 0.2
|  |  | 41.15 | 1.0503 | 1 | -0.146 | 45.437 | 20.078 | 0.002 | 0.2
|  |  | 41.2 | 1.0399 | 1 | -0.146 | 45.427 | 20.078 | 0.003 | 0.3
|  |  | 41.25 | 1.0295 | 1 | -0.146 | 45.417 | 20.078 | 0.003 | 0.3
|  |  | 41.3 | 1.0191 | 1 | -0.146 | 45.407 | 20.078 | 0.003 | 0.3
|  |  | 41.35 | 1.0087 | 1 | -0.146 | 45.397 | 20.078 | 0.003 | 0.3
|  |  | 41.4 | 0.99827 | 1 | -0.146 | 45.387 | 20.078 | 0.003 | 0.3
|  |  | 41.45 | 0.98788 | 1 | -0.146 | 45.377 | 20.078 | 0.003 | 0.3
|  |  | 41.5 | 0.97748 | 1 | -0.146 | 45.367 | 20.078 | 0.003 | 0.3
|  |  | 41.55 | 0.96709 | 1 | -0.146 | 45.357 | 20.078 | 0.003 | 0.3
|  |  | 41.6 | 0.9567 | 1 | -0.146 | 45.347 | 20.078 | 0.003 | 0.3
|  |  | 41.65 | 0.94631 | 1 | -0.146 | 45.337 | 20.078 | 0.003 | 0.3
|  |  | 41.7 | 0.93593 | 1 | -0.146 | 45.327 | 20.078 | 0.003 | 0.3
|  |  | 41.75 | 0.92555 | 1 | -0.146 | 45.317 | 20.078 | 0.003 | 0.3
|  |  | 41.8 | 0.91517 | 1 | -0.146 | 45.307 | 20.078 | 0.003 | 0.3
|  |  | 41.85 | 0.90689 | 1 | -0.146 | 45.297 | 20.078 | 0.003 | 0.3
|  |  | 41.9 | 0.89861 | 1 | -0.146 | 45.287 | 20.078 | 0.003 | 0.3
|  |  | 41.95 | 0.89033 | 1 | -0.146 | 45.277 | 20.078 | 0.004 | 0.4
|  |  | 42 | 0.88206 | 1 | -0.146 | 45.267 | 20.078 | 0.004 | 0.4
|  |  | 42.05 | 0.87379 | 1 | -0.146 | 45.257 | 20.078 | 0.004 | 0.4
|  |  | 42.1 | 0.86552 | 1 | -0.146 | 45.247 | 20.078 | 0.004 | 0.4
|  |  | 42.15 | 0.85725 | 1 | -0.146 | 45.237 | 20.078 | 0.004 | 0.4
|  |  | 42.2 | 0.84898 | 1 | -0.146 | 45.227 | 20.078 | 0.004 | 0.4
|  |  | 42.25 | 0.84072 | 1 | -0.146 | 45.217 | 20.078 | 0.004 | 0.4
|  |  | 42.3 | 0.83246 | 1 | -0.146 | 45.207 | 20.078 | 0.004 | 0.4
|  |  | 42.35 | 0.8242 | 1 | -0.146 | 45.197 | 20.078 | 0.004 | 0.4
|  |  | 42.4 | 0.81594 | 1 | -0.146 | 45.187 | 20.078 | 0.004 | 0.4
|  |  | 42.45 | 0.80768 | 1 | -0.146 | 45.177 | 20.078 | 0.004 | 0.4
|  |  | 42.5 | 0.79943 | 1 | -0.146 | 45.167 | 20.078 | 0.004 | 0.4
|  |  | 42.55 | 0.79117 | 1 | -0.146 | 45.157 | 20.078 | 0.004 | 0.4
|  |  | 42.6 | 0.78292 | 1 | -0.146 | 45.147 | 20.078 | 0.004 | 0.4
|  |  | 42.65 | 0.77589 | 1 | -0.146 | 45.137 | 20.078 | 0.004 | 0.4
|  |  | 42.7 | 0.76887 | 1 | -0.146 | 45.127 | 20.078 | 0.004 | 0.4
|  |  | 42.75 | 0.76184 | 1 | -0.146 | 45.117 | 20.078 | 0.004 | 0.4
|  |  | 42.8 | 0.75482 | 1 | -0.146 | 45.107 | 20.078 | 0.004 | 0.4
|  |  | 42.85 | 0.7478 | 1 | -0.146 | 45.097 | 20.078 | 0.004 | 0.4
|  |  | 42.9 | 0.74077 | 1 | -0.146 | 45.087 | 20.078 | 0.005 | 0.5
|  |  | 42.95 | 0.73375 | 1 | -0.146 | 45.077 | 20.078 | 0.005 | 0.5
|  |  | 43 | 0.72673 | 1 | -0.146 | 45.067 | 20.078 | 0.005 | 0.5
|  |  | 43.05 | 0.71972 | 1 | -0.146 | 45.057 | 20.078 | 0.005 | 0.5
|  |  | 43.1 | 0.71271 | 1 | -0.146 | 45.047 | 20.078 | 0.005 | 0.5
|  |  | 43.15 | 0.70569 | 1 | -0.146 | 45.037 | 20.078 | 0.005 | 0.5
|  |  | 43.2 | 0.69868 | 1 | -0.146 | 45.027 | 20.078 | 0.005 | 0.5
|  |  | 43.25 | 0.69167 | 1 | -0.146 | 45.017 | 20.078 | 0.005 | 0.5
|  |  | 43.3 | 0.68467 | 1 | -0.146 | 45.007 | 20.078 | 0.005 | 0.5
|  |  | 43.35 | 0.67766 | 1 | -0.146 | 44.997 | 20.078 | 0.005 | 0.5
|  |  | 43.4 | 0.67066 | 1 | -0.146 | 44.987 | 20.078 | 0.005 | 0.5
|  |  | 43.45 | 0.6646 | 1 | -0.146 | 44.977 | 20.078 | 0.005 | 0.5
|  |  | 43.5 | 0.65853 | 1 | -0.146 | 44.967 | 20.078 | 0.005 | 0.5
|  |  | 43.55 | 0.65247 | 1 | -0.146 | 44.957 | 20.078 | 0.005 | 0.5
|  |  | 43.6 | 0.6464 | 1 | -0.146 | 44.947 | 20.078 | 0.005 | 0.5
|  |  | 43.65 | 0.64035 | 1 | -0.146 | 44.937 | 20.078 | 0.005 | 0.5
|  |  | 43.7 | 0.63429 | 1 | -0.146 | 44.927 | 20.078 | 0.005 | 0.5
|  |  | 43.75 | 0.62823 | 1 | -0.146 | 44.917 | 20.078 | 0.005 | 0.5
|  |  | 43.8 | 0.62218 | 1 | -0.146 | 44.907 | 20.078 | 0.005 | 0.5
|  |  | 43.85 | 0.61613 | 1 | -0.146 | 44.897 | 20.078 | 0.005 | 0.5
|  |  | 43.9 | 0.61008 | 1 | -0.146 | 44.887 | 20.078 | 0.005 | 0.5
|  |  | 43.95 | 0.60403 | 1 | -0.146 | 44.877 | 20.078 | 0.005 | 0.5
|  |  | 44 | 0.59797 | 1 | -0.146 | 44.867 | 20.078 | 0.005 | 0.5
|  |  | 44.05 | 0.59193 | 1 | -0.146 | 44.857 | 20.078 | 0.005 | 0.5
|  |  | 44.1 | 0.58589 | 1 | -0.146 | 44.847 | 20.078 | 0.005 | 0.5
|  |  | 44.15 | 0.57984 | 1 | -0.146 | 44.837 | 20.078 | 0.005 | 0.5
|  |  | 44.2 | 0.5738 | 1 | -0.146 | 44.827 | 20.077 | 0.005 | 0.5
|  |  | 44.25 | 0.56929 | 1 | -0.146 | 44.817 | 20.077 | 0.005 | 0.5
|  |  | 44.3 | 0.56479 | 1 | -0.146 | 44.807 | 20.077 | 0.005 | 0.5
|  |  | 44.35 | 0.56029 | 1 | -0.146 | 44.797 | 20.077 | 0.005 | 0.5
|  |  | 44.4 | 0.55579 | 1 | -0.146 | 44.787 | 20.077 | 0.005 | 0.5
|  |  | 44.45 | 0.55129 | 1 | -0.146 | 44.777 | 20.077 | 0.005 | 0.5
|  |  | 44.5 | 0.54679 | 1 | -0.146 | 44.767 | 20.077 | 0.005 | 0.5
|  |  | 44.55 | 0.54229 | 1 | -0.146 | 44.757 | 20.077 | 0.005 | 0.5
|  |  | 44.6 | 0.53779 | 1 | -0.146 | 44.747 | 20.077 | 0.005 | 0.5
|  |  | 44.65 | 0.5333 | 1 | -0.146 | 44.737 | 20.077 | 0.005 | 0.5
|  |  | 44.7 | 0.5288 | 1 | -0.146 | 44.727 | 20.077 | 0.005 | 0.5
|  |  | 44.75 | 0.52431 | 1 | -0.146 | 44.717 | 20.077 | 0.005 | 0.5
|  |  | 44.8 | 0.51982 | 1 | -0.146 | 44.707 | 20.077 | 0.005 | 0.5
|  |  | 44.85 | 0.51533 | 1 | -0.146 | 44.697 | 20.077 | 0.005 | 0.5
|  |  | 44.9 | 0.51084 | 1 | -0.146 | 44.687 | 20.077 | 0.005 | 0.5
|  |  | 44.95 | 0.50635 | 1 | -0.146 | 44.677 | 20.077 | 0.005 | 0.5
|  |  | 45 | 0.50186 | 1 | -0.146 | 44.667 | 20.077 | 0.005 | 0.5
|  |  | 45.05 | 0.47491 | 1 | -0.146 | 44.657 | 20.077 | 0.005 | 0.5
|  |  | 45.1 | 0.44796 | 1 | -0.146 | 44.647 | 20.077 | 0.005 | 0.5
|  |  | 45.15 | 0.42101 | 1 | -0.146 | 44.637 | 20.077 | 0.005 | 0.5
|  |  | 45.2 | 0.39406 | 1 | -0.146 | 44.627 | 20.077 | 0.005 | 0.5
|  |  | 45.25 | 0.36715 | 1 | -0.146 | 44.617 | 20.077 | 0.005 | 0.5
|  |  | 45.3 | 0.34023 | 1 | -0.146 | 44.607 | 20.077 | 0.005 | 0.5
|  |  | 45.35 | 0.31332 | 1 | -0.146 | 44.597 | 20.077 | 0.005 | 0.5
|  |  | 45.4 | 0.28641 | 1 | -0.146 | 44.587 | 20.077 | 0.005 | 0.5
|  |  | 45.45 | 0.23402 | 1 | -0.146 | 44.577 | 20.077 | 0.005 | 0.5
|  |  | 45.5 | 0.18164 | 1 | -0.146 | 44.567 | 20.077 | 0.005 | 0.5
|  |  | 45.55 | 0.12926 | 1 | -0.146 | 44.557 | 20.077 | 0.005 | 0.5
|  |  | 45.6 | 0.076883 | 1 | -0.146 | 44.547 | 20.077 | 0.005 | 0.5
|  |  | 45.65 | 0.0755 | 1 | -0.146 | 44.537 | 20.077 | 0.005 | 0.5
|  |  | 45.7 | 0.074118 | 1 | -0.146 | 44.527 | 20.077 | 0.005 | 0.5
|  |  | 45.75 | 0.072735 | 1 | -0.146 | 44.517 | 20.077 | 0.005 | 0.5
|  |  | 45.8 | 0.071352 | 1 | -0.146 | 44.507 | 20.077 | 0.005 | 0.5
|  |  | 45.85 | 0.070208 | 1 | -0.146 | 44.497 | 20.077 | 0.005 | 0.5
|  |  | 45.9 | 0.069064 | 1 | -0.146 | 44.487 | 20.077 | 0.005 | 0.5
|  |  | 45.95 | 0.067919 | 1 | -0.146 | 44.477 | 20.077 | 0.005 | 0.5
|  |  | 46 | 0.066775 | 1 | -0.146 | 44.467 | 20.077 | 0.005 | 0.5
|  |  | 46.05 | 0.064972 | 1 | -0.146 | 44.457 | 20.077 | 0.005 | 0.5
|  |  | 46.1 | 0.063169 | 1 | -0.146 | 44.447 | 20.077 | 0.006 | 0.6
|  |  | 46.15 | 0.061365 | 1 | -0.146 | 44.437 | 20.077 | 0.006 | 0.6
|  |  | 46.2 | 0.059562 | 1 | -0.146 | 44.427 | 20.077 | 0.006 | 0.6
|  |  | 46.25 | 0.058895 | 1 | -0.146 | 44.417 | 20.077 | 0.006 | 0.6
|  |  | 46.3 | 0.058228 | 1 | -0.146 | 44.407 | 20.077 | 0.006 | 0.6
|  |  | 46.35 | 0.057561 | 1 | -0.146 | 44.397 | 20.077 | 0.006 | 0.6
|  |  | 46.4 | 0.056894 | 1 | -0.146 | 44.387 | 20.077 | 0.006 | 0.6
|  |  | 46.45 | 0.055599 | 1 | -0.146 | 44.377 | 20.077 | 0.006 | 0.6
|  |  | 46.5 | 0.054304 | 1 | -0.146 | 44.367 | 20.077 | 0.006 | 0.6
|  |  | 46.55 | 0.053009 | 1 | -0.146 | 44.357 | 20.077 | 0.006 | 0.6
|  |  | 46.6 | 0.051713 | 1 | -0.146 | 44.347 | 20.077 | 0.006 | 0.6
|  |  | 46.65 | 0.050853 | 1 | -0.146 | 44.337 | 20.077 | 0.006 | 0.6
|  |  | 46.7 | 0.049993 | 1 | -0.146 | 44.327 | 20.077 | 0.006 | 0.6
|  |  | 46.75 | 0.049133 | 1 | -0.146 | 44.317 | 20.077 | 0.006 | 0.6
|  |  | 46.8 | 0.048273 | 1 | -0.146 | 44.307 | 20.077 | 0.006 | 0.6
|  |  | 46.85 | 0.047667 | 1 | -0.146 | 44.297 | 20.077 | 0.006 | 0.6
|  |  | 46.9 | 0.047062 | 1 | -0.146 | 44.287 | 20.077 | 0.006 | 0.6
|  |  | 46.95 | 0.046457 | 1 | -0.146 | 44.277 | 20.077 | 0.006 | 0.6
|  |  | 47 | 0.045851 | 1 | -0.146 | 44.267 | 20.077 | 0.006 | 0.6
|  |  | 47.05 | 0.044771 | 1 | -0.146 | 44.257 | 20.077 | 0.006 | 0.6
|  |  | 47.1 | 0.043691 | 1 | -0.146 | 44.247 | 20.077 | 0.006 | 0.6
|  |  | 47.15 | 0.042611 | 1 | -0.146 | 44.237 | 20.077 | 0.006 | 0.6
|  |  | 47.2 | 0.041531 | 1 | -0.146 | 44.227 | 20.077 | 0.006 | 0.6
|  |  | 47.25 | 0.041121 | 1 | -0.146 | 44.217 | 20.077 | 0.006 | 0.6
|  |  | 47.3 | 0.040711 | 1 | -0.146 | 44.207 | 20.077 | 0.006 | 0.6
|  |  | 47.35 | 0.040302 | 1 | -0.146 | 44.197 | 20.077 | 0.006 | 0.6
|  |  | 47.4 | 0.039892 | 1 | -0.146 | 44.187 | 20.077 | 0.006 | 0.6
|  |  | 47.45 | 0.039164 | 1 | -0.146 | 44.177 | 20.077 | 0.006 | 0.6
|  |  | 47.5 | 0.038436 | 1 | -0.146 | 44.167 | 20.077 | 0.006 | 0.6
|  |  | 47.55 | 0.037708 | 1 | -0.146 | 44.157 | 20.077 | 0.006 | 0.6
|  |  | 47.6 | 0.03698 | 1 | -0.146 | 44.147 | 20.077 | 0.006 | 0.6
|  |  | 47.65 | 0.036405 | 1 | -0.146 | 44.137 | 20.076 | 0.006 | 0.6
|  |  | 47.7 | 0.035831 | 1 | -0.146 | 44.127 | 20.076 | 0.006 | 0.6
|  |  | 47.75 | 0.035256 | 1 | -0.146 | 44.117 | 20.076 | 0.006 | 0.6
|  |  | 47.8 | 0.034681 | 1 | -0.146 | 44.107 | 20.076 | 0.006 | 0.6
|  |  | 47.85 | 0.034331 | 1 | -0.146 | 44.097 | 20.076 | 0.006 | 0.6
|  |  | 47.9 | 0.03398 | 1 | -0.146 | 44.087 | 20.076 | 0.006 | 0.6
|  |  | 47.95 | 0.033629 | 1 | -0.146 | 44.077 | 20.076 | 0.006 | 0.6
|  |  | 48 | 0.033278 | 1 | -0.146 | 44.067 | 20.076 | 0.006 | 0.6
|  |  | 48.05 | 0.032618 | 1 | -0.146 | 44.057 | 20.076 | 0.006 | 0.6
|  |  | 48.1 | 0.031957 | 1 | -0.146 | 44.047 | 20.076 | 0.006 | 0.6
|  |  | 48.15 | 0.031296 | 1 | -0.146 | 44.037 | 20.076 | 0.006 | 0.6
|  |  | 48.2 | 0.030635 | 1 | -0.146 | 44.027 | 20.076 | 0.006 | 0.6
|  |  | 48.25 | 0.030365 | 1 | -0.146 | 44.017 | 20.076 | 0.006 | 0.6
|  |  | 48.3 | 0.030096 | 1 | -0.146 | 44.007 | 20.076 | 0.006 | 0.6
|  |  | 48.35 | 0.029826 | 1 | -0.146 | 43.997 | 20.076 | 0.006 | 0.6
|  |  | 48.4 | 0.029556 | 1 | -0.146 | 43.987 | 20.076 | 0.006 | 0.6
|  |  | 48.45 | 0.029138 | 1 | -0.146 | 43.977 | 20.076 | 0.006 | 0.6
|  |  | 48.5 | 0.028719 | 1 | -0.146 | 43.967 | 20.076 | 0.006 | 0.6
|  |  | 48.55 | 0.0283 | 1 | -0.146 | 43.957 | 20.076 | 0.006 | 0.6
|  |  | 48.6 | 0.027882 | 1 | -0.146 | 43.947 | 20.076 | 0.006 | 0.6
|  |  | 48.65 | 0.027494 | 1 | -0.146 | 43.937 | 20.076 | 0.006 | 0.6
|  |  | 48.7 | 0.027107 | 1 | -0.146 | 43.927 | 20.076 | 0.006 | 0.6
|  |  | 48.75 | 0.026719 | 1 | -0.146 | 43.917 | 20.076 | 0.006 | 0.6
|  |  | 48.8 | 0.026332 | 1 | -0.146 | 43.907 | 20.076 | 0.006 | 0.6
|  |  | 48.85 | 0.026126 | 1 | -0.146 | 43.897 | 20.076 | 0.006 | 0.6
|  |  | 48.9 | 0.025919 | 1 | -0.146 | 43.887 | 20.076 | 0.006 | 0.6
|  |  | 48.95 | 0.025712 | 1 | -0.146 | 43.877 | 20.076 | 0.006 | 0.6
|  |  | 49 | 0.025506 | 1 | -0.146 | 43.867 | 20.076 | 0.006 | 0.6
|  |  | 49.05 | 0.025083 | 1 | -0.146 | 43.857 | 20.076 | 0.006 | 0.6
|  |  | 49.1 | 0.024661 | 1 | -0.146 | 43.847 | 20.076 | 0.006 | 0.6
|  |  | 49.15 | 0.024238 | 1 | -0.146 | 43.837 | 20.076 | 0.006 | 0.6
|  |  | 49.2 | 0.023816 | 1 | -0.146 | 43.827 | 20.076 | 0.006 | 0.6
|  |  | 49.25 | 0.023625 | 1 | -0.146 | 43.817 | 20.076 | 0.006 | 0.6
|  |  | 49.3 | 0.023434 | 1 | -0.146 | 43.807 | 20.076 | 0.006 | 0.6
|  |  | 49.35 | 0.023243 | 1 | -0.146 | 43.797 | 20.076 | 0.006 | 0.6
|  |  | 49.4 | 0.023052 | 1 | -0.146 | 43.787 | 20.076 | 0.006 | 0.6
|  |  | 49.45 | 0.022792 | 1 | -0.146 | 43.777 | 20.076 | 0.006 | 0.6
|  |  | 49.5 | 0.022533 | 1 | -0.146 | 43.767 | 20.076 | 0.006 | 0.6
|  |  | 49.55 | 0.022273 | 1 | -0.146 | 43.757 | 20.076 | 0.006 | 0.6
|  |  | 49.6 | 0.022014 | 1 | -0.146 | 43.747 | 20.076 | 0.006 | 0.6
|  |  | 49.65 | 0.021723 | 1 | -0.146 | 43.737 | 20.076 | 0.006 | 0.6
|  |  | 49.7 | 0.021432 | 1 | -0.146 | 43.727 | 20.076 | 0.006 | 0.6
|  |  | 49.75 | 0.02114 | 1 | -0.146 | 43.717 | 20.076 | 0.006 | 0.6
|  |  | 49.8 | 0.020849 | 1 | -0.146 | 43.707 | 20.076 | 0.006 | 0.6
|  |  | 49.85 | 0.020708 | 1 | -0.146 | 43.697 | 20.076 | 0.006 | 0.6
|  |  | 49.9 | 0.020567 | 1 | -0.146 | 43.687 | 20.076 | 0.006 | 0.6
|  |  | 49.95 | 0.020426 | 1 | -0.146 | 43.677 | 20.076 | 0.006 | 0.6
|  |  | 50 | 0.020285 | 1 | -0.146 | 43.667 | 20.076 | 0.006 | 0.6
|  |  | 50.05 | 0.019985 | 1 | -0.146 | 43.657 | 20.076 | 0.006 | 0.6
|  |  | 50.1 | 0.019685 | 1 | -0.146 | 43.647 | 20.076 | 0.006 | 0.6
|  |  | 50.15 | 0.019385 | 1 | -0.146 | 43.637 | 20.076 | 0.006 | 0.6
|  |  | 50.2 | 0.019085 | 1 | -0.146 | 43.627 | 20.076 | 0.006 | 0.6
|  |  | 50.25 | 0.018933 | 1 | -0.146 | 43.617 | 20.076 | 0.006 | 0.6
|  |  | 50.3 | 0.018782 | 1 | -0.146 | 43.607 | 20.076 | 0.006 | 0.6
|  |  | 50.35 | 0.01863 | 1 | -0.146 | 43.597 | 20.076 | 0.006 | 0.6
|  |  | 50.4 | 0.018478 | 1 | -0.146 | 43.587 | 20.076 | 0.006 | 0.6
|  |  | 50.45 | 0.018302 | 1 | -0.146 | 43.577 | 20.076 | 0.006 | 0.6
|  |  | 50.5 | 0.018127 | 1 | -0.146 | 43.567 | 20.076 | 0.006 | 0.6
|  |  | 50.55 | 0.017951 | 1 | -0.146 | 43.557 | 20.076 | 0.007 | 0.7
|  |  | 50.6 | 0.017776 | 1 | -0.146 | 43.547 | 20.076 | 0.007 | 0.7
|  |  | 50.65 | 0.017542 | 1 | -0.146 | 43.537 | 20.076 | 0.007 | 0.7
|  |  | 50.7 | 0.017308 | 1 | -0.146 | 43.527 | 20.076 | 0.007 | 0.7
|  |  | 50.75 | 0.017074 | 1 | -0.146 | 43.517 | 20.076 | 0.007 | 0.7
|  |  | 50.8 | 0.01684 | 1 | -0.146 | 43.507 | 20.076 | 0.007 | 0.7
|  |  | 50.85 | 0.016736 | 1 | -0.146 | 43.497 | 20.076 | 0.007 | 0.7
|  |  | 50.9 | 0.016632 | 1 | -0.146 | 43.487 | 20.076 | 0.007 | 0.7
|  |  | 50.95 | 0.016528 | 1 | -0.146 | 43.477 | 20.076 | 0.007 | 0.7
|  |  | 51 | 0.016424 | 1 | -0.146 | 43.467 | 20.076 | 0.007 | 0.7
|  |  | 51.05 | 0.0162 | 1 | -0.146 | 43.457 | 20.076 | 0.007 | 0.7
|  |  | 51.1 | 0.015976 | 1 | -0.146 | 43.447 | 20.075 | 0.007 | 0.7
|  |  | 51.15 | 0.015751 | 1 | -0.146 | 43.437 | 20.075 | 0.007 | 0.7
|  |  | 51.2 | 0.015527 | 1 | -0.146 | 43.427 | 20.075 | 0.007 | 0.7
|  |  | 51.25 | 0.015396 | 1 | -0.146 | 43.417 | 20.075 | 0.007 | 0.7
|  |  | 51.3 | 0.015265 | 1 | -0.146 | 43.407 | 20.075 | 0.007 | 0.7
|  |  | 51.35 | 0.015134 | 1 | -0.146 | 43.397 | 20.075 | 0.007 | 0.7
|  |  | 51.4 | 0.015004 | 1 | -0.146 | 43.387 | 20.075 | 0.007 | 0.7
|  |  | 51.45 | 0.014874 | 1 | -0.146 | 43.377 | 20.075 | 0.007 | 0.7
|  |  | 51.5 | 0.014744 | 1 | -0.146 | 43.367 | 20.075 | 0.007 | 0.7
|  |  | 51.55 | 0.014614 | 1 | -0.146 | 43.357 | 20.075 | 0.007 | 0.7
|  |  | 51.6 | 0.014485 | 1 | -0.146 | 43.347 | 20.075 | 0.007 | 0.7
|  |  | 51.65 | 0.014288 | 1 | -0.146 | 43.337 | 20.075 | 0.007 | 0.7
|  |  | 51.7 | 0.014092 | 1 | -0.146 | 43.327 | 20.075 | 0.007 | 0.7
|  |  | 51.75 | 0.013896 | 1 | -0.146 | 43.317 | 20.075 | 0.007 | 0.7
|  |  | 51.8 | 0.013699 | 1 | -0.146 | 43.307 | 20.075 | 0.007 | 0.7
|  |  | 51.85 | 0.013615 | 1 | -0.146 | 43.297 | 20.075 | 0.007 | 0.7
|  |  | 51.9 | 0.013531 | 1 | -0.146 | 43.287 | 20.075 | 0.007 | 0.7
|  |  | 51.95 | 0.013447 | 1 | -0.146 | 43.277 | 20.075 | 0.007 | 0.7
|  |  | 52 | 0.013363 | 1 | -0.146 | 43.267 | 20.075 | 0.007 | 0.7
|  |  | 52.05 | 0.013186 | 1 | -0.146 | 43.257 | 20.075 | 0.007 | 0.7
|  |  | 52.1 | 0.013009 | 1 | -0.146 | 43.247 | 20.075 | 0.007 | 0.7
|  |  | 52.15 | 0.012832 | 1 | -0.146 | 43.237 | 20.075 | 0.007 | 0.7
|  |  | 52.2 | 0.012655 | 1 | -0.146 | 43.227 | 20.075 | 0.007 | 0.7
|  |  | 52.25 | 0.012538 | 1 | -0.146 | 43.217 | 20.075 | 0.007 | 0.7
|  |  | 52.3 | 0.012422 | 1 | -0.146 | 43.207 | 20.075 | 0.007 | 0.7
|  |  | 52.35 | 0.012305 | 1 | -0.146 | 43.197 | 20.075 | 0.007 | 0.7
|  |  | 52.4 | 0.012189 | 1 | -0.146 | 43.187 | 20.075 | 0.007 | 0.7
|  |  | 52.45 | 0.012088 | 1 | -0.146 | 43.177 | 20.075 | 0.007 | 0.7
|  |  | 52.5 | 0.011988 | 1 | -0.146 | 43.167 | 20.075 | 0.007 | 0.7
|  |  | 52.55 | 0.011888 | 1 | -0.146 | 43.157 | 20.075 | 0.007 | 0.7
|  |  | 52.6 | 0.011787 | 1 | -0.146 | 43.147 | 20.075 | 0.007 | 0.7
|  |  | 52.65 | 0.011619 | 1 | -0.146 | 43.137 | 20.075 | 0.007 | 0.7
|  |  | 52.7 | 0.011452 | 1 | -0.146 | 43.127 | 20.075 | 0.007 | 0.7
|  |  | 52.75 | 0.011284 | 1 | -0.146 | 43.117 | 20.075 | 0.007 | 0.7
|  |  | 52.8 | 0.011116 | 1 | -0.146 | 43.107 | 20.075 | 0.007 | 0.7
|  |  | 52.85 | 0.011044 | 1 | -0.146 | 43.097 | 20.075 | 0.007 | 0.7
|  |  | 52.9 | 0.010971 | 1 | -0.146 | 43.087 | 20.075 | 0.007 | 0.7
|  |  | 52.95 | 0.010899 | 1 | -0.146 | 43.077 | 20.075 | 0.007 | 0.7
|  |  | 53 | 0.010826 | 1 | -0.146 | 43.067 | 20.075 | 0.007 | 0.7
|  |  | 53.05 | 0.010686 | 1 | -0.146 | 43.057 | 20.075 | 0.007 | 0.7
|  |  | 53.1 | 0.010547 | 1 | -0.146 | 43.047 | 20.075 | 0.007 | 0.7
|  |  | 53.15 | 0.010407 | 1 | -0.146 | 43.037 | 20.075 | 0.007 | 0.7
|  |  | 53.2 | 0.010267 | 1 | -0.146 | 43.027 | 20.075 | 0.007 | 0.7
|  |  | 53.25 | 0.010161 | 1 | -0.146 | 43.017 | 20.075 | 0.007 | 0.7
|  |  | 53.3 | 0.010054 | 1 | -0.146 | 43.007 | 20.075 | 0.007 | 0.7
|  |  | 53.35 | 0.0099476 | 1 | -0.146 | 42.997 | 20.075 | 0.007 | 0.7
|  |  | 53.4 | 0.009841 | 1 | -0.146 | 42.987 | 20.075 | 0.007 | 0.7
|  |  | 53.45 | 0.0097631 | 1 | -0.146 | 42.978 | 20.075 | 0.007 | 0.7
|  |  | 53.5 | 0.0096852 | 1 | -0.146 | 42.968 | 20.075 | 0.007 | 0.7
|  |  | 53.55 | 0.0096073 | 1 | -0.146 | 42.958 | 20.075 | 0.007 | 0.7
|  |  | 53.6 | 0.0095295 | 1 | -0.146 | 42.948 | 20.075 | 0.007 | 0.7
|  |  | 53.65 | 0.0093872 | 1 | -0.146 | 42.938 | 20.075 | 0.007 | 0.7
|  |  | 53.7 | 0.009245 | 1 | -0.146 | 42.928 | 20.075 | 0.007 | 0.7
|  |  | 53.75 | 0.0091028 | 1 | -0.146 | 42.918 | 20.075 | 0.007 | 0.7
|  |  | 53.8 | 0.0089606 | 1 | -0.146 | 42.908 | 20.075 | 0.007 | 0.7
|  |  | 53.85 | 0.0088971 | 1 | -0.146 | 42.898 | 20.075 | 0.007 | 0.7
|  |  | 53.9 | 0.0088337 | 1 | -0.146 | 42.888 | 20.075 | 0.007 | 0.7
|  |  | 53.95 | 0.0087702 | 1 | -0.146 | 42.878 | 20.075 | 0.007 | 0.7
|  |  | 54 | 0.0087068 | 1 | -0.146 | 42.868 | 20.075 | 0.007 | 0.7
|  |  | 54.05 | 0.008597 | 1 | -0.146 | 42.858 | 20.075 | 0.007 | 0.7
|  |  | 54.1 | 0.0084873 | 1 | -0.146 | 42.848 | 20.075 | 0.007 | 0.7
|  |  | 54.15 | 0.0083775 | 1 | -0.146 | 42.838 | 20.075 | 0.007 | 0.7
|  |  | 54.2 | 0.0082678 | 1 | -0.146 | 42.828 | 20.075 | 0.008 | 0.8
|  |  | 54.25 | 0.0081715 | 1 | -0.146 | 42.818 | 20.075 | 0.008 | 0.8
|  |  | 54.3 | 0.0080753 | 1 | -0.146 | 42.808 | 20.075 | 0.008 | 0.8
|  |  | 54.35 | 0.0079791 | 1 | -0.146 | 42.798 | 20.075 | 0.008 | 0.8
|  |  | 54.4 | 0.0078829 | 1 | -0.146 | 42.788 | 20.075 | 0.008 | 0.8
|  |  | 54.45 | 0.0078205 | 1 | -0.146 | 42.778 | 20.075 | 0.008 | 0.8
|  |  | 54.5 | 0.0077581 | 1 | -0.146 | 42.768 | 20.075 | 0.008 | 0.8
|  |  | 54.55 | 0.0076957 | 1 | -0.146 | 42.758 | 20.074 | 0.008 | 0.8
|  |  | 54.6 | 0.0076334 | 1 | -0.146 | 42.748 | 20.074 | 0.008 | 0.8
|  |  | 54.65 | 0.0075144 | 1 | -0.146 | 42.738 | 20.074 | 0.008 | 0.8
|  |  | 54.7 | 0.0073954 | 1 | -0.146 | 42.728 | 20.074 | 0.008 | 0.8
|  |  | 54.75 | 0.0072764 | 1 | -0.146 | 42.718 | 20.074 | 0.008 | 0.8
|  |  | 54.8 | 0.0071574 | 1 | -0.146 | 42.708 | 20.074 | 0.008 | 0.8
|  |  | 54.85 | 0.0071259 | 1 | -0.146 | 42.698 | 20.074 | 0.008 | 0.8
|  |  | 54.9 | 0.0070944 | 1 | -0.146 | 42.688 | 20.074 | 0.008 | 0.8
|  |  | 54.95 | 0.0070629 | 1 | -0.146 | 42.678 | 20.074 | 0.008 | 0.8
|  | END_DATA |  |  |  | -0.146 | 42.668 | 20.074 | 0.008 | 0.8
| END_SCAN 1 |  |  |  |  | -0.146 | 42.658 | 20.074 | 0.008 | 0.8
END_SCAN_DATA | [5 empty] | -0.146 | 42.648 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.638 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.628 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.618 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.608 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.598 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.588 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.578 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.568 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.558 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.548 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.538 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.528 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.518 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.508 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.498 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.488 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.478 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.468 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.458 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.448 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.438 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.428 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.418 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.408 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.398 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.388 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.378 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.368 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.358 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.348 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.338 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.328 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.318 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.308 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.298 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.288 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.278 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.268 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.258 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.248 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.238 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.228 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.218 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.208 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.198 | 20.074 | 0.008 | 0.8
[6 empty] | -0.146 | 42.188 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.178 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.168 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.158 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.148 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.138 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.128 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.118 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.108 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.098 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.088 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.078 | 20.074 | 0.009 | 0.9
[6 empty] | -0.146 | 42.068 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.058 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.048 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.038 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.028 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.018 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 42.008 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.998 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.988 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.978 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.968 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.958 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.948 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.938 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.928 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.918 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.908 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.898 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.888 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.878 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.868 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.858 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.848 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.838 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.828 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.818 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.808 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.798 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.788 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.778 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.768 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.758 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.748 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.738 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.728 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.718 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.708 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.698 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.688 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.678 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.668 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.658 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.648 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.638 | 20.073 | 0.009 | 0.9
[6 empty] | -0.146 | 41.628 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.618 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.608 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.598 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.588 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.578 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.568 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.558 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.548 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.538 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.528 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.518 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.508 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.498 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.488 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.478 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.468 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.458 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.448 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.438 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.428 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.418 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.408 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.398 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.388 | 20.073 | 0.01 | 1
[6 empty] | -0.146 | 41.378 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.368 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.358 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.348 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.338 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.328 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.318 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.308 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.298 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.288 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.278 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.268 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.258 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.248 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.238 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.228 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.218 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.208 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.198 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.188 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.178 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.168 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.158 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.148 | 20.072 | 0.01 | 1
[6 empty] | -0.146 | 41.138 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.128 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.118 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.108 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.098 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.088 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.078 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.068 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.058 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.048 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.038 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.028 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.018 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 41.008 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.998 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.988 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.978 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.968 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.958 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.948 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.938 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.928 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.918 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.908 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.898 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.888 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.878 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.868 | 20.072 | 0.011 | 1.1
[6 empty] | -0.146 | 40.858 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.848 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.838 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.828 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.818 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.808 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.798 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.788 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.778 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.768 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.758 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.748 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.738 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.728 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.718 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.708 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.698 | 20.072 | 0.012 | 1.2
[6 empty] | -0.146 | 40.688 | 20.071 | 0.012 | 1.2
[6 empty] | -0.146 | 40.678 | 20.071 | 0.012 | 1.2
[6 empty] | -0.146 | 40.668 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.658 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.648 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.638 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.628 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.618 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.608 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.598 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.588 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.578 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.568 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.558 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.548 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.538 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.528 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.518 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.508 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.498 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.488 | 20.071 | 0.013 | 1.3
[6 empty] | -0.146 | 40.478 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.468 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.458 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.448 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.438 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.428 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.418 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.408 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.398 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.388 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.378 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.368 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.358 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.348 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.338 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.328 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.318 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.308 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.298 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.288 | 20.071 | 0.014 | 1.4
[6 empty] | -0.146 | 40.278 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.268 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.258 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.248 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.238 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.228 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.218 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.208 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.198 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.188 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.178 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.168 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.158 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.148 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.138 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.128 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.118 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.108 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.098 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.088 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.078 | 20.071 | 0.015 | 1.5
[6 empty] | -0.146 | 40.068 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.058 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.048 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.038 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.028 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.018 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 40.008 | 20.071 | 0.016 | 1.6
[6 empty] | -0.146 | 39.998 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.988 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.978 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.969 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.959 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.949 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.939 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.929 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.919 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.909 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.899 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.889 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.879 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.869 | 20.07 | 0.016 | 1.6
[6 empty] | -0.146 | 39.859 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.849 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.839 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.829 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.819 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.809 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.799 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.789 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.779 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.769 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.759 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.749 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.739 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.729 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.719 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.709 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.699 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.689 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.679 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.669 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.659 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.649 | 20.07 | 0.017 | 1.7
[6 empty] | -0.146 | 39.639 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.629 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.619 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.609 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.599 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.589 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.579 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.569 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.559 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.549 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.539 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.529 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.519 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.509 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.499 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.489 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.479 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.469 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.459 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.449 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.439 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.429 | 20.07 | 0.018 | 1.8
[6 empty] | -0.146 | 39.419 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.409 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.399 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.389 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.379 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.369 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.359 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.349 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.339 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.329 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.319 | 20.07 | 0.019 | 1.9
[6 empty] | -0.146 | 39.309 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.299 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.289 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.279 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.269 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.259 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.249 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.239 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.229 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.219 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.209 | 20.069 | 0.019 | 1.9
[6 empty] | -0.146 | 39.199 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.189 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.179 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.169 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.159 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.149 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.139 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.129 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.119 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.109 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.099 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.089 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.079 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.069 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.059 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.049 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.039 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.029 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.019 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 39.009 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 38.999 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 38.989 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 38.979 | 20.069 | 0.02 | 2
[6 empty] | -0.146 | 38.969 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.959 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.949 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.939 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.929 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.919 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.909 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.899 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.889 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.879 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.869 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.859 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.849 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.839 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.829 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.819 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.809 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.799 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.789 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.779 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.769 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.759 | 20.069 | 0.021 | 2.1
[6 empty] | -0.146 | 38.749 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.739 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.729 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.719 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.709 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.699 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.689 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.679 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.669 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.659 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.649 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.639 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.629 | 20.069 | 0.022 | 2.2
[6 empty] | -0.146 | 38.619 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.609 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.599 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.589 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.579 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.569 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.559 | 20.068 | 0.022 | 2.2
[6 empty] | -0.146 | 38.549 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.539 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.529 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.519 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.509 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.499 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.489 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.479 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.469 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.459 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.449 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.439 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.429 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.419 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.409 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.399 | 20.068 | 0.023 | 2.3
[6 empty] | -0.146 | 38.389 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.379 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.369 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.359 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.349 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.339 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.329 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.319 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.309 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.299 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.289 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.279 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.269 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.259 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.249 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.239 | 20.068 | 0.024 | 2.4
[6 empty] | -0.146 | 38.229 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.219 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.209 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.199 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.189 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.179 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.169 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.159 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.149 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.139 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.129 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.119 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.109 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.099 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.089 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.079 | 20.068 | 0.025 | 2.5
[6 empty] | -0.146 | 38.069 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.059 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.049 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.039 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.029 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.019 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 38.009 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.999 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.989 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.979 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.969 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.959 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.949 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.939 | 20.068 | 0.026 | 2.6
[6 empty] | -0.146 | 37.929 | 20.067 | 0.026 | 2.6
[6 empty] | -0.146 | 37.919 | 20.067 | 0.026 | 2.6
[6 empty] | -0.146 | 37.909 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.899 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.889 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.879 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.869 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.859 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.849 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.839 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.829 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.819 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.809 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.799 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.789 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.779 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.769 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.759 | 20.067 | 0.027 | 2.7
[6 empty] | -0.146 | 37.749 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.739 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.729 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.719 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.709 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.699 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.689 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.679 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.669 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.659 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.649 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.639 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.629 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.619 | 20.067 | 0.028 | 2.8
[6 empty] | -0.146 | 37.609 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.599 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.589 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.579 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.569 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.559 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.549 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.539 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.529 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.519 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.509 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.499 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.489 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.479 | 20.067 | 0.029 | 2.9
[6 empty] | -0.146 | 37.469 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.459 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.449 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.439 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.429 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.419 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.409 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.399 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.389 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.379 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.369 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.359 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.349 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.339 | 20.067 | 0.03 | 3
[6 empty] | -0.146 | 37.329 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.319 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.309 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.299 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.289 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.279 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.269 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.259 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.249 | 20.067 | 0.031 | 3.1
[6 empty] | -0.146 | 37.239 | 20.066 | 0.031 | 3.1
[6 empty] | -0.146 | 37.229 | 20.066 | 0.031 | 3.1
[6 empty] | -0.146 | 37.219 | 20.066 | 0.031 | 3.1
[6 empty] | -0.146 | 37.209 | 20.066 | 0.031 | 3.1
[6 empty] | -0.146 | 37.199 | 20.066 | 0.031 | 3.1
[6 empty] | -0.146 | 37.189 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.179 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.169 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.159 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.149 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.139 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.129 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.119 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.109 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.099 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.089 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.079 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.069 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.059 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.049 | 20.066 | 0.032 | 3.2
[6 empty] | -0.146 | 37.039 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 37.029 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 37.019 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 37.009 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.999 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.989 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.979 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.97 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.96 | 20.066 | 0.033 | 3.3
[6 empty] | -0.146 | 36.95 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.94 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.93 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.92 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.91 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.9 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.89 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.88 | 20.066 | 0.034 | 3.4
[6 empty] | -0.146 | 36.87 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.86 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.85 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.84 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.83 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.82 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.81 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.8 | 20.066 | 0.035 | 3.5
[6 empty] | -0.146 | 36.79 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.78 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.77 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.76 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.75 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.74 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.73 | 20.066 | 0.036 | 3.6
[6 empty] | -0.146 | 36.72 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.71 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.7 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.69 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.68 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.67 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.66 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.65 | 20.066 | 0.037 | 3.7
[6 empty] | -0.146 | 36.64 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.63 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.62 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.61 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.6 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.59 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.58 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.57 | 20.066 | 0.038 | 3.8
[6 empty] | -0.146 | 36.56 | 20.066 | 0.039 | 3.9
[6 empty] | -0.146 | 36.55 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.54 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.53 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.52 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.51 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.5 | 20.065 | 0.039 | 3.9
[6 empty] | -0.146 | 36.49 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.48 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.47 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.46 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.45 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.44 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.43 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.42 | 20.065 | 0.04 | 4
[6 empty] | -0.146 | 36.41 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.4 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.39 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.38 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.37 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.36 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.35 | 20.065 | 0.041 | 4.1
[6 empty] | -0.146 | 36.34 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.33 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.32 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.31 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.3 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.29 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.28 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.27 | 20.065 | 0.042 | 4.2
[6 empty] | -0.146 | 36.26 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.25 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.24 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.23 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.22 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.21 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.2 | 20.065 | 0.043 | 4.3
[6 empty] | -0.146 | 36.19 | 20.065 | 0.044 | 4.4
[6 empty] | -0.146 | 36.18 | 20.065 | 0.044 | 4.4
[6 empty] | -0.146 | 36.17 | 20.065 | 0.045 | 4.5
[6 empty] | -0.146 | 36.16 | 20.065 | 0.045 | 4.5
[6 empty] | -0.146 | 36.15 | 20.065 | 0.046 | 4.6
[6 empty] | -0.146 | 36.14 | 20.065 | 0.046 | 4.6
[6 empty] | -0.146 | 36.13 | 20.065 | 0.047 | 4.7
[6 empty] | -0.146 | 36.12 | 20.065 | 0.047 | 4.7
[6 empty] | -0.146 | 36.11 | 20.065 | 0.048 | 4.8
[6 empty] | -0.146 | 36.1 | 20.065 | 0.048 | 4.8
[6 empty] | -0.146 | 36.09 | 20.065 | 0.049 | 4.9
[6 empty] | -0.146 | 36.08 | 20.065 | 0.049 | 4.9
[6 empty] | -0.146 | 36.07 | 20.065 | 0.05 | 5
[6 empty] | -0.146 | 36.06 | 20.065 | 0.05 | 5
[6 empty] | -0.146 | 36.05 | 20.065 | 0.051 | 5.1
[6 empty] | -0.146 | 36.04 | 20.065 | 0.051 | 5.1
[6 empty] | -0.146 | 36.03 | 20.065 | 0.052 | 5.2
[6 empty] | -0.146 | 36.02 | 20.065 | 0.052 | 5.2
[6 empty] | -0.146 | 36.01 | 20.065 | 0.053 | 5.3
[6 empty] | -0.146 | 36 | 20.065 | 0.053 | 5.3
[6 empty] | -0.146 | 35.99 | 20.065 | 0.054 | 5.4
[6 empty] | -0.146 | 35.98 | 20.065 | 0.054 | 5.4
[6 empty] | -0.146 | 35.97 | 20.065 | 0.054 | 5.4
[6 empty] | -0.146 | 35.96 | 20.065 | 0.055 | 5.5
[6 empty] | -0.146 | 35.95 | 20.065 | 0.055 | 5.5
[6 empty] | -0.146 | 35.94 | 20.065 | 0.056 | 5.6
[6 empty] | -0.146 | 35.93 | 20.065 | 0.056 | 5.6
[6 empty] | -0.146 | 35.92 | 20.065 | 0.057 | 5.7
[6 empty] | -0.146 | 35.91 | 20.065 | 0.057 | 5.7
[6 empty] | -0.146 | 35.9 | 20.065 | 0.058 | 5.8
[6 empty] | -0.146 | 35.89 | 20.065 | 0.058 | 5.8
[6 empty] | -0.146 | 35.88 | 20.065 | 0.059 | 5.9
[6 empty] | -0.146 | 35.87 | 20.065 | 0.059 | 5.9
[6 empty] | -0.146 | 35.86 | 20.064 | 0.06 | 6
[6 empty] | -0.146 | 35.85 | 20.064 | 0.06 | 6
[6 empty] | -0.146 | 35.84 | 20.064 | 0.061 | 6.1
[6 empty] | -0.146 | 35.83 | 20.064 | 0.061 | 6.1
[6 empty] | -0.146 | 35.82 | 20.064 | 0.062 | 6.2
[6 empty] | -0.146 | 35.81 | 20.064 | 0.062 | 6.2
[6 empty] | -0.146 | 35.8 | 20.064 | 0.063 | 6.3
[6 empty] | -0.146 | 35.79 | 20.064 | 0.065 | 6.5
[6 empty] | -0.146 | 35.78 | 20.064 | 0.067 | 6.7
[6 empty] | -0.146 | 35.77 | 20.064 | 0.068 | 6.8
[6 empty] | -0.146 | 35.76 | 20.064 | 0.07 | 7
[6 empty] | -0.146 | 35.75 | 20.064 | 0.072 | 7.2
[6 empty] | -0.146 | 35.74 | 20.064 | 0.074 | 7.4
[6 empty] | -0.146 | 35.73 | 20.064 | 0.076 | 7.6
[6 empty] | -0.146 | 35.72 | 20.064 | 0.077 | 7.7
[6 empty] | -0.146 | 35.71 | 20.064 | 0.079 | 7.9
[6 empty] | -0.146 | 35.7 | 20.064 | 0.081 | 8.1
[6 empty] | -0.146 | 35.69 | 20.064 | 0.083 | 8.3
[6 empty] | -0.146 | 35.68 | 20.064 | 0.085 | 8.5
[6 empty] | -0.146 | 35.67 | 20.064 | 0.086 | 8.6
[6 empty] | -0.146 | 35.66 | 20.064 | 0.088 | 8.8
[6 empty] | -0.146 | 35.65 | 20.064 | 0.09 | 9
[6 empty] | -0.146 | 35.64 | 20.064 | 0.092 | 9.2
[6 empty] | -0.146 | 35.63 | 20.064 | 0.094 | 9.4
[6 empty] | -0.146 | 35.62 | 20.064 | 0.095 | 9.5
[6 empty] | -0.146 | 35.61 | 20.064 | 0.097 | 9.7
[6 empty] | -0.146 | 35.6 | 20.064 | 0.101 | 10.1
[6 empty] | -0.146 | 35.59 | 20.064 | 0.109 | 10.9
[6 empty] | -0.146 | 35.58 | 20.064 | 0.118 | 11.8
[6 empty] | -0.146 | 35.57 | 20.064 | 0.126 | 12.6
[6 empty] | -0.146 | 35.56 | 20.064 | 0.134 | 13.4
[6 empty] | -0.146 | 35.55 | 20.064 | 0.142 | 14.2
[6 empty] | -0.146 | 35.54 | 20.064 | 0.151 | 15.1
[6 empty] | -0.146 | 35.53 | 20.064 | 0.159 | 15.9
[6 empty] | -0.146 | 35.52 | 20.064 | 0.167 | 16.7
[6 empty] | -0.146 | 35.51 | 20.064 | 0.176 | 17.6
[6 empty] | -0.146 | 35.5 | 20.064 | 0.184 | 18.4
[6 empty] | -0.146 | 35.49 | 20.064 | 0.192 | 19.2
[6 empty] | -0.146 | 35.48 | 20.064 | 0.2 | 20
[6 empty] | -0.146 | 35.47 | 20.064 | 0.209 | 20.9
[6 empty] | -0.146 | 35.46 | 20.064 | 0.217 | 21.7
[6 empty] | -0.146 | 35.45 | 20.064 | 0.225 | 22.5
[6 empty] | -0.146 | 35.44 | 20.064 | 0.234 | 23.4
[6 empty] | -0.146 | 35.43 | 20.064 | 0.242 | 24.2
[6 empty] | -0.146 | 35.42 | 20.064 | 0.25 | 25
[6 empty] | -0.146 | 35.41 | 20.064 | 0.258 | 25.8
[6 empty] | -0.146 | 35.4 | 20.064 | 0.266 | 26.6
[6 empty] | -0.146 | 35.39 | 20.064 | 0.272 | 27.2
[6 empty] | -0.146 | 35.38 | 20.064 | 0.278 | 27.8
[6 empty] | -0.146 | 35.37 | 20.064 | 0.284 | 28.4
[6 empty] | -0.146 | 35.36 | 20.064 | 0.29 | 29
[6 empty] | -0.146 | 35.35 | 20.064 | 0.296 | 29.6
[6 empty] | -0.146 | 35.34 | 20.064 | 0.303 | 30.3
[6 empty] | -0.146 | 35.33 | 20.064 | 0.309 | 30.9
[6 empty] | -0.146 | 35.32 | 20.064 | 0.315 | 31.5
[6 empty] | -0.146 | 35.31 | 20.064 | 0.321 | 32.1
[6 empty] | -0.146 | 35.3 | 20.064 | 0.327 | 32.7
[6 empty] | -0.146 | 35.29 | 20.064 | 0.333 | 33.3
[6 empty] | -0.146 | 35.28 | 20.064 | 0.339 | 33.9
[6 empty] | -0.146 | 35.27 | 20.064 | 0.345 | 34.5
[6 empty] | -0.146 | 35.26 | 20.064 | 0.351 | 35.1
[6 empty] | -0.146 | 35.25 | 20.064 | 0.357 | 35.7
[6 empty] | -0.146 | 35.24 | 20.064 | 0.363 | 36.3
[6 empty] | -0.146 | 35.23 | 20.064 | 0.369 | 36.9
[6 empty] | -0.146 | 35.22 | 20.064 | 0.375 | 37.5
[6 empty] | -0.146 | 35.21 | 20.064 | 0.381 | 38.1
[6 empty] | -0.146 | 35.2 | 20.064 | 0.385 | 38.5
[6 empty] | -0.146 | 35.19 | 20.064 | 0.383 | 38.3
[6 empty] | -0.146 | 35.18 | 20.064 | 0.381 | 38.1
[6 empty] | -0.146 | 35.17 | 20.063 | 0.379 | 37.9
[6 empty] | -0.146 | 35.16 | 20.063 | 0.378 | 37.8
[6 empty] | -0.146 | 35.15 | 20.063 | 0.376 | 37.6
[6 empty] | -0.146 | 35.14 | 20.063 | 0.374 | 37.4
[6 empty] | -0.146 | 35.13 | 20.063 | 0.372 | 37.2
[6 empty] | -0.146 | 35.12 | 20.063 | 0.37 | 37
[6 empty] | -0.146 | 35.11 | 20.063 | 0.368 | 36.8
[6 empty] | -0.146 | 35.1 | 20.063 | 0.366 | 36.6
[6 empty] | -0.146 | 35.09 | 20.063 | 0.365 | 36.5
[6 empty] | -0.146 | 35.08 | 20.063 | 0.363 | 36.3
[6 empty] | -0.146 | 35.07 | 20.063 | 0.361 | 36.1
[6 empty] | -0.146 | 35.06 | 20.063 | 0.359 | 35.9
[6 empty] | -0.146 | 35.05 | 20.063 | 0.357 | 35.7
[6 empty] | -0.146 | 35.04 | 20.063 | 0.355 | 35.5
[6 empty] | -0.146 | 35.03 | 20.063 | 0.353 | 35.3
[6 empty] | -0.146 | 35.02 | 20.063 | 0.352 | 35.2
[6 empty] | -0.146 | 35.01 | 20.063 | 0.35 | 35
[6 empty] | -0.146 | 35 | 20.063 | 0.345 | 34.5
[6 empty] | -0.146 | 34.99 | 20.063 | 0.336 | 33.6
[6 empty] | -0.146 | 34.98 | 20.063 | 0.326 | 32.6
[6 empty] | -0.146 | 34.97 | 20.063 | 0.317 | 31.7
[6 empty] | -0.146 | 34.96 | 20.063 | 0.307 | 30.7
[6 empty] | -0.146 | 34.95 | 20.063 | 0.298 | 29.8
[6 empty] | -0.146 | 34.94 | 20.063 | 0.288 | 28.8
[6 empty] | -0.146 | 34.93 | 20.063 | 0.279 | 27.9
[6 empty] | -0.146 | 34.92 | 20.063 | 0.269 | 26.9
[6 empty] | -0.146 | 34.91 | 20.063 | 0.26 | 26
[6 empty] | -0.146 | 34.9 | 20.063 | 0.25 | 25
[6 empty] | -0.146 | 34.89 | 20.063 | 0.241 | 24.1
[6 empty] | -0.146 | 34.88 | 20.063 | 0.231 | 23.1
[6 empty] | -0.146 | 34.87 | 20.063 | 0.222 | 22.2
[6 empty] | -0.146 | 34.86 | 20.063 | 0.212 | 21.2
[6 empty] | -0.146 | 34.85 | 20.063 | 0.203 | 20.3
[6 empty] | -0.146 | 34.84 | 20.063 | 0.193 | 19.3
[6 empty] | -0.146 | 34.83 | 20.063 | 0.184 | 18.4
[6 empty] | -0.146 | 34.82 | 20.063 | 0.174 | 17.4
[6 empty] | -0.146 | 34.81 | 20.063 | 0.165 | 16.5
[6 empty] | -0.146 | 34.8 | 20.063 | 0.158 | 15.8
[6 empty] | -0.146 | 34.79 | 20.063 | 0.155 | 15.5
[6 empty] | -0.146 | 34.78 | 20.063 | 0.153 | 15.3
[6 empty] | -0.146 | 34.77 | 20.063 | 0.151 | 15.1
[6 empty] | -0.146 | 34.76 | 20.063 | 0.149 | 14.9
[6 empty] | -0.146 | 34.75 | 20.063 | 0.147 | 14.7
[6 empty] | -0.146 | 34.74 | 20.063 | 0.144 | 14.4
[6 empty] | -0.146 | 34.73 | 20.063 | 0.142 | 14.2
[6 empty] | -0.146 | 34.72 | 20.063 | 0.14 | 14
[6 empty] | -0.146 | 34.71 | 20.063 | 0.138 | 13.8
[6 empty] | -0.146 | 34.7 | 20.063 | 0.135 | 13.5
[6 empty] | -0.146 | 34.69 | 20.063 | 0.133 | 13.3
[6 empty] | -0.146 | 34.68 | 20.063 | 0.131 | 13.1
[6 empty] | -0.146 | 34.67 | 20.063 | 0.129 | 12.9
[6 empty] | -0.146 | 34.66 | 20.063 | 0.127 | 12.7
[6 empty] | -0.146 | 34.65 | 20.063 | 0.124 | 12.4
[6 empty] | -0.146 | 34.64 | 20.063 | 0.122 | 12.2
[6 empty] | -0.146 | 34.63 | 20.063 | 0.12 | 12
[6 empty] | -0.146 | 34.62 | 20.063 | 0.118 | 11.8
[6 empty] | -0.146 | 34.61 | 20.063 | 0.115 | 11.5
[6 empty] | -0.146 | 34.6 | 20.063 | 0.115 | 11.5
[6 empty] | -0.146 | 34.59 | 20.063 | 0.118 | 11.8
[6 empty] | -0.146 | 34.58 | 20.063 | 0.122 | 12.2
[6 empty] | -0.146 | 34.57 | 20.063 | 0.125 | 12.5
[6 empty] | -0.146 | 34.56 | 20.063 | 0.129 | 12.9
[6 empty] | -0.146 | 34.55 | 20.063 | 0.132 | 13.2
[6 empty] | -0.146 | 34.54 | 20.063 | 0.136 | 13.6
[6 empty] | -0.146 | 34.53 | 20.063 | 0.139 | 13.9
[6 empty] | -0.146 | 34.52 | 20.063 | 0.143 | 14.3
[6 empty] | -0.146 | 34.51 | 20.063 | 0.146 | 14.6
[6 empty] | -0.146 | 34.5 | 20.063 | 0.15 | 15
[6 empty] | -0.146 | 34.49 | 20.063 | 0.153 | 15.3
[6 empty] | -0.146 | 34.48 | 20.062 | 0.156 | 15.6
[6 empty] | -0.146 | 34.47 | 20.062 | 0.16 | 16
[6 empty] | -0.146 | 34.46 | 20.062 | 0.163 | 16.3
[6 empty] | -0.146 | 34.45 | 20.062 | 0.167 | 16.7
[6 empty] | -0.146 | 34.44 | 20.062 | 0.17 | 17
[6 empty] | -0.146 | 34.43 | 20.062 | 0.174 | 17.4
[6 empty] | -0.146 | 34.42 | 20.062 | 0.177 | 17.7
[6 empty] | -0.146 | 34.41 | 20.062 | 0.181 | 18.1
[6 empty] | -0.146 | 34.4 | 20.062 | 0.186 | 18.6
[6 empty] | -0.146 | 34.39 | 20.062 | 0.195 | 19.5
[6 empty] | -0.146 | 34.38 | 20.062 | 0.204 | 20.4
[6 empty] | -0.146 | 34.37 | 20.062 | 0.213 | 21.3
[6 empty] | -0.146 | 34.36 | 20.062 | 0.222 | 22.2
[6 empty] | -0.146 | 34.35 | 20.062 | 0.232 | 23.2
[6 empty] | -0.146 | 34.34 | 20.062 | 0.241 | 24.1
[6 empty] | -0.146 | 34.33 | 20.062 | 0.25 | 25
[6 empty] | -0.146 | 34.32 | 20.062 | 0.259 | 25.9
[6 empty] | -0.146 | 34.31 | 20.062 | 0.268 | 26.8
[6 empty] | -0.146 | 34.3 | 20.062 | 0.277 | 27.7
[6 empty] | -0.146 | 34.29 | 20.062 | 0.286 | 28.6
[6 empty] | -0.146 | 34.28 | 20.062 | 0.296 | 29.6
[6 empty] | -0.146 | 34.27 | 20.062 | 0.305 | 30.5
[6 empty] | -0.146 | 34.26 | 20.062 | 0.314 | 31.4
[6 empty] | -0.146 | 34.25 | 20.062 | 0.323 | 32.3
[6 empty] | -0.146 | 34.24 | 20.062 | 0.332 | 33.2
[6 empty] | -0.146 | 34.23 | 20.062 | 0.341 | 34.1
[6 empty] | -0.146 | 34.22 | 20.062 | 0.35 | 35
[6 empty] | -0.146 | 34.21 | 20.062 | 0.36 | 36
[6 empty] | -0.146 | 34.2 | 20.062 | 0.366 | 36.6
[6 empty] | -0.146 | 34.19 | 20.062 | 0.368 | 36.8
[6 empty] | -0.146 | 34.18 | 20.062 | 0.369 | 36.9
[6 empty] | -0.146 | 34.17 | 20.062 | 0.371 | 37.1
[6 empty] | -0.146 | 34.16 | 20.062 | 0.372 | 37.2
[6 empty] | -0.146 | 34.15 | 20.062 | 0.374 | 37.4
[6 empty] | -0.146 | 34.14 | 20.062 | 0.375 | 37.5
[6 empty] | -0.146 | 34.13 | 20.062 | 0.377 | 37.7
[6 empty] | -0.146 | 34.12 | 20.062 | 0.378 | 37.8
[6 empty] | -0.146 | 34.11 | 20.062 | 0.38 | 38
[6 empty] | -0.146 | 34.1 | 20.062 | 0.381 | 38.1
[6 empty] | -0.146 | 34.09 | 20.062 | 0.383 | 38.3
[6 empty] | -0.146 | 34.08 | 20.062 | 0.384 | 38.4
[6 empty] | -0.146 | 34.07 | 20.062 | 0.386 | 38.6
[6 empty] | -0.146 | 34.06 | 20.062 | 0.387 | 38.7
[6 empty] | -0.146 | 34.05 | 20.062 | 0.389 | 38.9
[6 empty] | -0.146 | 34.04 | 20.062 | 0.39 | 39
[6 empty] | -0.146 | 34.03 | 20.062 | 0.392 | 39.2
[6 empty] | -0.146 | 34.02 | 20.062 | 0.393 | 39.3
[6 empty] | -0.146 | 34.01 | 20.062 | 0.395 | 39.5
[6 empty] | -0.146 | 34 | 20.062 | 0.394 | 39.4
[6 empty] | -0.146 | 33.99 | 20.062 | 0.387 | 38.7
[6 empty] | -0.146 | 33.98 | 20.062 | 0.38 | 38
[6 empty] | -0.146 | 33.971 | 20.062 | 0.373 | 37.3
[6 empty] | -0.146 | 33.961 | 20.062 | 0.366 | 36.6
[6 empty] | -0.146 | 33.951 | 20.062 | 0.359 | 35.9
[6 empty] | -0.146 | 33.941 | 20.062 | 0.352 | 35.2
[6 empty] | -0.146 | 33.931 | 20.062 | 0.346 | 34.6
[6 empty] | -0.146 | 33.921 | 20.062 | 0.339 | 33.9
[6 empty] | -0.146 | 33.911 | 20.062 | 0.332 | 33.2
[6 empty] | -0.146 | 33.901 | 20.062 | 0.325 | 32.5
[6 empty] | -0.146 | 33.891 | 20.062 | 0.318 | 31.8
[6 empty] | -0.146 | 33.881 | 20.062 | 0.311 | 31.1
[6 empty] | -0.146 | 33.871 | 20.062 | 0.304 | 30.4
[6 empty] | -0.146 | 33.861 | 20.062 | 0.297 | 29.7
[6 empty] | -0.146 | 33.851 | 20.062 | 0.291 | 29.1
[6 empty] | -0.146 | 33.841 | 20.062 | 0.284 | 28.4
[6 empty] | -0.146 | 33.831 | 20.062 | 0.277 | 27.7
[6 empty] | -0.146 | 33.821 | 20.062 | 0.27 | 27
[6 empty] | -0.146 | 33.811 | 20.062 | 0.263 | 26.3
[6 empty] | -0.146 | 33.801 | 20.062 | 0.257 | 25.7
[6 empty] | -0.146 | 33.791 | 20.061 | 0.25 | 25
[6 empty] | -0.146 | 33.781 | 20.061 | 0.244 | 24.4
[6 empty] | -0.146 | 33.771 | 20.061 | 0.238 | 23.8
[6 empty] | -0.146 | 33.761 | 20.061 | 0.231 | 23.1
[6 empty] | -0.146 | 33.751 | 20.061 | 0.225 | 22.5
[6 empty] | -0.146 | 33.741 | 20.061 | 0.219 | 21.9
[6 empty] | -0.146 | 33.731 | 20.061 | 0.213 | 21.3
[6 empty] | -0.146 | 33.721 | 20.061 | 0.206 | 20.6
[6 empty] | -0.146 | 33.711 | 20.061 | 0.2 | 20
[6 empty] | -0.146 | 33.701 | 20.061 | 0.194 | 19.4
[6 empty] | -0.146 | 33.691 | 20.061 | 0.188 | 18.8
[6 empty] | -0.146 | 33.681 | 20.061 | 0.181 | 18.1
[6 empty] | -0.146 | 33.671 | 20.061 | 0.175 | 17.5
[6 empty] | -0.146 | 33.661 | 20.061 | 0.169 | 16.9
[6 empty] | -0.146 | 33.651 | 20.061 | 0.163 | 16.3
[6 empty] | -0.146 | 33.641 | 20.061 | 0.156 | 15.6
[6 empty] | -0.146 | 33.631 | 20.061 | 0.15 | 15
[6 empty] | -0.146 | 33.621 | 20.061 | 0.144 | 14.4
[6 empty] | -0.146 | 33.611 | 20.061 | 0.137 | 13.7
[6 empty] | -0.146 | 33.601 | 20.061 | 0.133 | 13.3
[6 empty] | -0.146 | 33.591 | 20.061 | 0.133 | 13.3
[6 empty] | -0.146 | 33.581 | 20.061 | 0.134 | 13.4
[6 empty] | -0.146 | 33.571 | 20.061 | 0.134 | 13.4
[6 empty] | -0.146 | 33.561 | 20.061 | 0.134 | 13.4
[6 empty] | -0.146 | 33.551 | 20.061 | 0.134 | 13.4
[6 empty] | -0.146 | 33.541 | 20.061 | 0.135 | 13.5
[6 empty] | -0.146 | 33.531 | 20.061 | 0.135 | 13.5
[6 empty] | -0.146 | 33.521 | 20.061 | 0.135 | 13.5
[6 empty] | -0.146 | 33.511 | 20.061 | 0.135 | 13.5
[6 empty] | -0.146 | 33.501 | 20.061 | 0.136 | 13.6
[6 empty] | -0.146 | 33.491 | 20.061 | 0.136 | 13.6
[6 empty] | -0.146 | 33.481 | 20.061 | 0.136 | 13.6
[6 empty] | -0.146 | 33.471 | 20.061 | 0.136 | 13.6
[6 empty] | -0.146 | 33.461 | 20.061 | 0.136 | 13.6
[6 empty] | -0.146 | 33.451 | 20.061 | 0.137 | 13.7
[6 empty] | -0.146 | 33.441 | 20.061 | 0.137 | 13.7
[6 empty] | -0.146 | 33.431 | 20.061 | 0.137 | 13.7
[6 empty] | -0.146 | 33.421 | 20.061 | 0.137 | 13.7
[6 empty] | -0.146 | 33.411 | 20.061 | 0.138 | 13.8
[6 empty] | -0.146 | 33.401 | 20.061 | 0.14 | 14
[6 empty] | -0.146 | 33.391 | 20.061 | 0.148 | 14.8
[6 empty] | -0.146 | 33.381 | 20.061 | 0.156 | 15.6
[6 empty] | -0.146 | 33.371 | 20.061 | 0.163 | 16.3
[6 empty] | -0.146 | 33.361 | 20.061 | 0.171 | 17.1
[6 empty] | -0.146 | 33.351 | 20.061 | 0.179 | 17.9
[6 empty] | -0.146 | 33.341 | 20.061 | 0.186 | 18.6
[6 empty] | -0.146 | 33.331 | 20.061 | 0.194 | 19.4
[6 empty] | -0.146 | 33.321 | 20.061 | 0.202 | 20.2
[6 empty] | -0.146 | 33.311 | 20.061 | 0.21 | 21
[6 empty] | -0.146 | 33.301 | 20.061 | 0.217 | 21.7
[6 empty] | -0.146 | 33.291 | 20.061 | 0.225 | 22.5
[6 empty] | -0.146 | 33.281 | 20.061 | 0.233 | 23.3
[6 empty] | -0.146 | 33.271 | 20.061 | 0.241 | 24.1
[6 empty] | -0.146 | 33.261 | 20.061 | 0.248 | 24.8
[6 empty] | -0.146 | 33.251 | 20.061 | 0.256 | 25.6
[6 empty] | -0.146 | 33.241 | 20.061 | 0.264 | 26.4
[6 empty] | -0.146 | 33.231 | 20.061 | 0.271 | 27.1
[6 empty] | -0.146 | 33.221 | 20.061 | 0.279 | 27.9
[6 empty] | -0.146 | 33.211 | 20.061 | 0.287 | 28.7
[6 empty] | -0.146 | 33.201 | 20.061 | 0.294 | 29.4
[6 empty] | -0.146 | 33.191 | 20.061 | 0.3 | 30
[6 empty] | -0.146 | 33.181 | 20.061 | 0.305 | 30.5
[6 empty] | -0.146 | 33.171 | 20.061 | 0.311 | 31.1
[6 empty] | -0.146 | 33.161 | 20.061 | 0.316 | 31.6
[6 empty] | -0.146 | 33.151 | 20.061 | 0.322 | 32.2
[6 empty] | -0.146 | 33.141 | 20.061 | 0.328 | 32.8
[6 empty] | -0.146 | 33.131 | 20.061 | 0.333 | 33.3
[6 empty] | -0.146 | 33.121 | 20.061 | 0.339 | 33.9
[6 empty] | -0.146 | 33.111 | 20.061 | 0.344 | 34.4
[6 empty] | -0.146 | 33.101 | 20.06 | 0.35 | 35
[6 empty] | -0.146 | 33.091 | 20.06 | 0.355 | 35.5
[6 empty] | -0.146 | 33.081 | 20.06 | 0.361 | 36.1
[6 empty] | -0.146 | 33.071 | 20.06 | 0.367 | 36.7
[6 empty] | -0.146 | 33.061 | 20.06 | 0.372 | 37.2
[6 empty] | -0.146 | 33.051 | 20.06 | 0.378 | 37.8
[6 empty] | -0.146 | 33.041 | 20.06 | 0.383 | 38.3
[6 empty] | -0.146 | 33.031 | 20.06 | 0.389 | 38.9
[6 empty] | -0.146 | 33.021 | 20.06 | 0.395 | 39.5
[6 empty] | -0.146 | 33.011 | 20.06 | 0.4 | 40
[6 empty] | -0.146 | 33.001 | 20.06 | 0.404 | 40.4
[6 empty] | -0.146 | 32.991 | 20.06 | 0.401 | 40.1
[6 empty] | -0.146 | 32.981 | 20.06 | 0.399 | 39.9
[6 empty] | -0.146 | 32.971 | 20.06 | 0.397 | 39.7
[6 empty] | -0.146 | 32.961 | 20.06 | 0.395 | 39.5
[6 empty] | -0.146 | 32.951 | 20.06 | 0.393 | 39.3
[6 empty] | -0.146 | 32.941 | 20.06 | 0.391 | 39.1
[6 empty] | -0.146 | 32.931 | 20.06 | 0.389 | 38.9
[6 empty] | -0.146 | 32.921 | 20.06 | 0.387 | 38.7
[6 empty] | -0.146 | 32.911 | 20.06 | 0.385 | 38.5
[6 empty] | -0.146 | 32.901 | 20.06 | 0.383 | 38.3
[6 empty] | -0.146 | 32.891 | 20.06 | 0.38 | 38
[6 empty] | -0.146 | 32.881 | 20.06 | 0.378 | 37.8
[6 empty] | -0.146 | 32.871 | 20.06 | 0.376 | 37.6
[6 empty] | -0.146 | 32.861 | 20.06 | 0.374 | 37.4
[6 empty] | -0.146 | 32.851 | 20.06 | 0.372 | 37.2
[6 empty] | -0.146 | 32.841 | 20.06 | 0.37 | 37
[6 empty] | -0.146 | 32.831 | 20.06 | 0.368 | 36.8
[6 empty] | -0.146 | 32.821 | 20.06 | 0.366 | 36.6
[6 empty] | -0.146 | 32.811 | 20.06 | 0.364 | 36.4
[6 empty] | -0.146 | 32.801 | 20.06 | 0.36 | 36
[6 empty] | -0.146 | 32.791 | 20.06 | 0.35 | 35
[6 empty] | -0.146 | 32.781 | 20.06 | 0.341 | 34.1
[6 empty] | -0.146 | 32.771 | 20.06 | 0.331 | 33.1
[6 empty] | -0.146 | 32.761 | 20.06 | 0.322 | 32.2
[6 empty] | -0.146 | 32.751 | 20.06 | 0.312 | 31.2
[6 empty] | -0.146 | 32.741 | 20.06 | 0.303 | 30.3
[6 empty] | -0.146 | 32.731 | 20.06 | 0.293 | 29.3
[6 empty] | -0.146 | 32.721 | 20.06 | 0.284 | 28.4
[6 empty] | -0.146 | 32.711 | 20.06 | 0.274 | 27.4
[6 empty] | -0.146 | 32.701 | 20.06 | 0.265 | 26.5
[6 empty] | -0.146 | 32.691 | 20.06 | 0.255 | 25.5
[6 empty] | -0.146 | 32.681 | 20.06 | 0.246 | 24.6
[6 empty] | -0.146 | 32.671 | 20.06 | 0.236 | 23.6
[6 empty] | -0.146 | 32.661 | 20.06 | 0.227 | 22.7
[6 empty] | -0.146 | 32.651 | 20.06 | 0.217 | 21.7
[6 empty] | -0.146 | 32.641 | 20.06 | 0.208 | 20.8
[6 empty] | -0.146 | 32.631 | 20.06 | 0.198 | 19.8
[6 empty] | -0.146 | 32.621 | 20.06 | 0.189 | 18.9
[6 empty] | -0.146 | 32.611 | 20.06 | 0.179 | 17.9
[6 empty] | -0.146 | 32.601 | 20.06 | 0.172 | 17.2
[6 empty] | -0.146 | 32.591 | 20.06 | 0.17 | 17
[6 empty] | -0.146 | 32.581 | 20.06 | 0.167 | 16.7
[6 empty] | -0.146 | 32.571 | 20.06 | 0.165 | 16.5
[6 empty] | -0.146 | 32.561 | 20.06 | 0.163 | 16.3
[6 empty] | -0.146 | 32.551 | 20.06 | 0.161 | 16.1
[6 empty] | -0.146 | 32.541 | 20.06 | 0.159 | 15.9
[6 empty] | -0.146 | 32.531 | 20.06 | 0.157 | 15.7
[6 empty] | -0.146 | 32.521 | 20.06 | 0.155 | 15.5
[6 empty] | -0.146 | 32.511 | 20.06 | 0.152 | 15.2
[6 empty] | -0.146 | 32.501 | 20.06 | 0.15 | 15
[6 empty] | -0.146 | 32.491 | 20.06 | 0.148 | 14.8
[6 empty] | -0.146 | 32.481 | 20.06 | 0.146 | 14.6
[6 empty] | -0.146 | 32.471 | 20.06 | 0.144 | 14.4
[6 empty] | -0.146 | 32.461 | 20.06 | 0.142 | 14.2
[6 empty] | -0.146 | 32.451 | 20.06 | 0.14 | 14
[6 empty] | -0.146 | 32.441 | 20.06 | 0.138 | 13.8
[6 empty] | -0.146 | 32.431 | 20.06 | 0.135 | 13.5
[6 empty] | -0.146 | 32.421 | 20.06 | 0.133 | 13.3
[6 empty] | -0.146 | 32.411 | 20.059 | 0.131 | 13.1
[6 empty] | -0.146 | 32.401 | 20.059 | 0.131 | 13.1
[6 empty] | -0.146 | 32.391 | 20.059 | 0.134 | 13.4
[6 empty] | -0.146 | 32.381 | 20.059 | 0.138 | 13.8
[6 empty] | -0.146 | 32.371 | 20.059 | 0.141 | 14.1
[6 empty] | -0.146 | 32.361 | 20.059 | 0.145 | 14.5
[6 empty] | -0.146 | 32.351 | 20.059 | 0.148 | 14.8
[6 empty] | -0.146 | 32.341 | 20.059 | 0.152 | 15.2
[6 empty] | -0.146 | 32.331 | 20.059 | 0.155 | 15.5
[6 empty] | -0.146 | 32.321 | 20.059 | 0.159 | 15.9
[6 empty] | -0.146 | 32.311 | 20.059 | 0.162 | 16.2
[6 empty] | -0.146 | 32.301 | 20.059 | 0.166 | 16.6
[6 empty] | -0.146 | 32.291 | 20.059 | 0.169 | 16.9
[6 empty] | -0.146 | 32.281 | 20.059 | 0.173 | 17.3
[6 empty] | -0.146 | 32.271 | 20.059 | 0.176 | 17.6
[6 empty] | -0.146 | 32.261 | 20.059 | 0.18 | 18
[6 empty] | -0.146 | 32.251 | 20.059 | 0.183 | 18.3
[6 empty] | -0.146 | 32.241 | 20.059 | 0.187 | 18.7
[6 empty] | -0.146 | 32.231 | 20.059 | 0.19 | 19
[6 empty] | -0.146 | 32.221 | 20.059 | 0.194 | 19.4
[6 empty] | -0.146 | 32.211 | 20.059 | 0.197 | 19.7
[6 empty] | -0.146 | 32.201 | 20.059 | 0.202 | 20.2
[6 empty] | -0.146 | 32.191 | 20.059 | 0.211 | 21.1
[6 empty] | -0.146 | 32.181 | 20.059 | 0.22 | 22
[6 empty] | -0.146 | 32.171 | 20.059 | 0.229 | 22.9
[6 empty] | -0.146 | 32.161 | 20.059 | 0.238 | 23.8
[6 empty] | -0.146 | 32.151 | 20.059 | 0.247 | 24.7
[6 empty] | -0.146 | 32.141 | 20.059 | 0.256 | 25.6
[6 empty] | -0.146 | 32.131 | 20.059 | 0.264 | 26.4
[6 empty] | -0.146 | 32.121 | 20.059 | 0.273 | 27.3
[6 empty] | -0.146 | 32.111 | 20.059 | 0.282 | 28.2
[6 empty] | -0.146 | 32.101 | 20.059 | 0.291 | 29.1
[6 empty] | -0.146 | 32.091 | 20.059 | 0.3 | 30
[6 empty] | -0.146 | 32.081 | 20.059 | 0.309 | 30.9
[6 empty] | -0.146 | 32.071 | 20.059 | 0.318 | 31.8
[6 empty] | -0.146 | 32.061 | 20.059 | 0.327 | 32.7
[6 empty] | -0.146 | 32.051 | 20.059 | 0.336 | 33.6
[6 empty] | -0.146 | 32.041 | 20.059 | 0.345 | 34.5
[6 empty] | -0.146 | 32.031 | 20.059 | 0.354 | 35.4
[6 empty] | -0.146 | 32.021 | 20.059 | 0.363 | 36.3
[6 empty] | -0.146 | 32.011 | 20.059 | 0.371 | 37.1
[6 empty] | -0.146 | 32.001 | 20.059 | 0.378 | 37.8
[6 empty] | -0.146 | 31.991 | 20.059 | 0.38 | 38
[6 empty] | -0.146 | 31.981 | 20.059 | 0.381 | 38.1
[6 empty] | -0.146 | 31.971 | 20.059 | 0.382 | 38.2
[6 empty] | -0.146 | 31.961 | 20.059 | 0.383 | 38.3
[6 empty] | -0.146 | 31.951 | 20.059 | 0.385 | 38.5
[6 empty] | -0.146 | 31.941 | 20.059 | 0.386 | 38.6
[6 empty] | -0.146 | 31.931 | 20.059 | 0.387 | 38.7
[6 empty] | -0.146 | 31.921 | 20.059 | 0.389 | 38.9
[6 empty] | -0.146 | 31.911 | 20.059 | 0.39 | 39
[6 empty] | -0.146 | 31.901 | 20.059 | 0.391 | 39.1
[6 empty] | -0.146 | 31.891 | 20.059 | 0.392 | 39.2
[6 empty] | -0.146 | 31.881 | 20.059 | 0.394 | 39.4
[6 empty] | -0.146 | 31.871 | 20.059 | 0.395 | 39.5
[6 empty] | -0.146 | 31.861 | 20.059 | 0.396 | 39.6
[6 empty] | -0.146 | 31.851 | 20.059 | 0.397 | 39.7
[6 empty] | -0.146 | 31.841 | 20.059 | 0.399 | 39.9
[6 empty] | -0.146 | 31.831 | 20.059 | 0.4 | 40
[6 empty] | -0.146 | 31.821 | 20.059 | 0.401 | 40.1
[6 empty] | -0.146 | 31.811 | 20.059 | 0.402 | 40.2
[6 empty] | -0.146 | 31.801 | 20.059 | 0.402 | 40.2
[6 empty] | -0.146 | 31.791 | 20.059 | 0.395 | 39.5
[6 empty] | -0.146 | 31.781 | 20.059 | 0.388 | 38.8
[6 empty] | -0.146 | 31.771 | 20.059 | 0.381 | 38.1
[6 empty] | -0.146 | 31.761 | 20.059 | 0.374 | 37.4
[6 empty] | -0.146 | 31.751 | 20.059 | 0.367 | 36.7
[6 empty] | -0.146 | 31.741 | 20.059 | 0.36 | 36
[6 empty] | -0.146 | 31.731 | 20.059 | 0.353 | 35.3
[6 empty] | -0.146 | 31.721 | 20.058 | 0.346 | 34.6
[6 empty] | -0.146 | 31.711 | 20.058 | 0.339 | 33.9
[6 empty] | -0.146 | 31.701 | 20.058 | 0.332 | 33.2
[6 empty] | -0.146 | 31.691 | 20.058 | 0.325 | 32.5
[6 empty] | -0.146 | 31.681 | 20.058 | 0.318 | 31.8
[6 empty] | -0.146 | 31.671 | 20.058 | 0.311 | 31.1
[6 empty] | -0.146 | 31.661 | 20.058 | 0.304 | 30.4
[6 empty] | -0.146 | 31.651 | 20.058 | 0.297 | 29.7
[6 empty] | -0.146 | 31.641 | 20.058 | 0.29 | 29
[6 empty] | -0.146 | 31.631 | 20.058 | 0.283 | 28.3
[6 empty] | -0.146 | 31.621 | 20.058 | 0.276 | 27.6
[6 empty] | -0.146 | 31.611 | 20.058 | 0.27 | 27
[6 empty] | -0.146 | 31.601 | 20.058 | 0.263 | 26.3
[6 empty] | -0.146 | 31.591 | 20.058 | 0.257 | 25.7
[6 empty] | -0.146 | 31.581 | 20.058 | 0.251 | 25.1
[6 empty] | -0.146 | 31.571 | 20.058 | 0.244 | 24.4
[6 empty] | -0.146 | 31.561 | 20.058 | 0.238 | 23.8
[6 empty] | -0.146 | 31.551 | 20.058 | 0.232 | 23.2
[6 empty] | -0.146 | 31.541 | 20.058 | 0.226 | 22.6
[6 empty] | -0.146 | 31.531 | 20.058 | 0.22 | 22
[6 empty] | -0.146 | 31.521 | 20.058 | 0.214 | 21.4
[6 empty] | -0.146 | 31.511 | 20.058 | 0.208 | 20.8
[6 empty] | -0.146 | 31.501 | 20.058 | 0.202 | 20.2
[6 empty] | -0.146 | 31.491 | 20.058 | 0.196 | 19.6
[6 empty] | -0.146 | 31.481 | 20.058 | 0.189 | 18.9
[6 empty] | -0.146 | 31.471 | 20.058 | 0.183 | 18.3
[6 empty] | -0.146 | 31.461 | 20.058 | 0.177 | 17.7
[6 empty] | -0.146 | 31.451 | 20.058 | 0.171 | 17.1
[6 empty] | -0.146 | 31.441 | 20.058 | 0.165 | 16.5
[6 empty] | -0.146 | 31.431 | 20.058 | 0.159 | 15.9
[6 empty] | -0.146 | 31.421 | 20.058 | 0.153 | 15.3
[6 empty] | -0.146 | 31.411 | 20.058 | 0.147 | 14.7
[6 empty] | -0.146 | 31.401 | 20.058 | 0.142 | 14.2
[6 empty] | -0.146 | 31.391 | 20.058 | 0.143 | 14.3
[6 empty] | -0.146 | 31.381 | 20.058 | 0.143 | 14.3
[6 empty] | -0.146 | 31.371 | 20.058 | 0.143 | 14.3
[6 empty] | -0.146 | 31.361 | 20.058 | 0.144 | 14.4
[6 empty] | -0.146 | 31.351 | 20.058 | 0.144 | 14.4
[6 empty] | -0.146 | 31.341 | 20.058 | 0.144 | 14.4
[6 empty] | -0.146 | 31.331 | 20.058 | 0.145 | 14.5
[6 empty] | -0.146 | 31.321 | 20.058 | 0.145 | 14.5
[6 empty] | -0.146 | 31.311 | 20.058 | 0.145 | 14.5
[6 empty] | -0.146 | 31.301 | 20.058 | 0.145 | 14.5
[6 empty] | -0.146 | 31.291 | 20.058 | 0.146 | 14.6
[6 empty] | -0.146 | 31.281 | 20.058 | 0.146 | 14.6
[6 empty] | -0.146 | 31.271 | 20.058 | 0.146 | 14.6
[6 empty] | -0.146 | 31.261 | 20.058 | 0.147 | 14.7
[6 empty] | -0.146 | 31.251 | 20.058 | 0.147 | 14.7
[6 empty] | -0.146 | 31.241 | 20.058 | 0.147 | 14.7
[6 empty] | -0.146 | 31.231 | 20.058 | 0.148 | 14.8
[6 empty] | -0.146 | 31.221 | 20.058 | 0.148 | 14.8
[6 empty] | -0.146 | 31.211 | 20.058 | 0.148 | 14.8
[6 empty] | -0.146 | 31.201 | 20.058 | 0.151 | 15.1
[6 empty] | -0.146 | 31.191 | 20.058 | 0.158 | 15.8
[6 empty] | -0.146 | 31.181 | 20.058 | 0.166 | 16.6
[6 empty] | -0.146 | 31.171 | 20.058 | 0.174 | 17.4
[6 empty] | -0.146 | 31.161 | 20.058 | 0.181 | 18.1
[6 empty] | -0.146 | 31.151 | 20.058 | 0.189 | 18.9
[6 empty] | -0.146 | 31.141 | 20.058 | 0.197 | 19.7
[6 empty] | -0.146 | 31.131 | 20.058 | 0.204 | 20.4
[6 empty] | -0.146 | 31.121 | 20.058 | 0.212 | 21.2
[6 empty] | -0.146 | 31.111 | 20.058 | 0.22 | 22
[6 empty] | -0.146 | 31.101 | 20.058 | 0.228 | 22.8
[6 empty] | -0.146 | 31.091 | 20.058 | 0.235 | 23.5
[6 empty] | -0.146 | 31.081 | 20.058 | 0.243 | 24.3
[6 empty] | -0.146 | 31.071 | 20.058 | 0.251 | 25.1
[6 empty] | -0.146 | 31.061 | 20.058 | 0.258 | 25.8
[6 empty] | -0.146 | 31.051 | 20.058 | 0.266 | 26.6
[6 empty] | -0.146 | 31.041 | 20.058 | 0.274 | 27.4
[6 empty] | -0.146 | 31.031 | 20.057 | 0.281 | 28.1
[6 empty] | -0.146 | 31.021 | 20.057 | 0.289 | 28.9
[6 empty] | -0.146 | 31.011 | 20.057 | 0.297 | 29.7
[6 empty] | -0.146 | 31.001 | 20.057 | 0.304 | 30.4
[6 empty] | -0.146 | 30.991 | 20.057 | 0.309 | 30.9
[6 empty] | -0.146 | 30.981 | 20.057 | 0.315 | 31.5
[6 empty] | -0.146 | 30.972 | 20.057 | 0.32 | 32
[6 empty] | -0.146 | 30.962 | 20.057 | 0.325 | 32.5
[6 empty] | -0.146 | 30.952 | 20.057 | 0.331 | 33.1
[6 empty] | -0.146 | 30.942 | 20.057 | 0.336 | 33.6
[6 empty] | -0.146 | 30.932 | 20.057 | 0.341 | 34.1
[6 empty] | -0.146 | 30.922 | 20.057 | 0.347 | 34.7
[6 empty] | -0.146 | 30.912 | 20.057 | 0.352 | 35.2
[6 empty] | -0.146 | 30.902 | 20.057 | 0.357 | 35.7
[6 empty] | -0.146 | 30.892 | 20.057 | 0.363 | 36.3
[6 empty] | -0.146 | 30.882 | 20.057 | 0.368 | 36.8
[6 empty] | -0.146 | 30.872 | 20.057 | 0.374 | 37.4
[6 empty] | -0.146 | 30.862 | 20.057 | 0.379 | 37.9
[6 empty] | -0.146 | 30.852 | 20.057 | 0.384 | 38.4
[6 empty] | -0.146 | 30.842 | 20.057 | 0.39 | 39
[6 empty] | -0.146 | 30.832 | 20.057 | 0.395 | 39.5
[6 empty] | -0.146 | 30.822 | 20.057 | 0.4 | 40
[6 empty] | -0.146 | 30.812 | 20.057 | 0.406 | 40.6
[6 empty] | -0.146 | 30.802 | 20.057 | 0.409 | 40.9
[6 empty] | -0.146 | 30.792 | 20.057 | 0.407 | 40.7
[6 empty] | -0.146 | 30.782 | 20.057 | 0.405 | 40.5
[6 empty] | -0.146 | 30.772 | 20.057 | 0.402 | 40.2
[6 empty] | -0.146 | 30.762 | 20.057 | 0.4 | 40
[6 empty] | -0.146 | 30.752 | 20.057 | 0.398 | 39.8
[6 empty] | -0.146 | 30.742 | 20.057 | 0.396 | 39.6
[6 empty] | -0.146 | 30.732 | 20.057 | 0.394 | 39.4
[6 empty] | -0.146 | 30.722 | 20.057 | 0.391 | 39.1
[6 empty] | -0.146 | 30.712 | 20.057 | 0.389 | 38.9
[6 empty] | -0.146 | 30.702 | 20.057 | 0.387 | 38.7
[6 empty] | -0.146 | 30.692 | 20.057 | 0.385 | 38.5
[6 empty] | -0.146 | 30.682 | 20.057 | 0.382 | 38.2
[6 empty] | -0.146 | 30.672 | 20.057 | 0.38 | 38
[6 empty] | -0.146 | 30.662 | 20.057 | 0.378 | 37.8
[6 empty] | -0.146 | 30.652 | 20.057 | 0.376 | 37.6
[6 empty] | -0.146 | 30.642 | 20.057 | 0.373 | 37.3
[6 empty] | -0.146 | 30.632 | 20.057 | 0.371 | 37.1
[6 empty] | -0.146 | 30.622 | 20.057 | 0.369 | 36.9
[6 empty] | -0.146 | 30.612 | 20.057 | 0.367 | 36.7
[6 empty] | -0.146 | 30.602 | 20.057 | 0.363 | 36.3
[6 empty] | -0.146 | 30.592 | 20.057 | 0.353 | 35.3
[6 empty] | -0.146 | 30.582 | 20.057 | 0.344 | 34.4
[6 empty] | -0.146 | 30.572 | 20.057 | 0.334 | 33.4
[6 empty] | -0.146 | 30.562 | 20.057 | 0.325 | 32.5
[6 empty] | -0.146 | 30.552 | 20.057 | 0.316 | 31.6
[6 empty] | -0.146 | 30.542 | 20.057 | 0.306 | 30.6
[6 empty] | -0.146 | 30.532 | 20.057 | 0.297 | 29.7
[6 empty] | -0.146 | 30.522 | 20.057 | 0.287 | 28.7
[6 empty] | -0.146 | 30.512 | 20.057 | 0.278 | 27.8
[6 empty] | -0.146 | 30.502 | 20.057 | 0.268 | 26.8
[6 empty] | -0.146 | 30.492 | 20.057 | 0.259 | 25.9
[6 empty] | -0.146 | 30.482 | 20.057 | 0.25 | 25
[6 empty] | -0.146 | 30.472 | 20.057 | 0.24 | 24
[6 empty] | -0.146 | 30.462 | 20.057 | 0.231 | 23.1
[6 empty] | -0.146 | 30.452 | 20.057 | 0.221 | 22.1
[6 empty] | -0.146 | 30.442 | 20.057 | 0.212 | 21.2
[6 empty] | -0.146 | 30.432 | 20.057 | 0.202 | 20.2
[6 empty] | -0.146 | 30.422 | 20.057 | 0.193 | 19.3
[6 empty] | -0.146 | 30.412 | 20.057 | 0.184 | 18.4
[6 empty] | -0.146 | 30.402 | 20.057 | 0.176 | 17.6
[6 empty] | -0.146 | 30.392 | 20.057 | 0.174 | 17.4
[6 empty] | -0.146 | 30.382 | 20.057 | 0.172 | 17.2
[6 empty] | -0.146 | 30.372 | 20.057 | 0.17 | 17
[6 empty] | -0.146 | 30.362 | 20.057 | 0.168 | 16.8
[6 empty] | -0.146 | 30.352 | 20.057 | 0.166 | 16.6
[6 empty] | -0.146 | 30.342 | 20.056 | 0.164 | 16.4
[6 empty] | -0.146 | 30.332 | 20.056 | 0.162 | 16.2
[6 empty] | -0.146 | 30.322 | 20.056 | 0.16 | 16
[6 empty] | -0.146 | 30.312 | 20.056 | 0.158 | 15.8
[6 empty] | -0.146 | 30.302 | 20.056 | 0.156 | 15.6
[6 empty] | -0.146 | 30.292 | 20.056 | 0.154 | 15.4
[6 empty] | -0.146 | 30.282 | 20.056 | 0.152 | 15.2
[6 empty] | -0.146 | 30.272 | 20.056 | 0.15 | 15
[6 empty] | -0.146 | 30.262 | 20.056 | 0.148 | 14.8
[6 empty] | -0.146 | 30.252 | 20.056 | 0.146 | 14.6
[6 empty] | -0.146 | 30.242 | 20.056 | 0.144 | 14.4
[6 empty] | -0.146 | 30.232 | 20.056 | 0.142 | 14.2
[6 empty] | -0.146 | 30.222 | 20.056 | 0.14 | 14
[6 empty] | -0.146 | 30.212 | 20.056 | 0.138 | 13.8
[6 empty] | -0.146 | 30.202 | 20.056 | 0.137 | 13.7
[6 empty] | -0.146 | 30.192 | 20.056 | 0.141 | 14.1
[6 empty] | -0.146 | 30.182 | 20.056 | 0.145 | 14.5
[6 empty] | -0.146 | 30.172 | 20.056 | 0.148 | 14.8
[6 empty] | -0.146 | 30.162 | 20.056 | 0.152 | 15.2
[6 empty] | -0.146 | 30.152 | 20.056 | 0.155 | 15.5
[6 empty] | -0.146 | 30.142 | 20.056 | 0.159 | 15.9
[6 empty] | -0.146 | 30.132 | 20.056 | 0.163 | 16.3
[6 empty] | -0.146 | 30.122 | 20.056 | 0.166 | 16.6
[6 empty] | -0.146 | 30.112 | 20.056 | 0.17 | 17
[6 empty] | -0.146 | 30.102 | 20.056 | 0.173 | 17.3
[6 empty] | -0.146 | 30.092 | 20.056 | 0.177 | 17.7
[6 empty] | -0.146 | 30.082 | 20.056 | 0.181 | 18.1
[6 empty] | -0.146 | 30.072 | 20.056 | 0.184 | 18.4
[6 empty] | -0.146 | 30.062 | 20.056 | 0.188 | 18.8
[6 empty] | -0.146 | 30.052 | 20.056 | 0.191 | 19.1
[6 empty] | -0.146 | 30.042 | 20.056 | 0.195 | 19.5
[6 empty] | -0.146 | 30.032 | 20.056 | 0.199 | 19.9
[6 empty] | -0.146 | 30.022 | 20.056 | 0.202 | 20.2
[6 empty] | -0.146 | 30.012 | 20.056 | 0.206 | 20.6
[6 empty] | -0.146 | 30.002 | 20.056 | 0.21 | 21
[6 empty] | -0.146 | 29.992 | 20.056 | 0.219 | 21.9
[6 empty] | -0.146 | 29.982 | 20.056 | 0.228 | 22.8
[6 empty] | -0.146 | 29.972 | 20.056 | 0.237 | 23.7
[6 empty] | -0.146 | 29.962 | 20.056 | 0.245 | 24.5
[6 empty] | -0.146 | 29.952 | 20.056 | 0.254 | 25.4
[6 empty] | -0.146 | 29.942 | 20.056 | 0.263 | 26.3
[6 empty] | -0.146 | 29.932 | 20.056 | 0.271 | 27.1
[6 empty] | -0.146 | 29.922 | 20.056 | 0.28 | 28
[6 empty] | -0.146 | 29.912 | 20.056 | 0.289 | 28.9
[6 empty] | -0.146 | 29.902 | 20.056 | 0.298 | 29.8
[6 empty] | -0.146 | 29.892 | 20.056 | 0.306 | 30.6
[6 empty] | -0.146 | 29.882 | 20.056 | 0.315 | 31.5
[6 empty] | -0.146 | 29.872 | 20.056 | 0.324 | 32.4
[6 empty] | -0.146 | 29.862 | 20.056 | 0.332 | 33.2
[6 empty] | -0.146 | 29.852 | 20.056 | 0.341 | 34.1
[6 empty] | -0.146 | 29.842 | 20.056 | 0.35 | 35
[6 empty] | -0.146 | 29.832 | 20.056 | 0.358 | 35.8
[6 empty] | -0.146 | 29.822 | 20.056 | 0.367 | 36.7
[6 empty] | -0.146 | 29.812 | 20.056 | 0.376 | 37.6
[6 empty] | -0.146 | 29.802 | 20.056 | 0.383 | 38.3
[6 empty] | -0.146 | 29.792 | 20.056 | 0.384 | 38.4
[6 empty] | -0.146 | 29.782 | 20.056 | 0.385 | 38.5
[6 empty] | -0.146 | 29.772 | 20.056 | 0.386 | 38.6
[6 empty] | -0.146 | 29.762 | 20.056 | 0.387 | 38.7
[6 empty] | -0.146 | 29.752 | 20.056 | 0.389 | 38.9
[6 empty] | -0.146 | 29.742 | 20.056 | 0.39 | 39
[6 empty] | -0.146 | 29.732 | 20.056 | 0.391 | 39.1
[6 empty] | -0.146 | 29.722 | 20.056 | 0.392 | 39.2
[6 empty] | -0.146 | 29.712 | 20.056 | 0.393 | 39.3
[6 empty] | -0.146 | 29.702 | 20.056 | 0.394 | 39.4
[6 empty] | -0.146 | 29.692 | 20.056 | 0.395 | 39.5
[6 empty] | -0.146 | 29.682 | 20.056 | 0.397 | 39.7
[6 empty] | -0.146 | 29.672 | 20.056 | 0.398 | 39.8
[6 empty] | -0.146 | 29.662 | 20.056 | 0.399 | 39.9
[6 empty] | -0.146 | 29.652 | 20.055 | 0.4 | 40
[6 empty] | -0.146 | 29.642 | 20.055 | 0.401 | 40.1
[6 empty] | -0.146 | 29.632 | 20.055 | 0.402 | 40.2
[6 empty] | -0.146 | 29.622 | 20.055 | 0.403 | 40.3
[6 empty] | -0.146 | 29.612 | 20.055 | 0.405 | 40.5
[6 empty] | -0.146 | 29.602 | 20.055 | 0.404 | 40.4
[6 empty] | -0.146 | 29.592 | 20.055 | 0.397 | 39.7
[6 empty] | -0.146 | 29.582 | 20.055 | 0.39 | 39
[6 empty] | -0.146 | 29.572 | 20.055 | 0.383 | 38.3
[6 empty] | -0.146 | 29.562 | 20.055 | 0.376 | 37.6
[6 empty] | -0.146 | 29.552 | 20.055 | 0.369 | 36.9
[6 empty] | -0.146 | 29.542 | 20.055 | 0.362 | 36.2
[6 empty] | -0.146 | 29.532 | 20.055 | 0.355 | 35.5
[6 empty] | -0.146 | 29.522 | 20.055 | 0.348 | 34.8
[6 empty] | -0.146 | 29.512 | 20.055 | 0.341 | 34.1
[6 empty] | -0.146 | 29.502 | 20.055 | 0.334 | 33.4
[6 empty] | -0.146 | 29.492 | 20.055 | 0.327 | 32.7
[6 empty] | -0.146 | 29.482 | 20.055 | 0.32 | 32
[6 empty] | -0.146 | 29.472 | 20.055 | 0.313 | 31.3
[6 empty] | -0.146 | 29.462 | 20.055 | 0.306 | 30.6
[6 empty] | -0.146 | 29.452 | 20.055 | 0.299 | 29.9
[6 empty] | -0.146 | 29.442 | 20.055 | 0.292 | 29.2
[6 empty] | -0.146 | 29.432 | 20.055 | 0.285 | 28.5
[6 empty] | -0.146 | 29.422 | 20.055 | 0.278 | 27.8
[6 empty] | -0.146 | 29.412 | 20.055 | 0.271 | 27.1
[6 empty] | -0.146 | 29.402 | 20.055 | 0.264 | 26.4
[6 empty] | -0.146 | 29.392 | 20.055 | 0.258 | 25.8
[6 empty] | -0.146 | 29.382 | 20.055 | 0.252 | 25.2
[6 empty] | -0.146 | 29.372 | 20.055 | 0.246 | 24.6
[6 empty] | -0.146 | 29.362 | 20.055 | 0.241 | 24.1
[6 empty] | -0.146 | 29.352 | 20.055 | 0.235 | 23.5
[6 empty] | -0.146 | 29.342 | 20.055 | 0.229 | 22.9
[6 empty] | -0.146 | 29.332 | 20.055 | 0.223 | 22.3
[6 empty] | -0.146 | 29.322 | 20.055 | 0.217 | 21.7
[6 empty] | -0.146 | 29.312 | 20.055 | 0.211 | 21.1
[6 empty] | -0.146 | 29.302 | 20.055 | 0.205 | 20.5
[6 empty] | -0.146 | 29.292 | 20.055 | 0.199 | 19.9
[6 empty] | -0.146 | 29.282 | 20.055 | 0.193 | 19.3
[6 empty] | -0.146 | 29.272 | 20.055 | 0.187 | 18.7
[6 empty] | -0.146 | 29.262 | 20.055 | 0.181 | 18.1
[6 empty] | -0.146 | 29.252 | 20.055 | 0.175 | 17.5
[6 empty] | -0.146 | 29.242 | 20.055 | 0.169 | 16.9
[6 empty] | -0.146 | 29.232 | 20.055 | 0.163 | 16.3
[6 empty] | -0.146 | 29.222 | 20.055 | 0.157 | 15.7
[6 empty] | -0.146 | 29.212 | 20.055 | 0.151 | 15.1
[6 empty] | -0.146 | 29.202 | 20.055 | 0.147 | 14.7
[6 empty] | -0.146 | 29.192 | 20.055 | 0.147 | 14.7
[6 empty] | -0.146 | 29.182 | 20.055 | 0.148 | 14.8
[6 empty] | -0.146 | 29.172 | 20.055 | 0.148 | 14.8
[6 empty] | -0.146 | 29.162 | 20.055 | 0.148 | 14.8
[6 empty] | -0.146 | 29.152 | 20.055 | 0.149 | 14.9
[6 empty] | -0.146 | 29.142 | 20.055 | 0.149 | 14.9
[6 empty] | -0.146 | 29.132 | 20.055 | 0.15 | 15
[6 empty] | -0.146 | 29.122 | 20.055 | 0.15 | 15
[6 empty] | -0.146 | 29.112 | 20.055 | 0.15 | 15
[6 empty] | -0.146 | 29.102 | 20.055 | 0.151 | 15.1
[6 empty] | -0.146 | 29.092 | 20.055 | 0.151 | 15.1
[6 empty] | -0.146 | 29.082 | 20.055 | 0.152 | 15.2
[6 empty] | -0.146 | 29.072 | 20.055 | 0.152 | 15.2
[6 empty] | -0.146 | 29.062 | 20.055 | 0.153 | 15.3
[6 empty] | -0.146 | 29.052 | 20.055 | 0.153 | 15.3
[6 empty] | -0.146 | 29.042 | 20.055 | 0.153 | 15.3
[6 empty] | -0.146 | 29.032 | 20.055 | 0.154 | 15.4
[6 empty] | -0.146 | 29.022 | 20.055 | 0.154 | 15.4
[6 empty] | -0.146 | 29.012 | 20.055 | 0.155 | 15.5
[6 empty] | -0.146 | 29.002 | 20.055 | 0.157 | 15.7
[6 empty] | -0.146 | 28.992 | 20.055 | 0.164 | 16.4
[6 empty] | -0.146 | 28.982 | 20.055 | 0.172 | 17.2
[6 empty] | -0.146 | 28.972 | 20.055 | 0.18 | 18
[6 empty] | -0.146 | 28.962 | 20.054 | 0.187 | 18.7
[6 empty] | -0.146 | 28.952 | 20.054 | 0.195 | 19.5
[6 empty] | -0.146 | 28.942 | 20.054 | 0.202 | 20.2
[6 empty] | -0.146 | 28.932 | 20.054 | 0.21 | 21
[6 empty] | -0.146 | 28.922 | 20.054 | 0.218 | 21.8
[6 empty] | -0.146 | 28.912 | 20.054 | 0.225 | 22.5
[6 empty] | -0.146 | 28.902 | 20.054 | 0.233 | 23.3
[6 empty] | -0.146 | 28.892 | 20.054 | 0.241 | 24.1
[6 empty] | -0.146 | 28.882 | 20.054 | 0.248 | 24.8
[6 empty] | -0.146 | 28.872 | 20.054 | 0.256 | 25.6
[6 empty] | -0.146 | 28.862 | 20.054 | 0.264 | 26.4
[6 empty] | -0.146 | 28.852 | 20.054 | 0.271 | 27.1
[6 empty] | -0.146 | 28.842 | 20.054 | 0.279 | 27.9
[6 empty] | -0.146 | 28.832 | 20.054 | 0.287 | 28.7
[6 empty] | -0.146 | 28.822 | 20.054 | 0.294 | 29.4
[6 empty] | -0.146 | 28.812 | 20.054 | 0.302 | 30.2
[6 empty] | -0.146 | 28.802 | 20.054 | 0.309 | 30.9
[6 empty] | -0.146 | 28.792 | 20.054 | 0.314 | 31.4
[6 empty] | -0.146 | 28.782 | 20.054 | 0.319 | 31.9
[6 empty] | -0.146 | 28.772 | 20.054 | 0.325 | 32.5
[6 empty] | -0.146 | 28.762 | 20.054 | 0.33 | 33
[6 empty] | -0.146 | 28.752 | 20.054 | 0.335 | 33.5
[6 empty] | -0.146 | 28.742 | 20.054 | 0.34 | 34
[6 empty] | -0.146 | 28.732 | 20.054 | 0.345 | 34.5
[6 empty] | -0.146 | 28.722 | 20.054 | 0.35 | 35
[6 empty] | -0.146 | 28.712 | 20.054 | 0.356 | 35.6
[6 empty] | -0.146 | 28.702 | 20.054 | 0.361 | 36.1
[6 empty] | -0.146 | 28.692 | 20.054 | 0.366 | 36.6
[6 empty] | -0.146 | 28.682 | 20.054 | 0.371 | 37.1
[6 empty] | -0.146 | 28.672 | 20.054 | 0.376 | 37.6
[6 empty] | -0.146 | 28.662 | 20.054 | 0.381 | 38.1
[6 empty] | -0.146 | 28.652 | 20.054 | 0.387 | 38.7
[6 empty] | -0.146 | 28.642 | 20.054 | 0.392 | 39.2
[6 empty] | -0.146 | 28.632 | 20.054 | 0.397 | 39.7
[6 empty] | -0.146 | 28.622 | 20.054 | 0.402 | 40.2
[6 empty] | -0.146 | 28.612 | 20.054 | 0.407 | 40.7
[6 empty] | -0.146 | 28.602 | 20.054 | 0.411 | 41.1
[6 empty] | -0.146 | 28.592 | 20.054 | 0.409 | 40.9
[6 empty] | -0.146 | 28.582 | 20.054 | 0.406 | 40.6
[6 empty] | -0.146 | 28.572 | 20.054 | 0.404 | 40.4
[6 empty] | -0.146 | 28.562 | 20.054 | 0.402 | 40.2
[6 empty] | -0.146 | 28.552 | 20.054 | 0.399 | 39.9
[6 empty] | -0.146 | 28.542 | 20.054 | 0.397 | 39.7
[6 empty] | -0.146 | 28.532 | 20.054 | 0.395 | 39.5
[6 empty] | -0.146 | 28.522 | 20.054 | 0.392 | 39.2
[6 empty] | -0.146 | 28.512 | 20.054 | 0.39 | 39
[6 empty] | -0.146 | 28.502 | 20.054 | 0.388 | 38.8
[6 empty] | -0.146 | 28.492 | 20.054 | 0.385 | 38.5
[6 empty] | -0.146 | 28.482 | 20.054 | 0.383 | 38.3
[6 empty] | -0.146 | 28.472 | 20.054 | 0.381 | 38.1
[6 empty] | -0.146 | 28.462 | 20.054 | 0.378 | 37.8
[6 empty] | -0.146 | 28.452 | 20.054 | 0.376 | 37.6
[6 empty] | -0.146 | 28.442 | 20.054 | 0.374 | 37.4
[6 empty] | -0.146 | 28.432 | 20.054 | 0.372 | 37.2
[6 empty] | -0.146 | 28.422 | 20.054 | 0.369 | 36.9
[6 empty] | -0.146 | 28.412 | 20.054 | 0.367 | 36.7
[6 empty] | -0.146 | 28.402 | 20.054 | 0.363 | 36.3
[6 empty] | -0.146 | 28.392 | 20.054 | 0.354 | 35.4
[6 empty] | -0.146 | 28.382 | 20.054 | 0.344 | 34.4
[6 empty] | -0.146 | 28.372 | 20.054 | 0.335 | 33.5
[6 empty] | -0.146 | 28.362 | 20.054 | 0.326 | 32.6
[6 empty] | -0.146 | 28.352 | 20.054 | 0.316 | 31.6
[6 empty] | -0.146 | 28.342 | 20.054 | 0.307 | 30.7
[6 empty] | -0.146 | 28.332 | 20.054 | 0.297 | 29.7
[6 empty] | -0.146 | 28.322 | 20.054 | 0.288 | 28.8
[6 empty] | -0.146 | 28.312 | 20.054 | 0.279 | 27.9
[6 empty] | -0.146 | 28.302 | 20.054 | 0.269 | 26.9
[6 empty] | -0.146 | 28.292 | 20.054 | 0.26 | 26
[6 empty] | -0.146 | 28.282 | 20.054 | 0.251 | 25.1
[6 empty] | -0.146 | 28.272 | 20.053 | 0.241 | 24.1
[6 empty] | -0.146 | 28.262 | 20.053 | 0.232 | 23.2
[6 empty] | -0.146 | 28.252 | 20.053 | 0.222 | 22.2
[6 empty] | -0.146 | 28.242 | 20.053 | 0.213 | 21.3
[6 empty] | -0.146 | 28.232 | 20.053 | 0.204 | 20.4
[6 empty] | -0.146 | 28.222 | 20.053 | 0.194 | 19.4
[6 empty] | -0.146 | 28.212 | 20.053 | 0.185 | 18.5
[6 empty] | -0.146 | 28.202 | 20.053 | 0.177 | 17.7
[6 empty] | -0.146 | 28.192 | 20.053 | 0.175 | 17.5
[6 empty] | -0.146 | 28.182 | 20.053 | 0.173 | 17.3
[6 empty] | -0.146 | 28.172 | 20.053 | 0.172 | 17.2
[6 empty] | -0.146 | 28.162 | 20.053 | 0.17 | 17
[6 empty] | -0.146 | 28.152 | 20.053 | 0.168 | 16.8
[6 empty] | -0.146 | 28.142 | 20.053 | 0.166 | 16.6
[6 empty] | -0.146 | 28.132 | 20.053 | 0.164 | 16.4
[6 empty] | -0.146 | 28.122 | 20.053 | 0.162 | 16.2
[6 empty] | -0.146 | 28.112 | 20.053 | 0.16 | 16
[6 empty] | -0.146 | 28.102 | 20.053 | 0.159 | 15.9
[6 empty] | -0.146 | 28.092 | 20.053 | 0.157 | 15.7
[6 empty] | -0.146 | 28.082 | 20.053 | 0.155 | 15.5
[6 empty] | -0.146 | 28.072 | 20.053 | 0.153 | 15.3
[6 empty] | -0.146 | 28.062 | 20.053 | 0.151 | 15.1
[6 empty] | -0.146 | 28.052 | 20.053 | 0.149 | 14.9
[6 empty] | -0.146 | 28.042 | 20.053 | 0.148 | 14.8
[6 empty] | -0.146 | 28.032 | 20.053 | 0.146 | 14.6
[6 empty] | -0.146 | 28.022 | 20.053 | 0.144 | 14.4
[6 empty] | -0.146 | 28.012 | 20.053 | 0.142 | 14.2
[6 empty] | -0.146 | 28.002 | 20.053 | 0.141 | 14.1
[6 empty] | -0.146 | 27.992 | 20.053 | 0.145 | 14.5
[6 empty] | -0.146 | 27.982 | 20.053 | 0.149 | 14.9
[6 empty] | -0.146 | 27.973 | 20.053 | 0.152 | 15.2
[6 empty] | -0.146 | 27.963 | 20.053 | 0.156 | 15.6
[6 empty] | -0.146 | 27.953 | 20.053 | 0.159 | 15.9
[6 empty] | -0.146 | 27.943 | 20.053 | 0.163 | 16.3
[6 empty] | -0.146 | 27.933 | 20.053 | 0.167 | 16.7
[6 empty] | -0.146 | 27.923 | 20.053 | 0.17 | 17
[6 empty] | -0.146 | 27.913 | 20.053 | 0.174 | 17.4
[6 empty] | -0.146 | 27.903 | 20.053 | 0.178 | 17.8
[6 empty] | -0.146 | 27.893 | 20.053 | 0.181 | 18.1
[6 empty] | -0.146 | 27.883 | 20.053 | 0.185 | 18.5
[6 empty] | -0.146 | 27.873 | 20.053 | 0.189 | 18.9
[6 empty] | -0.146 | 27.863 | 20.053 | 0.192 | 19.2
[6 empty] | -0.146 | 27.853 | 20.053 | 0.196 | 19.6
[6 empty] | -0.146 | 27.843 | 20.053 | 0.2 | 20
[6 empty] | -0.146 | 27.833 | 20.053 | 0.203 | 20.3
[6 empty] | -0.146 | 27.823 | 20.053 | 0.207 | 20.7
[6 empty] | -0.146 | 27.813 | 20.053 | 0.21 | 21
[6 empty] | -0.146 | 27.803 | 20.053 | 0.215 | 21.5
[6 empty] | -0.146 | 27.793 | 20.053 | 0.224 | 22.4
[6 empty] | -0.146 | 27.783 | 20.053 | 0.232 | 23.2
[6 empty] | -0.146 | 27.773 | 20.053 | 0.241 | 24.1
[6 empty] | -0.146 | 27.763 | 20.053 | 0.249 | 24.9
[6 empty] | -0.146 | 27.753 | 20.053 | 0.258 | 25.8
[6 empty] | -0.146 | 27.743 | 20.053 | 0.266 | 26.6
[6 empty] | -0.146 | 27.733 | 20.053 | 0.275 | 27.5
[6 empty] | -0.146 | 27.723 | 20.053 | 0.283 | 28.3
[6 empty] | -0.146 | 27.713 | 20.053 | 0.292 | 29.2
[6 empty] | -0.146 | 27.703 | 20.053 | 0.3 | 30
[6 empty] | -0.146 | 27.693 | 20.053 | 0.309 | 30.9
[6 empty] | -0.146 | 27.683 | 20.053 | 0.317 | 31.7
[6 empty] | -0.146 | 27.673 | 20.053 | 0.326 | 32.6
[6 empty] | -0.146 | 27.663 | 20.053 | 0.334 | 33.4
[6 empty] | -0.146 | 27.653 | 20.053 | 0.343 | 34.3
[6 empty] | -0.146 | 27.643 | 20.053 | 0.351 | 35.1
[6 empty] | -0.146 | 27.633 | 20.053 | 0.36 | 36
[6 empty] | -0.146 | 27.623 | 20.053 | 0.368 | 36.8
[6 empty] | -0.146 | 27.613 | 20.053 | 0.377 | 37.7
[6 empty] | -0.146 | 27.603 | 20.053 | 0.384 | 38.4
[6 empty] | -0.146 | 27.593 | 20.053 | 0.385 | 38.5
[6 empty] | -0.146 | 27.583 | 20.052 | 0.386 | 38.6
[6 empty] | -0.146 | 27.573 | 20.052 | 0.387 | 38.7
[6 empty] | -0.146 | 27.563 | 20.052 | 0.388 | 38.8
[6 empty] | -0.146 | 27.553 | 20.052 | 0.389 | 38.9
[6 empty] | -0.146 | 27.543 | 20.052 | 0.39 | 39
[6 empty] | -0.146 | 27.533 | 20.052 | 0.391 | 39.1
[6 empty] | -0.146 | 27.523 | 20.052 | 0.392 | 39.2
[6 empty] | -0.146 | 27.513 | 20.052 | 0.393 | 39.3
[6 empty] | -0.146 | 27.503 | 20.052 | 0.394 | 39.4
[6 empty] | -0.146 | 27.493 | 20.052 | 0.395 | 39.5
[6 empty] | -0.146 | 27.483 | 20.052 | 0.396 | 39.6
[6 empty] | -0.146 | 27.473 | 20.052 | 0.397 | 39.7
[6 empty] | -0.146 | 27.463 | 20.052 | 0.398 | 39.8
[6 empty] | -0.146 | 27.453 | 20.052 | 0.399 | 39.9
[6 empty] | -0.146 | 27.443 | 20.052 | 0.4 | 40
[6 empty] | -0.146 | 27.433 | 20.052 | 0.401 | 40.1
[6 empty] | -0.146 | 27.423 | 20.052 | 0.402 | 40.2
[6 empty] | -0.146 | 27.413 | 20.052 | 0.403 | 40.3
[6 empty] | -0.146 | 27.403 | 20.052 | 0.403 | 40.3
[6 empty] | -0.146 | 27.393 | 20.052 | 0.396 | 39.6
[6 empty] | -0.146 | 27.383 | 20.052 | 0.389 | 38.9
[6 empty] | -0.146 | 27.373 | 20.052 | 0.382 | 38.2
[6 empty] | -0.146 | 27.363 | 20.052 | 0.375 | 37.5
[6 empty] | -0.146 | 27.353 | 20.052 | 0.368 | 36.8
[6 empty] | -0.146 | 27.343 | 20.052 | 0.361 | 36.1
[6 empty] | -0.146 | 27.333 | 20.052 | 0.354 | 35.4
[6 empty] | -0.146 | 27.323 | 20.052 | 0.347 | 34.7
[6 empty] | -0.146 | 27.313 | 20.052 | 0.34 | 34
[6 empty] | -0.146 | 27.303 | 20.052 | 0.333 | 33.3
[6 empty] | -0.146 | 27.293 | 20.052 | 0.326 | 32.6
[6 empty] | -0.146 | 27.283 | 20.052 | 0.319 | 31.9
[6 empty] | -0.146 | 27.273 | 20.052 | 0.312 | 31.2
[6 empty] | -0.146 | 27.263 | 20.052 | 0.305 | 30.5
[6 empty] | -0.146 | 27.253 | 20.052 | 0.298 | 29.8
[6 empty] | -0.146 | 27.243 | 20.052 | 0.291 | 29.1
[6 empty] | -0.146 | 27.233 | 20.052 | 0.284 | 28.4
[6 empty] | -0.146 | 27.223 | 20.052 | 0.277 | 27.7
[6 empty] | -0.146 | 27.213 | 20.052 | 0.27 | 27
[6 empty] | -0.146 | 27.203 | 20.052 | 0.263 | 26.3
[6 empty] | -0.146 | 27.193 | 20.052 | 0.257 | 25.7
[6 empty] | -0.146 | 27.183 | 20.052 | 0.251 | 25.1
[6 empty] | -0.146 | 27.173 | 20.052 | 0.246 | 24.6
[6 empty] | -0.146 | 27.163 | 20.052 | 0.24 | 24
[6 empty] | -0.146 | 27.153 | 20.052 | 0.234 | 23.4
[6 empty] | -0.146 | 27.143 | 20.052 | 0.228 | 22.8
[6 empty] | -0.146 | 27.133 | 20.052 | 0.222 | 22.2
[6 empty] | -0.146 | 27.123 | 20.052 | 0.217 | 21.7
[6 empty] | -0.146 | 27.113 | 20.052 | 0.211 | 21.1
[6 empty] | -0.146 | 27.103 | 20.052 | 0.205 | 20.5
[6 empty] | -0.146 | 27.093 | 20.052 | 0.199 | 19.9
[6 empty] | -0.146 | 27.083 | 20.052 | 0.193 | 19.3
[6 empty] | -0.146 | 27.073 | 20.052 | 0.188 | 18.8
[6 empty] | -0.146 | 27.063 | 20.052 | 0.182 | 18.2
[6 empty] | -0.146 | 27.053 | 20.052 | 0.176 | 17.6
[6 empty] | -0.146 | 27.043 | 20.052 | 0.17 | 17
[6 empty] | -0.146 | 27.033 | 20.052 | 0.164 | 16.4
[6 empty] | -0.146 | 27.023 | 20.052 | 0.159 | 15.9
[6 empty] | -0.146 | 27.013 | 20.052 | 0.153 | 15.3
[6 empty] | -0.146 | 27.003 | 20.052 | 0.148 | 14.8
[6 empty] | -0.146 | 26.993 | 20.052 | 0.149 | 14.9
[6 empty] | -0.146 | 26.983 | 20.052 | 0.149 | 14.9
[6 empty] | -0.146 | 26.973 | 20.052 | 0.15 | 15
[6 empty] | -0.146 | 26.963 | 20.052 | 0.15 | 15
[6 empty] | -0.146 | 26.953 | 20.052 | 0.151 | 15.1
[6 empty] | -0.146 | 26.943 | 20.052 | 0.151 | 15.1
[6 empty] | -0.146 | 26.933 | 20.052 | 0.152 | 15.2
[6 empty] | -0.146 | 26.923 | 20.052 | 0.152 | 15.2
[6 empty] | -0.146 | 26.913 | 20.052 | 0.153 | 15.3
[6 empty] | -0.146 | 26.903 | 20.052 | 0.153 | 15.3
[6 empty] | -0.146 | 26.893 | 20.051 | 0.153 | 15.3
[6 empty] | -0.146 | 26.883 | 20.051 | 0.154 | 15.4
[6 empty] | -0.146 | 26.873 | 20.051 | 0.154 | 15.4
[6 empty] | -0.146 | 26.863 | 20.051 | 0.155 | 15.5
[6 empty] | -0.146 | 26.853 | 20.051 | 0.155 | 15.5
[6 empty] | -0.146 | 26.843 | 20.051 | 0.156 | 15.6
[6 empty] | -0.146 | 26.833 | 20.051 | 0.156 | 15.6
[6 empty] | -0.146 | 26.823 | 20.051 | 0.157 | 15.7
[6 empty] | -0.146 | 26.813 | 20.051 | 0.157 | 15.7
[6 empty] | -0.146 | 26.803 | 20.051 | 0.159 | 15.9
[6 empty] | -0.146 | 26.793 | 20.051 | 0.167 | 16.7
[6 empty] | -0.146 | 26.783 | 20.051 | 0.174 | 17.4
[6 empty] | -0.146 | 26.773 | 20.051 | 0.182 | 18.2
[6 empty] | -0.146 | 26.763 | 20.051 | 0.19 | 19
[6 empty] | -0.146 | 26.753 | 20.051 | 0.197 | 19.7
[6 empty] | -0.146 | 26.743 | 20.051 | 0.205 | 20.5
[6 empty] | -0.146 | 26.733 | 20.051 | 0.212 | 21.2
[6 empty] | -0.146 | 26.723 | 20.051 | 0.22 | 22
[6 empty] | -0.146 | 26.713 | 20.051 | 0.228 | 22.8
[6 empty] | -0.146 | 26.703 | 20.051 | 0.235 | 23.5
[6 empty] | -0.146 | 26.693 | 20.051 | 0.243 | 24.3
[6 empty] | -0.146 | 26.683 | 20.051 | 0.251 | 25.1
[6 empty] | -0.146 | 26.673 | 20.051 | 0.258 | 25.8
[6 empty] | -0.146 | 26.663 | 20.051 | 0.266 | 26.6
[6 empty] | -0.146 | 26.653 | 20.051 | 0.273 | 27.3
[6 empty] | -0.146 | 26.643 | 20.051 | 0.281 | 28.1
[6 empty] | -0.146 | 26.633 | 20.051 | 0.289 | 28.9
[6 empty] | -0.146 | 26.623 | 20.051 | 0.296 | 29.6
[6 empty] | -0.146 | 26.613 | 20.051 | 0.304 | 30.4
[6 empty] | -0.146 | 26.603 | 20.051 | 0.311 | 31.1
[6 empty] | -0.146 | 26.593 | 20.051 | 0.316 | 31.6
[6 empty] | -0.146 | 26.583 | 20.051 | 0.321 | 32.1
[6 empty] | -0.146 | 26.573 | 20.051 | 0.326 | 32.6
[6 empty] | -0.146 | 26.563 | 20.051 | 0.331 | 33.1
[6 empty] | -0.146 | 26.553 | 20.051 | 0.336 | 33.6
[6 empty] | -0.146 | 26.543 | 20.051 | 0.341 | 34.1
[6 empty] | -0.146 | 26.533 | 20.051 | 0.346 | 34.6
[6 empty] | -0.146 | 26.523 | 20.051 | 0.351 | 35.1
[6 empty] | -0.146 | 26.513 | 20.051 | 0.356 | 35.6
[6 empty] | -0.146 | 26.503 | 20.051 | 0.361 | 36.1
[6 empty] | -0.146 | 26.493 | 20.051 | 0.366 | 36.6
[6 empty] | -0.146 | 26.483 | 20.051 | 0.371 | 37.1
[6 empty] | -0.146 | 26.473 | 20.051 | 0.375 | 37.5
[6 empty] | -0.146 | 26.463 | 20.051 | 0.38 | 38
[6 empty] | -0.146 | 26.453 | 20.051 | 0.385 | 38.5
[6 empty] | -0.146 | 26.443 | 20.051 | 0.39 | 39
[6 empty] | -0.146 | 26.433 | 20.051 | 0.395 | 39.5
[6 empty] | -0.146 | 26.423 | 20.051 | 0.4 | 40
[6 empty] | -0.146 | 26.413 | 20.051 | 0.405 | 40.5
[6 empty] | -0.146 | 26.403 | 20.051 | 0.409 | 40.9
[6 empty] | -0.146 | 26.393 | 20.051 | 0.406 | 40.6
[6 empty] | -0.146 | 26.383 | 20.051 | 0.404 | 40.4
[6 empty] | -0.146 | 26.373 | 20.051 | 0.402 | 40.2
[6 empty] | -0.146 | 26.363 | 20.051 | 0.399 | 39.9
[6 empty] | -0.146 | 26.353 | 20.051 | 0.397 | 39.7
[6 empty] | -0.146 | 26.343 | 20.051 | 0.394 | 39.4
[6 empty] | -0.146 | 26.333 | 20.051 | 0.392 | 39.2
[6 empty] | -0.146 | 26.323 | 20.051 | 0.389 | 38.9
[6 empty] | -0.146 | 26.313 | 20.051 | 0.387 | 38.7
[6 empty] | -0.146 | 26.303 | 20.051 | 0.385 | 38.5
[6 empty] | -0.146 | 26.293 | 20.051 | 0.382 | 38.2
[6 empty] | -0.146 | 26.283 | 20.051 | 0.38 | 38
[6 empty] | -0.146 | 26.273 | 20.051 | 0.377 | 37.7
[6 empty] | -0.146 | 26.263 | 20.051 | 0.375 | 37.5
[6 empty] | -0.146 | 26.253 | 20.051 | 0.372 | 37.2
[6 empty] | -0.146 | 26.243 | 20.051 | 0.37 | 37
[6 empty] | -0.146 | 26.233 | 20.051 | 0.368 | 36.8
[6 empty] | -0.146 | 26.223 | 20.051 | 0.365 | 36.5
[6 empty] | -0.146 | 26.213 | 20.051 | 0.363 | 36.3
[6 empty] | -0.146 | 26.203 | 20.05 | 0.359 | 35.9
[6 empty] | -0.146 | 26.193 | 20.05 | 0.35 | 35
[6 empty] | -0.146 | 26.183 | 20.05 | 0.341 | 34.1
[6 empty] | -0.146 | 26.173 | 20.05 | 0.331 | 33.1
[6 empty] | -0.146 | 26.163 | 20.05 | 0.322 | 32.2
[6 empty] | -0.146 | 26.153 | 20.05 | 0.313 | 31.3
[6 empty] | -0.146 | 26.143 | 20.05 | 0.304 | 30.4
[6 empty] | -0.146 | 26.133 | 20.05 | 0.295 | 29.5
[6 empty] | -0.146 | 26.123 | 20.05 | 0.285 | 28.5
[6 empty] | -0.146 | 26.113 | 20.05 | 0.276 | 27.6
[6 empty] | -0.146 | 26.103 | 20.05 | 0.267 | 26.7
[6 empty] | -0.146 | 26.093 | 20.05 | 0.258 | 25.8
[6 empty] | -0.146 | 26.083 | 20.05 | 0.248 | 24.8
[6 empty] | -0.146 | 26.073 | 20.05 | 0.239 | 23.9
[6 empty] | -0.146 | 26.063 | 20.05 | 0.23 | 23
[6 empty] | -0.146 | 26.053 | 20.05 | 0.221 | 22.1
[6 empty] | -0.146 | 26.043 | 20.05 | 0.211 | 21.1
[6 empty] | -0.146 | 26.033 | 20.05 | 0.202 | 20.2
[6 empty] | -0.146 | 26.023 | 20.05 | 0.193 | 19.3
[6 empty] | -0.146 | 26.013 | 20.05 | 0.184 | 18.4
[6 empty] | -0.146 | 26.003 | 20.05 | 0.176 | 17.6
[6 empty] | -0.146 | 25.993 | 20.05 | 0.174 | 17.4
[6 empty] | -0.146 | 25.983 | 20.05 | 0.172 | 17.2
[6 empty] | -0.146 | 25.973 | 20.05 | 0.17 | 17
[6 empty] | -0.146 | 25.963 | 20.05 | 0.169 | 16.9
[6 empty] | -0.146 | 25.953 | 20.05 | 0.167 | 16.7
[6 empty] | -0.146 | 25.943 | 20.05 | 0.165 | 16.5
[6 empty] | -0.146 | 25.933 | 20.05 | 0.163 | 16.3
[6 empty] | -0.146 | 25.923 | 20.05 | 0.162 | 16.2
[6 empty] | -0.146 | 25.913 | 20.05 | 0.16 | 16
[6 empty] | -0.146 | 25.903 | 20.05 | 0.158 | 15.8
[6 empty] | -0.146 | 25.893 | 20.05 | 0.157 | 15.7
[6 empty] | -0.146 | 25.883 | 20.05 | 0.155 | 15.5
[6 empty] | -0.146 | 25.873 | 20.05 | 0.153 | 15.3
[6 empty] | -0.146 | 25.863 | 20.05 | 0.151 | 15.1
[6 empty] | -0.146 | 25.853 | 20.05 | 0.15 | 15
[6 empty] | -0.146 | 25.843 | 20.05 | 0.148 | 14.8
[6 empty] | -0.146 | 25.833 | 20.05 | 0.146 | 14.6
[6 empty] | -0.146 | 25.823 | 20.05 | 0.144 | 14.4
[6 empty] | -0.146 | 25.813 | 20.05 | 0.143 | 14.3
[6 empty] | -0.146 | 25.803 | 20.05 | 0.142 | 14.2
[6 empty] | -0.146 | 25.793 | 20.05 | 0.145 | 14.5
[6 empty] | -0.146 | 25.783 | 20.05 | 0.149 | 14.9
[6 empty] | -0.146 | 25.773 | 20.05 | 0.153 | 15.3
[6 empty] | -0.146 | 25.763 | 20.05 | 0.157 | 15.7
[6 empty] | -0.146 | 25.753 | 20.05 | 0.16 | 16
[6 empty] | -0.146 | 25.743 | 20.05 | 0.164 | 16.4
[6 empty] | -0.146 | 25.733 | 20.05 | 0.168 | 16.8
[6 empty] | -0.146 | 25.723 | 20.05 | 0.171 | 17.1
[6 empty] | -0.146 | 25.713 | 20.05 | 0.175 | 17.5
[6 empty] | -0.146 | 25.703 | 20.05 | 0.179 | 17.9
[6 empty] | -0.146 | 25.693 | 20.05 | 0.183 | 18.3
[6 empty] | -0.146 | 25.683 | 20.05 | 0.186 | 18.6
[6 empty] | -0.146 | 25.673 | 20.05 | 0.19 | 19
[6 empty] | -0.146 | 25.663 | 20.05 | 0.194 | 19.4
[6 empty] | -0.146 | 25.653 | 20.05 | 0.197 | 19.7
[6 empty] | -0.146 | 25.643 | 20.05 | 0.201 | 20.1
[6 empty] | -0.146 | 25.633 | 20.05 | 0.205 | 20.5
[6 empty] | -0.146 | 25.623 | 20.05 | 0.208 | 20.8
[6 empty] | -0.146 | 25.613 | 20.05 | 0.212 | 21.2
[6 empty] | -0.146 | 25.603 | 20.05 | 0.217 | 21.7
[6 empty] | -0.146 | 25.593 | 20.05 | 0.225 | 22.5
[6 empty] | -0.146 | 25.583 | 20.05 | 0.233 | 23.3
[6 empty] | -0.146 | 25.573 | 20.05 | 0.242 | 24.2
[6 empty] | -0.146 | 25.563 | 20.05 | 0.25 | 25
[6 empty] | -0.146 | 25.553 | 20.05 | 0.258 | 25.8
[6 empty] | -0.146 | 25.543 | 20.05 | 0.267 | 26.7
[6 empty] | -0.146 | 25.533 | 20.05 | 0.275 | 27.5
[6 empty] | -0.146 | 25.523 | 20.05 | 0.283 | 28.3
[6 empty] | -0.146 | 25.513 | 20.049 | 0.291 | 29.1
[6 empty] | -0.146 | 25.503 | 20.049 | 0.3 | 30
[6 empty] | -0.146 | 25.493 | 20.049 | 0.308 | 30.8
[6 empty] | -0.146 | 25.483 | 20.049 | 0.316 | 31.6
[6 empty] | -0.146 | 25.473 | 20.049 | 0.325 | 32.5
[6 empty] | -0.146 | 25.463 | 20.049 | 0.333 | 33.3
[6 empty] | -0.146 | 25.453 | 20.049 | 0.341 | 34.1
[6 empty] | -0.146 | 25.443 | 20.049 | 0.35 | 35
[6 empty] | -0.146 | 25.433 | 20.049 | 0.358 | 35.8
[6 empty] | -0.146 | 25.423 | 20.049 | 0.366 | 36.6
[6 empty] | -0.146 | 25.413 | 20.049 | 0.375 | 37.5
[6 empty] | -0.146 | 25.403 | 20.049 | 0.382 | 38.2
[6 empty] | -0.146 | 25.393 | 20.049 | 0.383 | 38.3
[6 empty] | -0.146 | 25.383 | 20.049 | 0.384 | 38.4
[6 empty] | -0.146 | 25.373 | 20.049 | 0.384 | 38.4
[6 empty] | -0.146 | 25.363 | 20.049 | 0.385 | 38.5
[6 empty] | -0.146 | 25.353 | 20.049 | 0.386 | 38.6
[6 empty] | -0.146 | 25.343 | 20.049 | 0.387 | 38.7
[6 empty] | -0.146 | 25.333 | 20.049 | 0.388 | 38.8
[6 empty] | -0.146 | 25.323 | 20.049 | 0.389 | 38.9
[6 empty] | -0.146 | 25.313 | 20.049 | 0.39 | 39
[6 empty] | -0.146 | 25.303 | 20.049 | 0.391 | 39.1
[6 empty] | -0.146 | 25.293 | 20.049 | 0.391 | 39.1
[6 empty] | -0.146 | 25.283 | 20.049 | 0.392 | 39.2
[6 empty] | -0.146 | 25.273 | 20.049 | 0.393 | 39.3
[6 empty] | -0.146 | 25.263 | 20.049 | 0.394 | 39.4
[6 empty] | -0.146 | 25.253 | 20.049 | 0.395 | 39.5
[6 empty] | -0.146 | 25.243 | 20.049 | 0.396 | 39.6
[6 empty] | -0.146 | 25.233 | 20.049 | 0.397 | 39.7
[6 empty] | -0.146 | 25.223 | 20.049 | 0.398 | 39.8
[6 empty] | -0.146 | 25.213 | 20.049 | 0.398 | 39.8
[6 empty] | -0.146 | 25.203 | 20.049 | 0.398 | 39.8
[6 empty] | -0.146 | 25.193 | 20.049 | 0.391 | 39.1
[6 empty] | -0.146 | 25.183 | 20.049 | 0.384 | 38.4
[6 empty] | -0.146 | 25.173 | 20.049 | 0.377 | 37.7
[6 empty] | -0.146 | 25.163 | 20.049 | 0.37 | 37
[6 empty] | -0.146 | 25.153 | 20.049 | 0.363 | 36.3
[6 empty] | -0.146 | 25.143 | 20.049 | 0.356 | 35.6
[6 empty] | -0.146 | 25.133 | 20.049 | 0.349 | 34.9
[6 empty] | -0.146 | 25.123 | 20.049 | 0.342 | 34.2
[6 empty] | -0.146 | 25.113 | 20.049 | 0.335 | 33.5
[6 empty] | -0.146 | 25.103 | 20.049 | 0.328 | 32.8
[6 empty] | -0.146 | 25.093 | 20.049 | 0.321 | 32.1
[6 empty] | -0.146 | 25.083 | 20.049 | 0.314 | 31.4
[6 empty] | -0.146 | 25.073 | 20.049 | 0.307 | 30.7
[6 empty] | -0.146 | 25.063 | 20.049 | 0.3 | 30
[6 empty] | -0.146 | 25.053 | 20.049 | 0.293 | 29.3
[6 empty] | -0.146 | 25.043 | 20.049 | 0.286 | 28.6
[6 empty] | -0.146 | 25.033 | 20.049 | 0.279 | 27.9
[6 empty] | -0.146 | 25.023 | 20.049 | 0.272 | 27.2
[6 empty] | -0.146 | 25.013 | 20.049 | 0.266 | 26.6
[6 empty] | -0.146 | 25.003 | 20.049 | 0.259 | 25.9
[6 empty] | -0.146 | 24.993 | 20.049 | 0.253 | 25.3
[6 empty] | -0.146 | 24.983 | 20.049 | 0.247 | 24.7
[6 empty] | -0.146 | 24.974 | 20.049 | 0.242 | 24.2
[6 empty] | -0.146 | 24.964 | 20.049 | 0.236 | 23.6
[6 empty] | -0.146 | 24.954 | 20.049 | 0.231 | 23.1
[6 empty] | -0.146 | 24.944 | 20.049 | 0.225 | 22.5
[6 empty] | -0.146 | 24.934 | 20.049 | 0.219 | 21.9
[6 empty] | -0.146 | 24.924 | 20.049 | 0.214 | 21.4
[6 empty] | -0.146 | 24.914 | 20.049 | 0.208 | 20.8
[6 empty] | -0.146 | 24.904 | 20.049 | 0.202 | 20.2
[6 empty] | -0.146 | 24.894 | 20.049 | 0.197 | 19.7
[6 empty] | -0.146 | 24.884 | 20.049 | 0.191 | 19.1
[6 empty] | -0.146 | 24.874 | 20.049 | 0.185 | 18.5
[6 empty] | -0.146 | 24.864 | 20.049 | 0.18 | 18
[6 empty] | -0.146 | 24.854 | 20.049 | 0.174 | 17.4
[6 empty] | -0.146 | 24.844 | 20.049 | 0.169 | 16.9
[6 empty] | -0.146 | 24.834 | 20.049 | 0.163 | 16.3
[6 empty] | -0.146 | 24.824 | 20.048 | 0.157 | 15.7
[6 empty] | -0.146 | 24.814 | 20.048 | 0.152 | 15.2
[6 empty] | -0.146 | 24.804 | 20.048 | 0.147 | 14.7
[6 empty] | -0.146 | 24.794 | 20.048 | 0.148 | 14.8
[6 empty] | -0.146 | 24.784 | 20.048 | 0.148 | 14.8
[6 empty] | -0.146 | 24.774 | 20.048 | 0.149 | 14.9
[6 empty] | -0.146 | 24.764 | 20.048 | 0.149 | 14.9
[6 empty] | -0.146 | 24.754 | 20.048 | 0.15 | 15
[6 empty] | -0.146 | 24.744 | 20.048 | 0.151 | 15.1
[6 empty] | -0.146 | 24.734 | 20.048 | 0.151 | 15.1
[6 empty] | -0.146 | 24.724 | 20.048 | 0.152 | 15.2
[6 empty] | -0.146 | 24.714 | 20.048 | 0.152 | 15.2
[6 empty] | -0.146 | 24.704 | 20.048 | 0.153 | 15.3
[6 empty] | -0.146 | 24.694 | 20.048 | 0.154 | 15.4
[6 empty] | -0.146 | 24.684 | 20.048 | 0.154 | 15.4
[6 empty] | -0.146 | 24.674 | 20.048 | 0.155 | 15.5
[6 empty] | -0.146 | 24.664 | 20.048 | 0.155 | 15.5
[6 empty] | -0.146 | 24.654 | 20.048 | 0.156 | 15.6
[6 empty] | -0.146 | 24.644 | 20.048 | 0.157 | 15.7
[6 empty] | -0.146 | 24.634 | 20.048 | 0.157 | 15.7
[6 empty] | -0.146 | 24.624 | 20.048 | 0.158 | 15.8
[6 empty] | -0.146 | 24.614 | 20.048 | 0.158 | 15.8
[6 empty] | -0.146 | 24.604 | 20.048 | 0.16 | 16
[6 empty] | -0.146 | 24.594 | 20.048 | 0.168 | 16.8
[6 empty] | -0.146 | 24.584 | 20.048 | 0.175 | 17.5
[6 empty] | -0.146 | 24.574 | 20.048 | 0.183 | 18.3
[6 empty] | -0.146 | 24.564 | 20.048 | 0.19 | 19
[6 empty] | -0.146 | 24.554 | 20.048 | 0.198 | 19.8
[6 empty] | -0.146 | 24.544 | 20.048 | 0.206 | 20.6
[6 empty] | -0.146 | 24.534 | 20.048 | 0.213 | 21.3
[6 empty] | -0.146 | 24.524 | 20.048 | 0.221 | 22.1
[6 empty] | -0.146 | 24.514 | 20.048 | 0.228 | 22.8
[6 empty] | -0.146 | 24.504 | 20.048 | 0.236 | 23.6
[6 empty] | -0.146 | 24.494 | 20.048 | 0.244 | 24.4
[6 empty] | -0.146 | 24.484 | 20.048 | 0.251 | 25.1
[6 empty] | -0.146 | 24.474 | 20.048 | 0.259 | 25.9
[6 empty] | -0.146 | 24.464 | 20.048 | 0.266 | 26.6
[6 empty] | -0.146 | 24.454 | 20.048 | 0.274 | 27.4
[6 empty] | -0.146 | 24.444 | 20.048 | 0.282 | 28.2
[6 empty] | -0.146 | 24.434 | 20.048 | 0.289 | 28.9
[6 empty] | -0.146 | 24.424 | 20.048 | 0.297 | 29.7
[6 empty] | -0.146 | 24.414 | 20.048 | 0.304 | 30.4
[6 empty] | -0.146 | 24.404 | 20.048 | 0.312 | 31.2
[6 empty] | -0.146 | 24.394 | 20.048 | 0.317 | 31.7
[6 empty] | -0.146 | 24.384 | 20.048 | 0.321 | 32.1
[6 empty] | -0.146 | 24.374 | 20.048 | 0.326 | 32.6
[6 empty] | -0.146 | 24.364 | 20.048 | 0.331 | 33.1
[6 empty] | -0.146 | 24.354 | 20.048 | 0.336 | 33.6
[6 empty] | -0.146 | 24.344 | 20.048 | 0.341 | 34.1
[6 empty] | -0.146 | 24.334 | 20.048 | 0.346 | 34.6
[6 empty] | -0.146 | 24.324 | 20.048 | 0.351 | 35.1
[6 empty] | -0.146 | 24.314 | 20.048 | 0.355 | 35.5
[6 empty] | -0.146 | 24.304 | 20.048 | 0.36 | 36
[6 empty] | -0.146 | 24.294 | 20.048 | 0.365 | 36.5
[6 empty] | -0.146 | 24.284 | 20.048 | 0.37 | 37
[6 empty] | -0.146 | 24.274 | 20.048 | 0.375 | 37.5
[6 empty] | -0.146 | 24.264 | 20.048 | 0.38 | 38
[6 empty] | -0.146 | 24.254 | 20.048 | 0.385 | 38.5
[6 empty] | -0.146 | 24.244 | 20.048 | 0.39 | 39
[6 empty] | -0.146 | 24.234 | 20.048 | 0.394 | 39.4
[6 empty] | -0.146 | 24.224 | 20.048 | 0.399 | 39.9
[6 empty] | -0.146 | 24.214 | 20.048 | 0.404 | 40.4
[6 empty] | -0.146 | 24.204 | 20.048 | 0.408 | 40.8
[6 empty] | -0.146 | 24.194 | 20.048 | 0.405 | 40.5
[6 empty] | -0.146 | 24.184 | 20.048 | 0.403 | 40.3
[6 empty] | -0.146 | 24.174 | 20.048 | 0.4 | 40
[6 empty] | -0.146 | 24.164 | 20.048 | 0.398 | 39.8
[6 empty] | -0.146 | 24.154 | 20.048 | 0.395 | 39.5
[6 empty] | -0.146 | 24.144 | 20.048 | 0.393 | 39.3
[6 empty] | -0.146 | 24.134 | 20.047 | 0.39 | 39
[6 empty] | -0.146 | 24.124 | 20.047 | 0.388 | 38.8
[6 empty] | -0.146 | 24.114 | 20.047 | 0.385 | 38.5
[6 empty] | -0.146 | 24.104 | 20.047 | 0.382 | 38.2
[6 empty] | -0.146 | 24.094 | 20.047 | 0.38 | 38
[6 empty] | -0.146 | 24.084 | 20.047 | 0.377 | 37.7
[6 empty] | -0.146 | 24.074 | 20.047 | 0.375 | 37.5
[6 empty] | -0.146 | 24.064 | 20.047 | 0.372 | 37.2
[6 empty] | -0.146 | 24.054 | 20.047 | 0.37 | 37
[6 empty] | -0.146 | 24.044 | 20.047 | 0.367 | 36.7
[6 empty] | -0.146 | 24.034 | 20.047 | 0.365 | 36.5
[6 empty] | -0.146 | 24.024 | 20.047 | 0.362 | 36.2
[6 empty] | -0.146 | 24.014 | 20.047 | 0.36 | 36
[6 empty] | -0.146 | 24.004 | 20.047 | 0.356 | 35.6
[6 empty] | -0.146 | 23.994 | 20.047 | 0.347 | 34.7
[6 empty] | -0.146 | 23.984 | 20.047 | 0.338 | 33.8
[6 empty] | -0.146 | 23.974 | 20.047 | 0.328 | 32.8
[6 empty] | -0.146 | 23.964 | 20.047 | 0.319 | 31.9
[6 empty] | -0.146 | 23.954 | 20.047 | 0.31 | 31
[6 empty] | -0.146 | 23.944 | 20.047 | 0.301 | 30.1
[6 empty] | -0.146 | 23.934 | 20.047 | 0.292 | 29.2
[6 empty] | -0.146 | 23.924 | 20.047 | 0.282 | 28.2
[6 empty] | -0.146 | 23.914 | 20.047 | 0.273 | 27.3
[6 empty] | -0.146 | 23.904 | 20.047 | 0.264 | 26.4
[6 empty] | -0.146 | 23.894 | 20.047 | 0.255 | 25.5
[6 empty] | -0.146 | 23.884 | 20.047 | 0.245 | 24.5
[6 empty] | -0.146 | 23.874 | 20.047 | 0.236 | 23.6
[6 empty] | -0.146 | 23.864 | 20.047 | 0.227 | 22.7
[6 empty] | -0.146 | 23.854 | 20.047 | 0.218 | 21.8
[6 empty] | -0.146 | 23.844 | 20.047 | 0.208 | 20.8
[6 empty] | -0.146 | 23.834 | 20.047 | 0.199 | 19.9
[6 empty] | -0.146 | 23.824 | 20.047 | 0.19 | 19
[6 empty] | -0.146 | 23.814 | 20.047 | 0.181 | 18.1
[6 empty] | -0.146 | 23.804 | 20.047 | 0.173 | 17.3
[6 empty] | -0.146 | 23.794 | 20.047 | 0.171 | 17.1
[6 empty] | -0.146 | 23.784 | 20.047 | 0.169 | 16.9
[6 empty] | -0.146 | 23.774 | 20.047 | 0.168 | 16.8
[6 empty] | -0.146 | 23.764 | 20.047 | 0.166 | 16.6
[6 empty] | -0.146 | 23.754 | 20.047 | 0.165 | 16.5
[6 empty] | -0.146 | 23.744 | 20.047 | 0.163 | 16.3
[6 empty] | -0.146 | 23.734 | 20.047 | 0.161 | 16.1
[6 empty] | -0.146 | 23.724 | 20.047 | 0.16 | 16
[6 empty] | -0.146 | 23.714 | 20.047 | 0.158 | 15.8
[6 empty] | -0.146 | 23.704 | 20.047 | 0.157 | 15.7
[6 empty] | -0.146 | 23.694 | 20.047 | 0.155 | 15.5
[6 empty] | -0.146 | 23.684 | 20.047 | 0.153 | 15.3
[6 empty] | -0.146 | 23.674 | 20.047 | 0.152 | 15.2
[6 empty] | -0.146 | 23.664 | 20.047 | 0.15 | 15
[6 empty] | -0.146 | 23.654 | 20.047 | 0.149 | 14.9
[6 empty] | -0.146 | 23.644 | 20.047 | 0.147 | 14.7
[6 empty] | -0.146 | 23.634 | 20.047 | 0.145 | 14.5
[6 empty] | -0.146 | 23.624 | 20.047 | 0.144 | 14.4
[6 empty] | -0.146 | 23.614 | 20.047 | 0.142 | 14.2
[6 empty] | -0.146 | 23.604 | 20.047 | 0.141 | 14.1
[6 empty] | -0.146 | 23.594 | 20.047 | 0.145 | 14.5
[6 empty] | -0.146 | 23.584 | 20.047 | 0.149 | 14.9
[6 empty] | -0.146 | 23.574 | 20.047 | 0.153 | 15.3
[6 empty] | -0.146 | 23.564 | 20.047 | 0.156 | 15.6
[6 empty] | -0.146 | 23.554 | 20.047 | 0.16 | 16
[6 empty] | -0.146 | 23.544 | 20.047 | 0.164 | 16.4
[6 empty] | -0.146 | 23.534 | 20.047 | 0.168 | 16.8
[6 empty] | -0.146 | 23.524 | 20.047 | 0.172 | 17.2
[6 empty] | -0.146 | 23.514 | 20.047 | 0.176 | 17.6
[6 empty] | -0.146 | 23.504 | 20.047 | 0.179 | 17.9
[6 empty] | -0.146 | 23.494 | 20.047 | 0.183 | 18.3
[6 empty] | -0.146 | 23.484 | 20.047 | 0.187 | 18.7
[6 empty] | -0.146 | 23.474 | 20.047 | 0.191 | 19.1
[6 empty] | -0.146 | 23.464 | 20.047 | 0.195 | 19.5
[6 empty] | -0.146 | 23.454 | 20.047 | 0.198 | 19.8
[6 empty] | -0.146 | 23.444 | 20.046 | 0.202 | 20.2
[6 empty] | -0.146 | 23.434 | 20.046 | 0.206 | 20.6
[6 empty] | -0.146 | 23.424 | 20.046 | 0.21 | 21
[6 empty] | -0.146 | 23.414 | 20.046 | 0.214 | 21.4
[6 empty] | -0.146 | 23.404 | 20.046 | 0.218 | 21.8
[6 empty] | -0.146 | 23.394 | 20.046 | 0.226 | 22.6
[6 empty] | -0.146 | 23.384 | 20.046 | 0.235 | 23.5
[6 empty] | -0.146 | 23.374 | 20.046 | 0.243 | 24.3
[6 empty] | -0.146 | 23.364 | 20.046 | 0.251 | 25.1
[6 empty] | -0.146 | 23.354 | 20.046 | 0.26 | 26
[6 empty] | -0.146 | 23.344 | 20.046 | 0.268 | 26.8
[6 empty] | -0.146 | 23.334 | 20.046 | 0.277 | 27.7
[6 empty] | -0.146 | 23.324 | 20.046 | 0.285 | 28.5
[6 empty] | -0.146 | 23.314 | 20.046 | 0.293 | 29.3
[6 empty] | -0.146 | 23.304 | 20.046 | 0.302 | 30.2
[6 empty] | -0.146 | 23.294 | 20.046 | 0.31 | 31
[6 empty] | -0.146 | 23.284 | 20.046 | 0.318 | 31.8
[6 empty] | -0.146 | 23.274 | 20.046 | 0.327 | 32.7
[6 empty] | -0.146 | 23.264 | 20.046 | 0.335 | 33.5
[6 empty] | -0.146 | 23.254 | 20.046 | 0.343 | 34.3
[6 empty] | -0.146 | 23.244 | 20.046 | 0.352 | 35.2
[6 empty] | -0.146 | 23.234 | 20.046 | 0.36 | 36
[6 empty] | -0.146 | 23.224 | 20.046 | 0.369 | 36.9
[6 empty] | -0.146 | 23.214 | 20.046 | 0.377 | 37.7
[6 empty] | -0.146 | 23.204 | 20.046 | 0.384 | 38.4
[6 empty] | -0.146 | 23.194 | 20.046 | 0.385 | 38.5
[6 empty] | -0.146 | 23.184 | 20.046 | 0.386 | 38.6
[6 empty] | -0.146 | 23.174 | 20.046 | 0.387 | 38.7
[6 empty] | -0.146 | 23.164 | 20.046 | 0.388 | 38.8
[6 empty] | -0.146 | 23.154 | 20.046 | 0.388 | 38.8
[6 empty] | -0.146 | 23.144 | 20.046 | 0.389 | 38.9
[6 empty] | -0.146 | 23.134 | 20.046 | 0.39 | 39
[6 empty] | -0.146 | 23.124 | 20.046 | 0.391 | 39.1
[6 empty] | -0.146 | 23.114 | 20.046 | 0.392 | 39.2
[6 empty] | -0.146 | 23.104 | 20.046 | 0.392 | 39.2
[6 empty] | -0.146 | 23.094 | 20.046 | 0.393 | 39.3
[6 empty] | -0.146 | 23.084 | 20.046 | 0.394 | 39.4
[6 empty] | -0.146 | 23.074 | 20.046 | 0.395 | 39.5
[6 empty] | -0.146 | 23.064 | 20.046 | 0.396 | 39.6
[6 empty] | -0.146 | 23.054 | 20.046 | 0.396 | 39.6
[6 empty] | -0.146 | 23.044 | 20.046 | 0.397 | 39.7
[6 empty] | -0.146 | 23.034 | 20.046 | 0.398 | 39.8
[6 empty] | -0.146 | 23.024 | 20.046 | 0.399 | 39.9
[6 empty] | -0.146 | 23.014 | 20.046 | 0.4 | 40
[6 empty] | -0.146 | 23.004 | 20.046 | 0.399 | 39.9
[6 empty] | -0.146 | 22.994 | 20.046 | 0.392 | 39.2
[6 empty] | -0.146 | 22.984 | 20.046 | 0.385 | 38.5
[6 empty] | -0.146 | 22.974 | 20.046 | 0.378 | 37.8
[6 empty] | -0.146 | 22.964 | 20.046 | 0.371 | 37.1
[6 empty] | -0.146 | 22.954 | 20.046 | 0.364 | 36.4
[6 empty] | -0.146 | 22.944 | 20.046 | 0.357 | 35.7
[6 empty] | -0.146 | 22.934 | 20.046 | 0.349 | 34.9
[6 empty] | -0.146 | 22.924 | 20.046 | 0.342 | 34.2
[6 empty] | -0.146 | 22.914 | 20.046 | 0.335 | 33.5
[6 empty] | -0.146 | 22.904 | 20.046 | 0.328 | 32.8
[6 empty] | -0.146 | 22.894 | 20.046 | 0.321 | 32.1
[6 empty] | -0.146 | 22.884 | 20.046 | 0.314 | 31.4
[6 empty] | -0.146 | 22.874 | 20.046 | 0.307 | 30.7
[6 empty] | -0.146 | 22.864 | 20.046 | 0.299 | 29.9
[6 empty] | -0.146 | 22.854 | 20.046 | 0.292 | 29.2
[6 empty] | -0.146 | 22.844 | 20.046 | 0.285 | 28.5
[6 empty] | -0.146 | 22.834 | 20.046 | 0.278 | 27.8
[6 empty] | -0.146 | 22.824 | 20.046 | 0.271 | 27.1
[6 empty] | -0.146 | 22.814 | 20.046 | 0.264 | 26.4
[6 empty] | -0.146 | 22.804 | 20.046 | 0.257 | 25.7
[6 empty] | -0.146 | 22.794 | 20.046 | 0.251 | 25.1
[6 empty] | -0.146 | 22.784 | 20.046 | 0.246 | 24.6
[6 empty] | -0.146 | 22.774 | 20.046 | 0.24 | 24
[6 empty] | -0.146 | 22.764 | 20.046 | 0.234 | 23.4
[6 empty] | -0.146 | 22.754 | 20.045 | 0.229 | 22.9
[6 empty] | -0.146 | 22.744 | 20.045 | 0.223 | 22.3
[6 empty] | -0.146 | 22.734 | 20.045 | 0.218 | 21.8
[6 empty] | -0.146 | 22.724 | 20.045 | 0.212 | 21.2
[6 empty] | -0.146 | 22.714 | 20.045 | 0.206 | 20.6
[6 empty] | -0.146 | 22.704 | 20.045 | 0.201 | 20.1
[6 empty] | -0.146 | 22.694 | 20.045 | 0.195 | 19.5
[6 empty] | -0.146 | 22.684 | 20.045 | 0.189 | 18.9
[6 empty] | -0.146 | 22.674 | 20.045 | 0.184 | 18.4
[6 empty] | -0.146 | 22.664 | 20.045 | 0.178 | 17.8
[6 empty] | -0.146 | 22.654 | 20.045 | 0.173 | 17.3
[6 empty] | -0.146 | 22.644 | 20.045 | 0.167 | 16.7
[6 empty] | -0.146 | 22.634 | 20.045 | 0.161 | 16.1
[6 empty] | -0.146 | 22.624 | 20.045 | 0.156 | 15.6
[6 empty] | -0.146 | 22.614 | 20.045 | 0.15 | 15
[6 empty] | -0.146 | 22.604 | 20.045 | 0.145 | 14.5
[6 empty] | -0.146 | 22.594 | 20.045 | 0.146 | 14.6
[6 empty] | -0.146 | 22.584 | 20.045 | 0.147 | 14.7
[6 empty] | -0.146 | 22.574 | 20.045 | 0.147 | 14.7
[6 empty] | -0.146 | 22.564 | 20.045 | 0.148 | 14.8
[6 empty] | -0.146 | 22.554 | 20.045 | 0.149 | 14.9
[6 empty] | -0.146 | 22.544 | 20.045 | 0.149 | 14.9
[6 empty] | -0.146 | 22.534 | 20.045 | 0.15 | 15
[6 empty] | -0.146 | 22.524 | 20.045 | 0.151 | 15.1
[6 empty] | -0.146 | 22.514 | 20.045 | 0.151 | 15.1
[6 empty] | -0.146 | 22.504 | 20.045 | 0.152 | 15.2
[6 empty] | -0.146 | 22.494 | 20.045 | 0.153 | 15.3
[6 empty] | -0.146 | 22.484 | 20.045 | 0.153 | 15.3
[6 empty] | -0.146 | 22.474 | 20.045 | 0.154 | 15.4
[6 empty] | -0.146 | 22.464 | 20.045 | 0.155 | 15.5
[6 empty] | -0.146 | 22.454 | 20.045 | 0.155 | 15.5
[6 empty] | -0.146 | 22.444 | 20.045 | 0.156 | 15.6
[6 empty] | -0.146 | 22.434 | 20.045 | 0.157 | 15.7
[6 empty] | -0.146 | 22.424 | 20.045 | 0.157 | 15.7
[6 empty] | -0.146 | 22.414 | 20.045 | 0.158 | 15.8
[6 empty] | -0.146 | 22.404 | 20.045 | 0.16 | 16
[6 empty] | -0.146 | 22.394 | 20.045 | 0.167 | 16.7
[6 empty] | -0.146 | 22.384 | 20.045 | 0.175 | 17.5
[6 empty] | -0.146 | 22.374 | 20.045 | 0.183 | 18.3
[6 empty] | -0.146 | 22.364 | 20.045 | 0.191 | 19.1
[6 empty] | -0.146 | 22.354 | 20.045 | 0.199 | 19.9
[6 empty] | -0.146 | 22.344 | 20.045 | 0.206 | 20.6
[6 empty] | -0.146 | 22.334 | 20.045 | 0.214 | 21.4
[6 empty] | -0.146 | 22.324 | 20.045 | 0.222 | 22.2
[6 empty] | -0.146 | 22.314 | 20.045 | 0.23 | 23
[6 empty] | -0.146 | 22.304 | 20.045 | 0.237 | 23.7
[6 empty] | -0.146 | 22.294 | 20.045 | 0.245 | 24.5
[6 empty] | -0.146 | 22.284 | 20.045 | 0.253 | 25.3
[6 empty] | -0.146 | 22.274 | 20.045 | 0.261 | 26.1
[6 empty] | -0.146 | 22.264 | 20.045 | 0.269 | 26.9
[6 empty] | -0.146 | 22.254 | 20.045 | 0.276 | 27.6
[6 empty] | -0.146 | 22.244 | 20.045 | 0.284 | 28.4
[6 empty] | -0.146 | 22.234 | 20.045 | 0.292 | 29.2
[6 empty] | -0.146 | 22.224 | 20.045 | 0.3 | 30
[6 empty] | -0.146 | 22.214 | 20.045 | 0.308 | 30.8
[6 empty] | -0.146 | 22.204 | 20.045 | 0.315 | 31.5
[6 empty] | -0.146 | 22.194 | 20.045 | 0.32 | 32
[6 empty] | -0.146 | 22.184 | 20.045 | 0.325 | 32.5
[6 empty] | -0.146 | 22.174 | 20.045 | 0.329 | 32.9
[6 empty] | -0.146 | 22.164 | 20.045 | 0.334 | 33.4
[6 empty] | -0.146 | 22.154 | 20.045 | 0.339 | 33.9
[6 empty] | -0.146 | 22.144 | 20.045 | 0.344 | 34.4
[6 empty] | -0.146 | 22.134 | 20.045 | 0.349 | 34.9
[6 empty] | -0.146 | 22.124 | 20.045 | 0.353 | 35.3
[6 empty] | -0.146 | 22.114 | 20.045 | 0.358 | 35.8
[6 empty] | -0.146 | 22.104 | 20.045 | 0.363 | 36.3
[6 empty] | -0.146 | 22.094 | 20.045 | 0.368 | 36.8
[6 empty] | -0.146 | 22.084 | 20.045 | 0.373 | 37.3
[6 empty] | -0.146 | 22.074 | 20.045 | 0.377 | 37.7
[6 empty] | -0.146 | 22.064 | 20.044 | 0.382 | 38.2
[6 empty] | -0.146 | 22.054 | 20.044 | 0.387 | 38.7
[6 empty] | -0.146 | 22.044 | 20.044 | 0.392 | 39.2
[6 empty] | -0.146 | 22.034 | 20.044 | 0.397 | 39.7
[6 empty] | -0.146 | 22.024 | 20.044 | 0.401 | 40.1
[6 empty] | -0.146 | 22.014 | 20.044 | 0.406 | 40.6
[6 empty] | -0.146 | 22.004 | 20.044 | 0.41 | 41
[6 empty] | -0.146 | 21.994 | 20.044 | 0.408 | 40.8
[6 empty] | -0.146 | 21.984 | 20.044 | 0.405 | 40.5
[6 empty] | -0.146 | 21.975 | 20.044 | 0.402 | 40.2
[6 empty] | -0.146 | 21.965 | 20.044 | 0.4 | 40
[6 empty] | -0.146 | 21.955 | 20.044 | 0.397 | 39.7
[6 empty] | -0.146 | 21.945 | 20.044 | 0.394 | 39.4
[6 empty] | -0.146 | 21.935 | 20.044 | 0.392 | 39.2
[6 empty] | -0.146 | 21.925 | 20.044 | 0.389 | 38.9
[6 empty] | -0.146 | 21.915 | 20.044 | 0.386 | 38.6
[6 empty] | -0.146 | 21.905 | 20.044 | 0.383 | 38.3
[6 empty] | -0.146 | 21.895 | 20.044 | 0.381 | 38.1
[6 empty] | -0.146 | 21.885 | 20.044 | 0.378 | 37.8
[6 empty] | -0.146 | 21.875 | 20.044 | 0.375 | 37.5
[6 empty] | -0.146 | 21.865 | 20.044 | 0.373 | 37.3
[6 empty] | -0.146 | 21.855 | 20.044 | 0.37 | 37
[6 empty] | -0.146 | 21.845 | 20.044 | 0.367 | 36.7
[6 empty] | -0.146 | 21.835 | 20.044 | 0.365 | 36.5
[6 empty] | -0.146 | 21.825 | 20.044 | 0.362 | 36.2
[6 empty] | -0.146 | 21.815 | 20.044 | 0.359 | 35.9
[6 empty] | -0.146 | 21.805 | 20.044 | 0.356 | 35.6
[6 empty] | -0.146 | 21.795 | 20.044 | 0.347 | 34.7
[6 empty] | -0.146 | 21.785 | 20.044 | 0.337 | 33.7
[6 empty] | -0.146 | 21.775 | 20.044 | 0.328 | 32.8
[6 empty] | -0.146 | 21.765 | 20.044 | 0.319 | 31.9
[6 empty] | -0.146 | 21.755 | 20.044 | 0.309 | 30.9
[6 empty] | -0.146 | 21.745 | 20.044 | 0.3 | 30
[6 empty] | -0.146 | 21.735 | 20.044 | 0.291 | 29.1
[6 empty] | -0.146 | 21.725 | 20.044 | 0.281 | 28.1
[6 empty] | -0.146 | 21.715 | 20.044 | 0.272 | 27.2
[6 empty] | -0.146 | 21.705 | 20.044 | 0.263 | 26.3
[6 empty] | -0.146 | 21.695 | 20.044 | 0.253 | 25.3
[6 empty] | -0.146 | 21.685 | 20.044 | 0.244 | 24.4
[6 empty] | -0.146 | 21.675 | 20.044 | 0.235 | 23.5
[6 empty] | -0.146 | 21.665 | 20.044 | 0.225 | 22.5
[6 empty] | -0.146 | 21.655 | 20.044 | 0.216 | 21.6
[6 empty] | -0.146 | 21.645 | 20.044 | 0.207 | 20.7
[6 empty] | -0.146 | 21.635 | 20.044 | 0.198 | 19.8
[6 empty] | -0.146 | 21.625 | 20.044 | 0.188 | 18.8
[6 empty] | -0.146 | 21.615 | 20.044 | 0.179 | 17.9
[6 empty] | -0.146 | 21.605 | 20.044 | 0.17 | 17
[6 empty] | -0.146 | 21.595 | 20.044 | 0.169 | 16.9
[6 empty] | -0.146 | 21.585 | 20.044 | 0.167 | 16.7
[6 empty] | -0.146 | 21.575 | 20.044 | 0.166 | 16.6
[6 empty] | -0.146 | 21.565 | 20.044 | 0.164 | 16.4
[6 empty] | -0.146 | 21.555 | 20.044 | 0.163 | 16.3
[6 empty] | -0.146 | 21.545 | 20.044 | 0.161 | 16.1
[6 empty] | -0.146 | 21.535 | 20.044 | 0.16 | 16
[6 empty] | -0.146 | 21.525 | 20.044 | 0.158 | 15.8
[6 empty] | -0.146 | 21.515 | 20.044 | 0.157 | 15.7
[6 empty] | -0.146 | 21.505 | 20.044 | 0.155 | 15.5
[6 empty] | -0.146 | 21.495 | 20.044 | 0.153 | 15.3
[6 empty] | -0.146 | 21.485 | 20.044 | 0.152 | 15.2
[6 empty] | -0.146 | 21.475 | 20.044 | 0.15 | 15
[6 empty] | -0.146 | 21.465 | 20.044 | 0.149 | 14.9
[6 empty] | -0.146 | 21.455 | 20.044 | 0.147 | 14.7
[6 empty] | -0.146 | 21.445 | 20.044 | 0.146 | 14.6
[6 empty] | -0.146 | 21.435 | 20.044 | 0.144 | 14.4
[6 empty] | -0.146 | 21.425 | 20.044 | 0.143 | 14.3
[6 empty] | -0.146 | 21.415 | 20.044 | 0.141 | 14.1
[6 empty] | -0.146 | 21.405 | 20.044 | 0.14 | 14
[6 empty] | -0.146 | 21.395 | 20.044 | 0.144 | 14.4
[6 empty] | -0.146 | 21.385 | 20.044 | 0.148 | 14.8
[6 empty] | -0.146 | 21.375 | 20.043 | 0.152 | 15.2
[6 empty] | -0.146 | 21.365 | 20.043 | 0.156 | 15.6
[6 empty] | -0.146 | 21.355 | 20.043 | 0.16 | 16
[6 empty] | -0.146 | 21.345 | 20.043 | 0.164 | 16.4
[6 empty] | -0.146 | 21.335 | 20.043 | 0.168 | 16.8
[6 empty] | -0.146 | 21.325 | 20.043 | 0.172 | 17.2
[6 empty] | -0.146 | 21.315 | 20.043 | 0.176 | 17.6
[6 empty] | -0.146 | 21.305 | 20.043 | 0.18 | 18
[6 empty] | -0.146 | 21.295 | 20.043 | 0.183 | 18.3
[6 empty] | -0.146 | 21.285 | 20.043 | 0.187 | 18.7
[6 empty] | -0.146 | 21.275 | 20.043 | 0.191 | 19.1
[6 empty] | -0.146 | 21.265 | 20.043 | 0.195 | 19.5
[6 empty] | -0.146 | 21.255 | 20.043 | 0.199 | 19.9
[6 empty] | -0.146 | 21.245 | 20.043 | 0.203 | 20.3
[6 empty] | -0.146 | 21.235 | 20.043 | 0.207 | 20.7
[6 empty] | -0.146 | 21.225 | 20.043 | 0.211 | 21.1
[6 empty] | -0.146 | 21.215 | 20.043 | 0.215 | 21.5
[6 empty] | -0.146 | 21.205 | 20.043 | 0.219 | 21.9
[6 empty] | -0.146 | 21.195 | 20.043 | 0.228 | 22.8
[6 empty] | -0.146 | 21.185 | 20.043 | 0.236 | 23.6
[6 empty] | -0.146 | 21.175 | 20.043 | 0.244 | 24.4
[6 empty] | -0.146 | 21.165 | 20.043 | 0.253 | 25.3
[6 empty] | -0.146 | 21.155 | 20.043 | 0.261 | 26.1
[6 empty] | -0.146 | 21.145 | 20.043 | 0.27 | 27
[6 empty] | -0.146 | 21.135 | 20.043 | 0.278 | 27.8
[6 empty] | -0.146 | 21.125 | 20.043 | 0.287 | 28.7
[6 empty] | -0.146 | 21.115 | 20.043 | 0.295 | 29.5
[6 empty] | -0.146 | 21.105 | 20.043 | 0.303 | 30.3
[6 empty] | -0.146 | 21.095 | 20.043 | 0.312 | 31.2
[6 empty] | -0.146 | 21.085 | 20.043 | 0.32 | 32
[6 empty] | -0.146 | 21.075 | 20.043 | 0.329 | 32.9
[6 empty] | -0.146 | 21.065 | 20.043 | 0.337 | 33.7
[6 empty] | -0.146 | 21.055 | 20.043 | 0.345 | 34.5
[6 empty] | -0.146 | 21.045 | 20.043 | 0.354 | 35.4
[6 empty] | -0.146 | 21.035 | 20.043 | 0.362 | 36.2
[6 empty] | -0.146 | 21.025 | 20.043 | 0.371 | 37.1
[6 empty] | -0.146 | 21.015 | 20.043 | 0.379 | 37.9
[6 empty] | -0.146 | 21.005 | 20.043 | 0.387 | 38.7
[6 empty] | -0.146 | 20.995 | 20.043 | 0.388 | 38.8
[6 empty] | -0.146 | 20.985 | 20.043 | 0.388 | 38.8
[6 empty] | -0.146 | 20.975 | 20.043 | 0.389 | 38.9
[6 empty] | -0.146 | 20.965 | 20.043 | 0.39 | 39
[6 empty] | -0.146 | 20.955 | 20.043 | 0.39 | 39
[6 empty] | -0.146 | 20.945 | 20.043 | 0.391 | 39.1
[6 empty] | -0.146 | 20.935 | 20.043 | 0.392 | 39.2
[6 empty] | -0.146 | 20.925 | 20.043 | 0.393 | 39.3
[6 empty] | -0.146 | 20.915 | 20.043 | 0.393 | 39.3
[6 empty] | -0.146 | 20.905 | 20.043 | 0.394 | 39.4
[6 empty] | -0.146 | 20.895 | 20.043 | 0.395 | 39.5
[6 empty] | -0.146 | 20.885 | 20.043 | 0.395 | 39.5
[6 empty] | -0.146 | 20.875 | 20.043 | 0.396 | 39.6
[6 empty] | -0.146 | 20.865 | 20.043 | 0.397 | 39.7
[6 empty] | -0.146 | 20.855 | 20.043 | 0.397 | 39.7
[6 empty] | -0.146 | 20.845 | 20.043 | 0.398 | 39.8
[6 empty] | -0.146 | 20.835 | 20.043 | 0.399 | 39.9
[6 empty] | -0.146 | 20.825 | 20.043 | 0.4 | 40
[6 empty] | -0.146 | 20.815 | 20.043 | 0.4 | 40
[6 empty] | -0.146 | 20.805 | 20.043 | 0.4 | 40
[6 empty] | -0.146 | 20.795 | 20.043 | 0.393 | 39.3
[6 empty] | -0.146 | 20.785 | 20.043 | 0.386 | 38.6
[6 empty] | -0.146 | 20.775 | 20.043 | 0.378 | 37.8
[6 empty] | -0.146 | 20.765 | 20.043 | 0.371 | 37.1
[6 empty] | -0.146 | 20.755 | 20.043 | 0.364 | 36.4
[6 empty] | -0.146 | 20.745 | 20.043 | 0.357 | 35.7
[6 empty] | -0.146 | 20.735 | 20.043 | 0.349 | 34.9
[6 empty] | -0.146 | 20.725 | 20.043 | 0.342 | 34.2
[6 empty] | -0.146 | 20.715 | 20.043 | 0.335 | 33.5
[6 empty] | -0.146 | 20.705 | 20.043 | 0.327 | 32.7
[6 empty] | -0.146 | 20.695 | 20.043 | 0.32 | 32
[6 empty] | -0.146 | 20.685 | 20.042 | 0.313 | 31.3
[6 empty] | -0.146 | 20.675 | 20.042 | 0.305 | 30.5
[6 empty] | -0.146 | 20.665 | 20.042 | 0.298 | 29.8
[6 empty] | -0.146 | 20.655 | 20.042 | 0.291 | 29.1
[6 empty] | -0.146 | 20.645 | 20.042 | 0.284 | 28.4
[6 empty] | -0.146 | 20.635 | 20.042 | 0.276 | 27.6
[6 empty] | -0.146 | 20.625 | 20.042 | 0.269 | 26.9
[6 empty] | -0.146 | 20.615 | 20.042 | 0.262 | 26.2
[6 empty] | -0.146 | 20.605 | 20.042 | 0.254 | 25.4
[6 empty] | -0.146 | 20.595 | 20.042 | 0.249 | 24.9
[6 empty] | -0.146 | 20.585 | 20.042 | 0.243 | 24.3
[6 empty] | -0.146 | 20.575 | 20.042 | 0.238 | 23.8
[6 empty] | -0.146 | 20.565 | 20.042 | 0.232 | 23.2
[6 empty] | -0.146 | 20.555 | 20.042 | 0.226 | 22.6
[6 empty] | -0.146 | 20.545 | 20.042 | 0.221 | 22.1
[6 empty] | -0.146 | 20.535 | 20.042 | 0.215 | 21.5
[6 empty] | -0.146 | 20.525 | 20.042 | 0.209 | 20.9
[6 empty] | -0.146 | 20.515 | 20.042 | 0.204 | 20.4
[6 empty] | -0.146 | 20.505 | 20.042 | 0.198 | 19.8
[6 empty] | -0.146 | 20.495 | 20.042 | 0.193 | 19.3
[6 empty] | -0.146 | 20.485 | 20.042 | 0.187 | 18.7
[6 empty] | -0.146 | 20.475 | 20.042 | 0.181 | 18.1
[6 empty] | -0.146 | 20.465 | 20.042 | 0.176 | 17.6
[6 empty] | -0.146 | 20.455 | 20.042 | 0.17 | 17
[6 empty] | -0.146 | 20.445 | 20.042 | 0.164 | 16.4
[6 empty] | -0.146 | 20.435 | 20.042 | 0.159 | 15.9
[6 empty] | -0.146 | 20.425 | 20.042 | 0.153 | 15.3
[6 empty] | -0.146 | 20.415 | 20.042 | 0.147 | 14.7
[6 empty] | -0.146 | 20.405 | 20.042 | 0.142 | 14.2
[6 empty] | -0.146 | 20.395 | 20.042 | 0.143 | 14.3
[6 empty] | -0.146 | 20.385 | 20.042 | 0.144 | 14.4
[6 empty] | -0.146 | 20.375 | 20.042 | 0.145 | 14.5
[6 empty] | -0.146 | 20.365 | 20.042 | 0.145 | 14.5
[6 empty] | -0.146 | 20.355 | 20.042 | 0.146 | 14.6
[6 empty] | -0.146 | 20.345 | 20.042 | 0.147 | 14.7
[6 empty] | -0.146 | 20.335 | 20.042 | 0.148 | 14.8
[6 empty] | -0.146 | 20.325 | 20.042 | 0.148 | 14.8
[6 empty] | -0.146 | 20.315 | 20.042 | 0.149 | 14.9
[6 empty] | -0.146 | 20.305 | 20.042 | 0.15 | 15
[6 empty] | -0.146 | 20.295 | 20.042 | 0.151 | 15.1
[6 empty] | -0.146 | 20.285 | 20.042 | 0.151 | 15.1
[6 empty] | -0.146 | 20.275 | 20.042 | 0.152 | 15.2
[6 empty] | -0.146 | 20.265 | 20.042 | 0.153 | 15.3
[6 empty] | -0.146 | 20.255 | 20.042 | 0.154 | 15.4
[6 empty] | -0.146 | 20.245 | 20.042 | 0.154 | 15.4
[6 empty] | -0.146 | 20.235 | 20.042 | 0.155 | 15.5
[6 empty] | -0.146 | 20.225 | 20.042 | 0.156 | 15.6
[6 empty] | -0.146 | 20.215 | 20.042 | 0.157 | 15.7
[6 empty] | -0.146 | 20.205 | 20.042 | 0.158 | 15.8
[6 empty] | -0.146 | 20.195 | 20.042 | 0.166 | 16.6
[6 empty] | -0.146 | 20.185 | 20.042 | 0.174 | 17.4
[6 empty] | -0.146 | 20.175 | 20.042 | 0.182 | 18.2
[6 empty] | -0.146 | 20.165 | 20.042 | 0.19 | 19
[6 empty] | -0.146 | 20.155 | 20.042 | 0.198 | 19.8
[6 empty] | -0.146 | 20.145 | 20.042 | 0.206 | 20.6
[6 empty] | -0.146 | 20.135 | 20.042 | 0.213 | 21.3
[6 empty] | -0.146 | 20.125 | 20.042 | 0.221 | 22.1
[6 empty] | -0.146 | 20.115 | 20.042 | 0.229 | 22.9
[6 empty] | -0.146 | 20.105 | 20.042 | 0.237 | 23.7
[6 empty] | -0.146 | 20.095 | 20.042 | 0.245 | 24.5
[6 empty] | -0.146 | 20.085 | 20.042 | 0.253 | 25.3
[6 empty] | -0.146 | 20.075 | 20.042 | 0.261 | 26.1
[6 empty] | -0.146 | 20.065 | 20.042 | 0.269 | 26.9
[6 empty] | -0.146 | 20.055 | 20.042 | 0.277 | 27.7
[6 empty] | -0.146 | 20.045 | 20.042 | 0.285 | 28.5
[6 empty] | -0.146 | 20.035 | 20.042 | 0.293 | 29.3
[6 empty] | -0.146 | 20.025 | 20.042 | 0.301 | 30.1
[6 empty] | -0.146 | 20.015 | 20.042 | 0.309 | 30.9
[6 empty] | -0.146 | 20.005 | 20.042 | 0.316 | 31.6
[6 empty] | -0.146 | 19.995 | 20.041 | 0.321 | 32.1
[6 empty] | -0.146 | 19.985 | 20.041 | 0.326 | 32.6
[6 empty] | -0.146 | 19.975 | 20.041 | 0.331 | 33.1
[6 empty] | -0.146 | 19.965 | 20.041 | 0.335 | 33.5
[6 empty] | -0.146 | 19.955 | 20.041 | 0.34 | 34
[6 empty] | -0.146 | 19.945 | 20.041 | 0.345 | 34.5
[6 empty] | -0.146 | 19.935 | 20.041 | 0.35 | 35
[6 empty] | -0.146 | 19.925 | 20.041 | 0.354 | 35.4
[6 empty] | -0.146 | 19.915 | 20.041 | 0.359 | 35.9
[6 empty] | -0.146 | 19.905 | 20.041 | 0.364 | 36.4
[6 empty] | -0.146 | 19.895 | 20.041 | 0.369 | 36.9
[6 empty] | -0.146 | 19.885 | 20.041 | 0.373 | 37.3
[6 empty] | -0.146 | 19.875 | 20.041 | 0.378 | 37.8
[6 empty] | -0.146 | 19.865 | 20.041 | 0.383 | 38.3
[6 empty] | -0.146 | 19.855 | 20.041 | 0.387 | 38.7
[6 empty] | -0.146 | 19.845 | 20.041 | 0.392 | 39.2
[6 empty] | -0.146 | 19.835 | 20.041 | 0.397 | 39.7
[6 empty] | -0.146 | 19.825 | 20.041 | 0.402 | 40.2
[6 empty] | -0.146 | 19.815 | 20.041 | 0.406 | 40.6
[6 empty] | -0.146 | 19.805 | 20.041 | 0.411 | 41.1
[6 empty] | -0.146 | 19.795 | 20.041 | 0.408 | 40.8
[6 empty] | -0.146 | 19.785 | 20.041 | 0.405 | 40.5
[6 empty] | -0.146 | 19.775 | 20.041 | 0.402 | 40.2
[6 empty] | -0.146 | 19.765 | 20.041 | 0.399 | 39.9
[6 empty] | -0.146 | 19.755 | 20.041 | 0.397 | 39.7
[6 empty] | -0.146 | 19.745 | 20.041 | 0.394 | 39.4
[6 empty] | -0.146 | 19.735 | 20.041 | 0.391 | 39.1
[6 empty] | -0.146 | 19.725 | 20.041 | 0.388 | 38.8
[6 empty] | -0.146 | 19.715 | 20.041 | 0.385 | 38.5
[6 empty] | -0.146 | 19.705 | 20.041 | 0.382 | 38.2
[6 empty] | -0.146 | 19.695 | 20.041 | 0.379 | 37.9
[6 empty] | -0.146 | 19.685 | 20.041 | 0.377 | 37.7
[6 empty] | -0.146 | 19.675 | 20.041 | 0.374 | 37.4
[6 empty] | -0.146 | 19.665 | 20.041 | 0.371 | 37.1
[6 empty] | -0.146 | 19.655 | 20.041 | 0.368 | 36.8
[6 empty] | -0.146 | 19.645 | 20.041 | 0.365 | 36.5
[6 empty] | -0.146 | 19.635 | 20.041 | 0.362 | 36.2
[6 empty] | -0.146 | 19.625 | 20.041 | 0.36 | 36
[6 empty] | -0.146 | 19.615 | 20.041 | 0.357 | 35.7
[6 empty] | -0.146 | 19.605 | 20.041 | 0.354 | 35.4
[6 empty] | -0.146 | 19.595 | 20.041 | 0.344 | 34.4
[6 empty] | -0.146 | 19.585 | 20.041 | 0.335 | 33.5
[6 empty] | -0.146 | 19.575 | 20.041 | 0.325 | 32.5
[6 empty] | -0.146 | 19.565 | 20.041 | 0.316 | 31.6
[6 empty] | -0.146 | 19.555 | 20.041 | 0.307 | 30.7
[6 empty] | -0.146 | 19.545 | 20.041 | 0.297 | 29.7
[6 empty] | -0.146 | 19.535 | 20.041 | 0.288 | 28.8
[6 empty] | -0.146 | 19.525 | 20.041 | 0.278 | 27.8
[6 empty] | -0.146 | 19.515 | 20.041 | 0.269 | 26.9
[6 empty] | -0.146 | 19.505 | 20.041 | 0.26 | 26
[6 empty] | -0.146 | 19.495 | 20.041 | 0.25 | 25
[6 empty] | -0.146 | 19.485 | 20.041 | 0.241 | 24.1
[6 empty] | -0.146 | 19.475 | 20.041 | 0.231 | 23.1
[6 empty] | -0.146 | 19.465 | 20.041 | 0.222 | 22.2
[6 empty] | -0.146 | 19.455 | 20.041 | 0.213 | 21.3
[6 empty] | -0.146 | 19.445 | 20.041 | 0.203 | 20.3
[6 empty] | -0.146 | 19.435 | 20.041 | 0.194 | 19.4
[6 empty] | -0.146 | 19.425 | 20.041 | 0.184 | 18.4
[6 empty] | -0.146 | 19.415 | 20.041 | 0.175 | 17.5
[6 empty] | -0.146 | 19.405 | 20.041 | 0.166 | 16.6
[6 empty] | -0.146 | 19.395 | 20.041 | 0.165 | 16.5
[6 empty] | -0.146 | 19.385 | 20.041 | 0.163 | 16.3
[6 empty] | -0.146 | 19.375 | 20.041 | 0.162 | 16.2
[6 empty] | -0.146 | 19.365 | 20.041 | 0.16 | 16
[6 empty] | -0.146 | 19.355 | 20.041 | 0.159 | 15.9
[6 empty] | -0.146 | 19.345 | 20.041 | 0.157 | 15.7
[6 empty] | -0.146 | 19.335 | 20.041 | 0.156 | 15.6
[6 empty] | -0.146 | 19.325 | 20.041 | 0.154 | 15.4
[6 empty] | -0.146 | 19.315 | 20.041 | 0.153 | 15.3
[6 empty] | -0.146 | 19.305 | 20.04 | 0.151 | 15.1
[6 empty] | -0.146 | 19.295 | 20.04 | 0.15 | 15
[6 empty] | -0.146 | 19.285 | 20.04 | 0.148 | 14.8
[6 empty] | -0.146 | 19.275 | 20.04 | 0.147 | 14.7
[6 empty] | -0.146 | 19.265 | 20.04 | 0.145 | 14.5
[6 empty] | -0.146 | 19.255 | 20.04 | 0.144 | 14.4
[6 empty] | -0.146 | 19.245 | 20.04 | 0.142 | 14.2
[6 empty] | -0.146 | 19.235 | 20.04 | 0.141 | 14.1
[6 empty] | -0.146 | 19.225 | 20.04 | 0.139 | 13.9
[6 empty] | -0.146 | 19.215 | 20.04 | 0.138 | 13.8
[6 empty] | -0.146 | 19.205 | 20.04 | 0.137 | 13.7
[6 empty] | -0.146 | 19.195 | 20.04 | 0.141 | 14.1
[6 empty] | -0.146 | 19.185 | 20.04 | 0.145 | 14.5
[6 empty] | -0.146 | 19.175 | 20.04 | 0.149 | 14.9
[6 empty] | -0.146 | 19.165 | 20.04 | 0.153 | 15.3
[6 empty] | -0.146 | 19.155 | 20.04 | 0.157 | 15.7
[6 empty] | -0.146 | 19.145 | 20.04 | 0.161 | 16.1
[6 empty] | -0.146 | 19.135 | 20.04 | 0.165 | 16.5
[6 empty] | -0.146 | 19.125 | 20.04 | 0.169 | 16.9
[6 empty] | -0.146 | 19.115 | 20.04 | 0.173 | 17.3
[6 empty] | -0.146 | 19.105 | 20.04 | 0.177 | 17.7
[6 empty] | -0.146 | 19.095 | 20.04 | 0.181 | 18.1
[6 empty] | -0.146 | 19.085 | 20.04 | 0.185 | 18.5
[6 empty] | -0.146 | 19.075 | 20.04 | 0.189 | 18.9
[6 empty] | -0.146 | 19.065 | 20.04 | 0.193 | 19.3
[6 empty] | -0.146 | 19.055 | 20.04 | 0.197 | 19.7
[6 empty] | -0.146 | 19.045 | 20.04 | 0.201 | 20.1
[6 empty] | -0.146 | 19.035 | 20.04 | 0.205 | 20.5
[6 empty] | -0.146 | 19.025 | 20.04 | 0.209 | 20.9
[6 empty] | -0.146 | 19.015 | 20.04 | 0.213 | 21.3
[6 empty] | -0.146 | 19.005 | 20.04 | 0.217 | 21.7
[6 empty] | -0.146 | 18.995 | 20.04 | 0.226 | 22.6
[6 empty] | -0.146 | 18.985 | 20.04 | 0.234 | 23.4
[6 empty] | -0.146 | 18.976 | 20.04 | 0.243 | 24.3
[6 empty] | -0.146 | 18.966 | 20.04 | 0.251 | 25.1
[6 empty] | -0.146 | 18.956 | 20.04 | 0.26 | 26
[6 empty] | -0.146 | 18.946 | 20.04 | 0.268 | 26.8
[6 empty] | -0.146 | 18.936 | 20.04 | 0.277 | 27.7
[6 empty] | -0.146 | 18.926 | 20.04 | 0.285 | 28.5
[6 empty] | -0.146 | 18.916 | 20.04 | 0.294 | 29.4
[6 empty] | -0.146 | 18.906 | 20.04 | 0.302 | 30.2
[6 empty] | -0.146 | 18.896 | 20.04 | 0.311 | 31.1
[6 empty] | -0.146 | 18.886 | 20.04 | 0.319 | 31.9
[6 empty] | -0.146 | 18.876 | 20.04 | 0.328 | 32.8
[6 empty] | -0.146 | 18.866 | 20.04 | 0.336 | 33.6
[6 empty] | -0.146 | 18.856 | 20.04 | 0.345 | 34.5
[6 empty] | -0.146 | 18.846 | 20.04 | 0.353 | 35.3
[6 empty] | -0.146 | 18.836 | 20.04 | 0.362 | 36.2
[6 empty] | -0.146 | 18.826 | 20.04 | 0.37 | 37
[6 empty] | -0.146 | 18.816 | 20.04 | 0.379 | 37.9
[6 empty] | -0.146 | 18.806 | 20.04 | 0.387 | 38.7
[6 empty] | -0.146 | 18.796 | 20.04 | 0.388 | 38.8
[6 empty] | -0.146 | 18.786 | 20.04 | 0.388 | 38.8
[6 empty] | -0.146 | 18.776 | 20.04 | 0.389 | 38.9
[6 empty] | -0.146 | 18.766 | 20.04 | 0.389 | 38.9
[6 empty] | -0.146 | 18.756 | 20.04 | 0.39 | 39
[6 empty] | -0.146 | 18.746 | 20.04 | 0.39 | 39
[6 empty] | -0.146 | 18.736 | 20.04 | 0.391 | 39.1
[6 empty] | -0.146 | 18.726 | 20.04 | 0.392 | 39.2
[6 empty] | -0.146 | 18.716 | 20.04 | 0.392 | 39.2
[6 empty] | -0.146 | 18.706 | 20.04 | 0.393 | 39.3
[6 empty] | -0.146 | 18.696 | 20.04 | 0.393 | 39.3
[6 empty] | -0.146 | 18.686 | 20.04 | 0.394 | 39.4
[6 empty] | -0.146 | 18.676 | 20.04 | 0.394 | 39.4
[6 empty] | -0.146 | 18.666 | 20.04 | 0.395 | 39.5
[6 empty] | -0.146 | 18.656 | 20.04 | 0.396 | 39.6
[6 empty] | -0.146 | 18.646 | 20.04 | 0.396 | 39.6
[6 empty] | -0.146 | 18.636 | 20.04 | 0.397 | 39.7
[6 empty] | -0.146 | 18.626 | 20.04 | 0.397 | 39.7
[6 empty] | -0.146 | 18.616 | 20.039 | 0.398 | 39.8
[6 empty] | -0.146 | 18.606 | 20.039 | 0.398 | 39.8
[6 empty] | -0.146 | 18.596 | 20.039 | 0.391 | 39.1
[6 empty] | -0.146 | 18.586 | 20.039 | 0.383 | 38.3
[6 empty] | -0.146 | 18.576 | 20.039 | 0.376 | 37.6
[6 empty] | -0.146 | 18.566 | 20.039 | 0.368 | 36.8
[6 empty] | -0.146 | 18.556 | 20.039 | 0.361 | 36.1
[6 empty] | -0.146 | 18.546 | 20.039 | 0.354 | 35.4
[6 empty] | -0.146 | 18.536 | 20.039 | 0.346 | 34.6
[6 empty] | -0.146 | 18.526 | 20.039 | 0.339 | 33.9
[6 empty] | -0.146 | 18.516 | 20.039 | 0.331 | 33.1
[6 empty] | -0.146 | 18.506 | 20.039 | 0.324 | 32.4
[6 empty] | -0.146 | 18.496 | 20.039 | 0.316 | 31.6
[6 empty] | -0.146 | 18.486 | 20.039 | 0.309 | 30.9
[6 empty] | -0.146 | 18.476 | 20.039 | 0.301 | 30.1
[6 empty] | -0.146 | 18.466 | 20.039 | 0.294 | 29.4
[6 empty] | -0.146 | 18.456 | 20.039 | 0.287 | 28.7
[6 empty] | -0.146 | 18.446 | 20.039 | 0.279 | 27.9
[6 empty] | -0.146 | 18.436 | 20.039 | 0.272 | 27.2
[6 empty] | -0.146 | 18.426 | 20.039 | 0.264 | 26.4
[6 empty] | -0.146 | 18.416 | 20.039 | 0.257 | 25.7
[6 empty] | -0.146 | 18.406 | 20.039 | 0.249 | 24.9
[6 empty] | -0.146 | 18.396 | 20.039 | 0.244 | 24.4
[6 empty] | -0.146 | 18.386 | 20.039 | 0.238 | 23.8
[6 empty] | -0.146 | 18.376 | 20.039 | 0.233 | 23.3
[6 empty] | -0.146 | 18.366 | 20.039 | 0.227 | 22.7
[6 empty] | -0.146 | 18.356 | 20.039 | 0.221 | 22.1
[6 empty] | -0.146 | 18.346 | 20.039 | 0.216 | 21.6
[6 empty] | -0.146 | 18.336 | 20.039 | 0.21 | 21
[6 empty] | -0.146 | 18.326 | 20.039 | 0.204 | 20.4
[6 empty] | -0.146 | 18.316 | 20.039 | 0.199 | 19.9
[6 empty] | -0.146 | 18.306 | 20.039 | 0.193 | 19.3
[6 empty] | -0.146 | 18.296 | 20.039 | 0.187 | 18.7
[6 empty] | -0.146 | 18.286 | 20.039 | 0.182 | 18.2
[6 empty] | -0.146 | 18.276 | 20.039 | 0.176 | 17.6
[6 empty] | -0.146 | 18.266 | 20.039 | 0.17 | 17
[6 empty] | -0.146 | 18.256 | 20.039 | 0.165 | 16.5
[6 empty] | -0.146 | 18.246 | 20.039 | 0.159 | 15.9
[6 empty] | -0.146 | 18.236 | 20.039 | 0.153 | 15.3
[6 empty] | -0.146 | 18.226 | 20.039 | 0.148 | 14.8
[6 empty] | -0.146 | 18.216 | 20.039 | 0.142 | 14.2
[6 empty] | -0.146 | 18.206 | 20.039 | 0.137 | 13.7
[6 empty] | -0.146 | 18.196 | 20.039 | 0.138 | 13.8
[6 empty] | -0.146 | 18.186 | 20.039 | 0.138 | 13.8
[6 empty] | -0.146 | 18.176 | 20.039 | 0.139 | 13.9
[6 empty] | -0.146 | 18.166 | 20.039 | 0.14 | 14
[6 empty] | -0.146 | 18.156 | 20.039 | 0.141 | 14.1
[6 empty] | -0.146 | 18.146 | 20.039 | 0.142 | 14.2
[6 empty] | -0.146 | 18.136 | 20.039 | 0.142 | 14.2
[6 empty] | -0.146 | 18.126 | 20.039 | 0.143 | 14.3
[6 empty] | -0.146 | 18.116 | 20.039 | 0.144 | 14.4
[6 empty] | -0.146 | 18.106 | 20.039 | 0.145 | 14.5
[6 empty] | -0.146 | 18.096 | 20.039 | 0.146 | 14.6
[6 empty] | -0.146 | 18.086 | 20.039 | 0.146 | 14.6
[6 empty] | -0.146 | 18.076 | 20.039 | 0.147 | 14.7
[6 empty] | -0.146 | 18.066 | 20.039 | 0.148 | 14.8
[6 empty] | -0.146 | 18.056 | 20.039 | 0.149 | 14.9
[6 empty] | -0.146 | 18.046 | 20.039 | 0.15 | 15
[6 empty] | -0.146 | 18.036 | 20.039 | 0.15 | 15
[6 empty] | -0.146 | 18.026 | 20.039 | 0.151 | 15.1
[6 empty] | -0.146 | 18.016 | 20.039 | 0.152 | 15.2
[6 empty] | -0.146 | 18.006 | 20.039 | 0.153 | 15.3
[6 empty] | -0.146 | 17.996 | 20.039 | 0.161 | 16.1
[6 empty] | -0.146 | 17.986 | 20.039 | 0.169 | 16.9
[6 empty] | -0.146 | 17.976 | 20.039 | 0.177 | 17.7
[6 empty] | -0.146 | 17.966 | 20.039 | 0.185 | 18.5
[6 empty] | -0.146 | 17.956 | 20.039 | 0.193 | 19.3
[6 empty] | -0.146 | 17.946 | 20.039 | 0.202 | 20.2
[6 empty] | -0.146 | 17.936 | 20.039 | 0.21 | 21
[6 empty] | -0.146 | 17.926 | 20.038 | 0.218 | 21.8
[6 empty] | -0.146 | 17.916 | 20.038 | 0.226 | 22.6
[6 empty] | -0.146 | 17.906 | 20.038 | 0.234 | 23.4
[6 empty] | -0.146 | 17.896 | 20.038 | 0.242 | 24.2
[6 empty] | -0.146 | 17.886 | 20.038 | 0.25 | 25
[6 empty] | -0.146 | 17.876 | 20.038 | 0.258 | 25.8
[6 empty] | -0.146 | 17.866 | 20.038 | 0.266 | 26.6
[6 empty] | -0.146 | 17.856 | 20.038 | 0.274 | 27.4
[6 empty] | -0.146 | 17.846 | 20.038 | 0.282 | 28.2
[6 empty] | -0.146 | 17.836 | 20.038 | 0.291 | 29.1
[6 empty] | -0.146 | 17.826 | 20.038 | 0.299 | 29.9
[6 empty] | -0.146 | 17.816 | 20.038 | 0.307 | 30.7
[6 empty] | -0.146 | 17.806 | 20.038 | 0.315 | 31.5
[6 empty] | -0.146 | 17.796 | 20.038 | 0.319 | 31.9
[6 empty] | -0.146 | 17.786 | 20.038 | 0.324 | 32.4
[6 empty] | -0.146 | 17.776 | 20.038 | 0.329 | 32.9
[6 empty] | -0.146 | 17.766 | 20.038 | 0.333 | 33.3
[6 empty] | -0.146 | 17.756 | 20.038 | 0.338 | 33.8
[6 empty] | -0.146 | 17.746 | 20.038 | 0.343 | 34.3
[6 empty] | -0.146 | 17.736 | 20.038 | 0.347 | 34.7
[6 empty] | -0.146 | 17.726 | 20.038 | 0.352 | 35.2
[6 empty] | -0.146 | 17.716 | 20.038 | 0.357 | 35.7
[6 empty] | -0.146 | 17.706 | 20.038 | 0.361 | 36.1
[6 empty] | -0.146 | 17.696 | 20.038 | 0.366 | 36.6
[6 empty] | -0.146 | 17.686 | 20.038 | 0.371 | 37.1
[6 empty] | -0.146 | 17.676 | 20.038 | 0.375 | 37.5
[6 empty] | -0.146 | 17.666 | 20.038 | 0.38 | 38
[6 empty] | -0.146 | 17.656 | 20.038 | 0.385 | 38.5
[6 empty] | -0.146 | 17.646 | 20.038 | 0.389 | 38.9
[6 empty] | -0.146 | 17.636 | 20.038 | 0.394 | 39.4
[6 empty] | -0.146 | 17.626 | 20.038 | 0.399 | 39.9
[6 empty] | -0.146 | 17.616 | 20.038 | 0.403 | 40.3
[6 empty] | -0.146 | 17.606 | 20.038 | 0.408 | 40.8
[6 empty] | -0.146 | 17.596 | 20.038 | 0.405 | 40.5
[6 empty] | -0.146 | 17.586 | 20.038 | 0.402 | 40.2
[6 empty] | -0.146 | 17.576 | 20.038 | 0.399 | 39.9
[6 empty] | -0.146 | 17.566 | 20.038 | 0.396 | 39.6
[6 empty] | -0.146 | 17.556 | 20.038 | 0.393 | 39.3
[6 empty] | -0.146 | 17.546 | 20.038 | 0.39 | 39
[6 empty] | -0.146 | 17.536 | 20.038 | 0.387 | 38.7
[6 empty] | -0.146 | 17.526 | 20.038 | 0.384 | 38.4
[6 empty] | -0.146 | 17.516 | 20.038 | 0.381 | 38.1
[6 empty] | -0.146 | 17.506 | 20.038 | 0.378 | 37.8
[6 empty] | -0.146 | 17.496 | 20.038 | 0.375 | 37.5
[6 empty] | -0.146 | 17.486 | 20.038 | 0.372 | 37.2
[6 empty] | -0.146 | 17.476 | 20.038 | 0.369 | 36.9
[6 empty] | -0.146 | 17.466 | 20.038 | 0.366 | 36.6
[6 empty] | -0.146 | 17.456 | 20.038 | 0.363 | 36.3
[6 empty] | -0.146 | 17.446 | 20.038 | 0.36 | 36
[6 empty] | -0.146 | 17.436 | 20.038 | 0.357 | 35.7
[6 empty] | -0.146 | 17.426 | 20.038 | 0.354 | 35.4
[6 empty] | -0.146 | 17.416 | 20.038 | 0.35 | 35
[6 empty] | -0.146 | 17.406 | 20.038 | 0.347 | 34.7
[6 empty] | -0.146 | 17.396 | 20.038 | 0.338 | 33.8
[6 empty] | -0.146 | 17.386 | 20.038 | 0.328 | 32.8
[6 empty] | -0.146 | 17.376 | 20.038 | 0.319 | 31.9
[6 empty] | -0.146 | 17.366 | 20.038 | 0.309 | 30.9
[6 empty] | -0.146 | 17.356 | 20.038 | 0.3 | 30
[6 empty] | -0.146 | 17.346 | 20.038 | 0.291 | 29.1
[6 empty] | -0.146 | 17.336 | 20.038 | 0.281 | 28.1
[6 empty] | -0.146 | 17.326 | 20.038 | 0.272 | 27.2
[6 empty] | -0.146 | 17.316 | 20.038 | 0.262 | 26.2
[6 empty] | -0.146 | 17.306 | 20.038 | 0.253 | 25.3
[6 empty] | -0.146 | 17.296 | 20.038 | 0.243 | 24.3
[6 empty] | -0.146 | 17.286 | 20.038 | 0.234 | 23.4
[6 empty] | -0.146 | 17.276 | 20.038 | 0.224 | 22.4
[6 empty] | -0.146 | 17.266 | 20.038 | 0.215 | 21.5
[6 empty] | -0.146 | 17.256 | 20.038 | 0.205 | 20.5
[6 empty] | -0.146 | 17.246 | 20.038 | 0.196 | 19.6
[6 empty] | -0.146 | 17.236 | 20.038 | 0.187 | 18.7
[6 empty] | -0.146 | 17.226 | 20.037 | 0.177 | 17.7
[6 empty] | -0.146 | 17.216 | 20.037 | 0.168 | 16.8
[6 empty] | -0.146 | 17.206 | 20.037 | 0.158 | 15.8
[6 empty] | -0.146 | 17.196 | 20.037 | 0.157 | 15.7
[6 empty] | -0.146 | 17.186 | 20.037 | 0.155 | 15.5
[6 empty] | -0.146 | 17.176 | 20.037 | 0.154 | 15.4
[6 empty] | -0.146 | 17.166 | 20.037 | 0.153 | 15.3
[6 empty] | -0.146 | 17.156 | 20.037 | 0.151 | 15.1
[6 empty] | -0.146 | 17.146 | 20.037 | 0.15 | 15
[6 empty] | -0.146 | 17.136 | 20.037 | 0.148 | 14.8
[6 empty] | -0.146 | 17.126 | 20.037 | 0.147 | 14.7
[6 empty] | -0.146 | 17.116 | 20.037 | 0.145 | 14.5
[6 empty] | -0.146 | 17.106 | 20.037 | 0.144 | 14.4
[6 empty] | -0.146 | 17.096 | 20.037 | 0.143 | 14.3
[6 empty] | -0.146 | 17.086 | 20.037 | 0.141 | 14.1
[6 empty] | -0.146 | 17.076 | 20.037 | 0.14 | 14
[6 empty] | -0.146 | 17.066 | 20.037 | 0.138 | 13.8
[6 empty] | -0.146 | 17.056 | 20.037 | 0.137 | 13.7
[6 empty] | -0.146 | 17.046 | 20.037 | 0.135 | 13.5
[6 empty] | -0.146 | 17.036 | 20.037 | 0.134 | 13.4
[6 empty] | -0.146 | 17.026 | 20.037 | 0.133 | 13.3
[6 empty] | -0.146 | 17.016 | 20.037 | 0.131 | 13.1
[6 empty] | -0.146 | 17.006 | 20.037 | 0.13 | 13
[6 empty] | -0.146 | 16.996 | 20.037 | 0.134 | 13.4
[6 empty] | -0.146 | 16.986 | 20.037 | 0.138 | 13.8
[6 empty] | -0.146 | 16.976 | 20.037 | 0.142 | 14.2
[6 empty] | -0.146 | 16.966 | 20.037 | 0.146 | 14.6
[6 empty] | -0.146 | 16.956 | 20.037 | 0.151 | 15.1
[6 empty] | -0.146 | 16.946 | 20.037 | 0.155 | 15.5
[6 empty] | -0.146 | 16.936 | 20.037 | 0.159 | 15.9
[6 empty] | -0.146 | 16.926 | 20.037 | 0.163 | 16.3
[6 empty] | -0.146 | 16.916 | 20.037 | 0.167 | 16.7
[6 empty] | -0.146 | 16.906 | 20.037 | 0.171 | 17.1
[6 empty] | -0.146 | 16.896 | 20.037 | 0.175 | 17.5
[6 empty] | -0.146 | 16.886 | 20.037 | 0.18 | 18
[6 empty] | -0.146 | 16.876 | 20.037 | 0.184 | 18.4
[6 empty] | -0.146 | 16.866 | 20.037 | 0.188 | 18.8
[6 empty] | -0.146 | 16.856 | 20.037 | 0.192 | 19.2
[6 empty] | -0.146 | 16.846 | 20.037 | 0.196 | 19.6
[6 empty] | -0.146 | 16.836 | 20.037 | 0.2 | 20
[6 empty] | -0.146 | 16.826 | 20.037 | 0.205 | 20.5
[6 empty] | -0.146 | 16.816 | 20.037 | 0.209 | 20.9
[6 empty] | -0.146 | 16.806 | 20.037 | 0.213 | 21.3
[6 empty] | -0.146 | 16.796 | 20.037 | 0.221 | 22.1
[6 empty] | -0.146 | 16.786 | 20.037 | 0.23 | 23
[6 empty] | -0.146 | 16.776 | 20.037 | 0.238 | 23.8
[6 empty] | -0.146 | 16.766 | 20.037 | 0.247 | 24.7
[6 empty] | -0.146 | 16.756 | 20.037 | 0.256 | 25.6
[6 empty] | -0.146 | 16.746 | 20.037 | 0.264 | 26.4
[6 empty] | -0.146 | 16.736 | 20.037 | 0.273 | 27.3
[6 empty] | -0.146 | 16.726 | 20.037 | 0.281 | 28.1
[6 empty] | -0.146 | 16.716 | 20.037 | 0.29 | 29
[6 empty] | -0.146 | 16.706 | 20.037 | 0.298 | 29.8
[6 empty] | -0.146 | 16.696 | 20.037 | 0.307 | 30.7
[6 empty] | -0.146 | 16.686 | 20.037 | 0.315 | 31.5
[6 empty] | -0.146 | 16.676 | 20.037 | 0.324 | 32.4
[6 empty] | -0.146 | 16.666 | 20.037 | 0.332 | 33.2
[6 empty] | -0.146 | 16.656 | 20.037 | 0.341 | 34.1
[6 empty] | -0.146 | 16.646 | 20.037 | 0.349 | 34.9
[6 empty] | -0.146 | 16.636 | 20.037 | 0.358 | 35.8
[6 empty] | -0.146 | 16.626 | 20.037 | 0.366 | 36.6
[6 empty] | -0.146 | 16.616 | 20.037 | 0.375 | 37.5
[6 empty] | -0.146 | 16.606 | 20.037 | 0.383 | 38.3
[6 empty] | -0.146 | 16.596 | 20.037 | 0.384 | 38.4
[6 empty] | -0.146 | 16.586 | 20.037 | 0.384 | 38.4
[6 empty] | -0.146 | 16.576 | 20.037 | 0.385 | 38.5
[6 empty] | -0.146 | 16.566 | 20.037 | 0.385 | 38.5
[6 empty] | -0.146 | 16.556 | 20.037 | 0.386 | 38.6
[6 empty] | -0.146 | 16.546 | 20.037 | 0.386 | 38.6
[6 empty] | -0.146 | 16.536 | 20.036 | 0.386 | 38.6
[6 empty] | -0.146 | 16.526 | 20.036 | 0.387 | 38.7
[6 empty] | -0.146 | 16.516 | 20.036 | 0.387 | 38.7
[6 empty] | -0.146 | 16.506 | 20.036 | 0.388 | 38.8
[6 empty] | -0.146 | 16.496 | 20.036 | 0.388 | 38.8
[6 empty] | -0.146 | 16.486 | 20.036 | 0.389 | 38.9
[6 empty] | -0.146 | 16.476 | 20.036 | 0.389 | 38.9
[6 empty] | -0.146 | 16.466 | 20.036 | 0.39 | 39
[6 empty] | -0.146 | 16.456 | 20.036 | 0.39 | 39
[6 empty] | -0.146 | 16.446 | 20.036 | 0.391 | 39.1
[6 empty] | -0.146 | 16.436 | 20.036 | 0.391 | 39.1
[6 empty] | -0.146 | 16.426 | 20.036 | 0.391 | 39.1
[6 empty] | -0.146 | 16.416 | 20.036 | 0.392 | 39.2
[6 empty] | -0.146 | 16.406 | 20.036 | 0.392 | 39.2
[6 empty] | -0.146 | 16.396 | 20.036 | 0.385 | 38.5
[6 empty] | -0.146 | 16.386 | 20.036 | 0.377 | 37.7
[6 empty] | -0.146 | 16.376 | 20.036 | 0.369 | 36.9
[6 empty] | -0.146 | 16.366 | 20.036 | 0.362 | 36.2
[6 empty] | -0.146 | 16.356 | 20.036 | 0.354 | 35.4
[6 empty] | -0.146 | 16.346 | 20.036 | 0.346 | 34.6
[6 empty] | -0.146 | 16.336 | 20.036 | 0.339 | 33.9
[6 empty] | -0.146 | 16.326 | 20.036 | 0.331 | 33.1
[6 empty] | -0.146 | 16.316 | 20.036 | 0.324 | 32.4
[6 empty] | -0.146 | 16.306 | 20.036 | 0.316 | 31.6
[6 empty] | -0.146 | 16.296 | 20.036 | 0.308 | 30.8
[6 empty] | -0.146 | 16.286 | 20.036 | 0.301 | 30.1
[6 empty] | -0.146 | 16.276 | 20.036 | 0.293 | 29.3
[6 empty] | -0.146 | 16.266 | 20.036 | 0.285 | 28.5
[6 empty] | -0.146 | 16.256 | 20.036 | 0.278 | 27.8
[6 empty] | -0.146 | 16.246 | 20.036 | 0.27 | 27
[6 empty] | -0.146 | 16.236 | 20.036 | 0.262 | 26.2
[6 empty] | -0.146 | 16.226 | 20.036 | 0.255 | 25.5
[6 empty] | -0.146 | 16.216 | 20.036 | 0.247 | 24.7
[6 empty] | -0.146 | 16.206 | 20.036 | 0.24 | 24
[6 empty] | -0.146 | 16.196 | 20.036 | 0.234 | 23.4
[6 empty] | -0.146 | 16.186 | 20.036 | 0.228 | 22.8
[6 empty] | -0.146 | 16.176 | 20.036 | 0.223 | 22.3
[6 empty] | -0.146 | 16.166 | 20.036 | 0.217 | 21.7
[6 empty] | -0.146 | 16.156 | 20.036 | 0.211 | 21.1
[6 empty] | -0.146 | 16.146 | 20.036 | 0.206 | 20.6
[6 empty] | -0.146 | 16.136 | 20.036 | 0.2 | 20
[6 empty] | -0.146 | 16.126 | 20.036 | 0.194 | 19.4
[6 empty] | -0.146 | 16.116 | 20.036 | 0.189 | 18.9
[6 empty] | -0.146 | 16.106 | 20.036 | 0.183 | 18.3
[6 empty] | -0.146 | 16.096 | 20.036 | 0.177 | 17.7
[6 empty] | -0.146 | 16.086 | 20.036 | 0.171 | 17.1
[6 empty] | -0.146 | 16.076 | 20.036 | 0.166 | 16.6
[6 empty] | -0.146 | 16.066 | 20.036 | 0.16 | 16
[6 empty] | -0.146 | 16.056 | 20.036 | 0.154 | 15.4
[6 empty] | -0.146 | 16.046 | 20.036 | 0.149 | 14.9
[6 empty] | -0.146 | 16.036 | 20.036 | 0.143 | 14.3
[6 empty] | -0.146 | 16.026 | 20.036 | 0.137 | 13.7
[6 empty] | -0.146 | 16.016 | 20.036 | 0.132 | 13.2
[6 empty] | -0.146 | 16.006 | 20.036 | 0.126 | 12.6
[6 empty] | -0.146 | 15.996 | 20.036 | 0.127 | 12.7
[6 empty] | -0.146 | 15.986 | 20.036 | 0.128 | 12.8
[6 empty] | -0.146 | 15.977 | 20.036 | 0.129 | 12.9
[6 empty] | -0.146 | 15.967 | 20.036 | 0.13 | 13
[6 empty] | -0.146 | 15.957 | 20.036 | 0.13 | 13
[6 empty] | -0.146 | 15.947 | 20.036 | 0.131 | 13.1
[6 empty] | -0.146 | 15.937 | 20.036 | 0.132 | 13.2
[6 empty] | -0.146 | 15.927 | 20.036 | 0.133 | 13.3
[6 empty] | -0.146 | 15.917 | 20.036 | 0.134 | 13.4
[6 empty] | -0.146 | 15.907 | 20.036 | 0.135 | 13.5
[6 empty] | -0.146 | 15.897 | 20.036 | 0.136 | 13.6
[6 empty] | -0.146 | 15.887 | 20.036 | 0.136 | 13.6
[6 empty] | -0.146 | 15.877 | 20.036 | 0.137 | 13.7
[6 empty] | -0.146 | 15.867 | 20.036 | 0.138 | 13.8
[6 empty] | -0.146 | 15.857 | 20.036 | 0.139 | 13.9
[6 empty] | -0.146 | 15.847 | 20.035 | 0.14 | 14
[6 empty] | -0.146 | 15.837 | 20.035 | 0.141 | 14.1
[6 empty] | -0.146 | 15.827 | 20.035 | 0.142 | 14.2
[6 empty] | -0.146 | 15.817 | 20.035 | 0.143 | 14.3
[6 empty] | -0.146 | 15.807 | 20.035 | 0.143 | 14.3
[6 empty] | -0.146 | 15.797 | 20.035 | 0.152 | 15.2
[6 empty] | -0.146 | 15.787 | 20.035 | 0.16 | 16
[6 empty] | -0.146 | 15.777 | 20.035 | 0.168 | 16.8
[6 empty] | -0.146 | 15.767 | 20.035 | 0.176 | 17.6
[6 empty] | -0.146 | 15.757 | 20.035 | 0.184 | 18.4
[6 empty] | -0.146 | 15.747 | 20.035 | 0.193 | 19.3
[6 empty] | -0.146 | 15.737 | 20.035 | 0.201 | 20.1
[6 empty] | -0.146 | 15.727 | 20.035 | 0.209 | 20.9
[6 empty] | -0.146 | 15.717 | 20.035 | 0.217 | 21.7
[6 empty] | -0.146 | 15.707 | 20.035 | 0.225 | 22.5
[6 empty] | -0.146 | 15.697 | 20.035 | 0.234 | 23.4
[6 empty] | -0.146 | 15.687 | 20.035 | 0.242 | 24.2
[6 empty] | -0.146 | 15.677 | 20.035 | 0.25 | 25
[6 empty] | -0.146 | 15.667 | 20.035 | 0.258 | 25.8
[6 empty] | -0.146 | 15.657 | 20.035 | 0.266 | 26.6
[6 empty] | -0.146 | 15.647 | 20.035 | 0.275 | 27.5
[6 empty] | -0.146 | 15.637 | 20.035 | 0.283 | 28.3
[6 empty] | -0.146 | 15.627 | 20.035 | 0.291 | 29.1
[6 empty] | -0.146 | 15.617 | 20.035 | 0.299 | 29.9
[6 empty] | -0.146 | 15.607 | 20.035 | 0.308 | 30.8
[6 empty] | -0.146 | 15.597 | 20.035 | 0.312 | 31.2
[6 empty] | -0.146 | 15.587 | 20.035 | 0.317 | 31.7
[6 empty] | -0.146 | 15.577 | 20.035 | 0.321 | 32.1
[6 empty] | -0.146 | 15.567 | 20.035 | 0.326 | 32.6
[6 empty] | -0.146 | 15.557 | 20.035 | 0.33 | 33
[6 empty] | -0.146 | 15.547 | 20.035 | 0.335 | 33.5
[6 empty] | -0.146 | 15.537 | 20.035 | 0.339 | 33.9
[6 empty] | -0.146 | 15.527 | 20.035 | 0.344 | 34.4
[6 empty] | -0.146 | 15.517 | 20.035 | 0.348 | 34.8
[6 empty] | -0.146 | 15.507 | 20.035 | 0.353 | 35.3
[6 empty] | -0.146 | 15.497 | 20.035 | 0.358 | 35.8
[6 empty] | -0.146 | 15.487 | 20.035 | 0.362 | 36.2
[6 empty] | -0.146 | 15.477 | 20.035 | 0.367 | 36.7
[6 empty] | -0.146 | 15.467 | 20.035 | 0.371 | 37.1
[6 empty] | -0.146 | 15.457 | 20.035 | 0.376 | 37.6
[6 empty] | -0.146 | 15.447 | 20.035 | 0.38 | 38
[6 empty] | -0.146 | 15.437 | 20.035 | 0.385 | 38.5
[6 empty] | -0.146 | 15.427 | 20.035 | 0.389 | 38.9
[6 empty] | -0.146 | 15.417 | 20.035 | 0.394 | 39.4
[6 empty] | -0.146 | 15.407 | 20.035 | 0.398 | 39.8
[6 empty] | -0.146 | 15.397 | 20.035 | 0.395 | 39.5
[6 empty] | -0.146 | 15.387 | 20.035 | 0.392 | 39.2
[6 empty] | -0.146 | 15.377 | 20.035 | 0.389 | 38.9
[6 empty] | -0.146 | 15.367 | 20.035 | 0.386 | 38.6
[6 empty] | -0.146 | 15.357 | 20.035 | 0.382 | 38.2
[6 empty] | -0.146 | 15.347 | 20.035 | 0.379 | 37.9
[6 empty] | -0.146 | 15.337 | 20.035 | 0.376 | 37.6
[6 empty] | -0.146 | 15.327 | 20.035 | 0.373 | 37.3
[6 empty] | -0.146 | 15.317 | 20.035 | 0.369 | 36.9
[6 empty] | -0.146 | 15.307 | 20.035 | 0.366 | 36.6
[6 empty] | -0.146 | 15.297 | 20.035 | 0.363 | 36.3
[6 empty] | -0.146 | 15.287 | 20.035 | 0.36 | 36
[6 empty] | -0.146 | 15.277 | 20.035 | 0.357 | 35.7
[6 empty] | -0.146 | 15.267 | 20.035 | 0.353 | 35.3
[6 empty] | -0.146 | 15.257 | 20.035 | 0.35 | 35
[6 empty] | -0.146 | 15.247 | 20.035 | 0.347 | 34.7
[6 empty] | -0.146 | 15.237 | 20.035 | 0.344 | 34.4
[6 empty] | -0.146 | 15.227 | 20.035 | 0.34 | 34
[6 empty] | -0.146 | 15.217 | 20.035 | 0.337 | 33.7
[6 empty] | -0.146 | 15.207 | 20.035 | 0.334 | 33.4
[6 empty] | -0.146 | 15.197 | 20.035 | 0.324 | 32.4
[6 empty] | -0.146 | 15.187 | 20.035 | 0.315 | 31.5
[6 empty] | -0.146 | 15.177 | 20.035 | 0.305 | 30.5
[6 empty] | -0.146 | 15.167 | 20.035 | 0.296 | 29.6
[6 empty] | -0.146 | 15.157 | 20.034 | 0.286 | 28.6
[6 empty] | -0.146 | 15.147 | 20.034 | 0.276 | 27.6
[6 empty] | -0.146 | 15.137 | 20.034 | 0.267 | 26.7
[6 empty] | -0.146 | 15.127 | 20.034 | 0.257 | 25.7
[6 empty] | -0.146 | 15.117 | 20.034 | 0.248 | 24.8
[6 empty] | -0.146 | 15.107 | 20.034 | 0.238 | 23.8
[6 empty] | -0.146 | 15.097 | 20.034 | 0.228 | 22.8
[6 empty] | -0.146 | 15.087 | 20.034 | 0.219 | 21.9
[6 empty] | -0.146 | 15.077 | 20.034 | 0.209 | 20.9
[6 empty] | -0.146 | 15.067 | 20.034 | 0.2 | 20
[6 empty] | -0.146 | 15.057 | 20.034 | 0.19 | 19
[6 empty] | -0.146 | 15.047 | 20.034 | 0.18 | 18
[6 empty] | -0.146 | 15.037 | 20.034 | 0.171 | 17.1
[6 empty] | -0.146 | 15.027 | 20.034 | 0.161 | 16.1
[6 empty] | -0.146 | 15.017 | 20.034 | 0.152 | 15.2
[6 empty] | -0.146 | 15.007 | 20.034 | 0.142 | 14.2
[6 empty] | -0.146 | 14.997 | 20.034 | 0.14 | 14
[6 empty] | -0.146 | 14.987 | 20.034 | 0.139 | 13.9
[6 empty] | -0.146 | 14.977 | 20.034 | 0.138 | 13.8
[6 empty] | -0.146 | 14.967 | 20.034 | 0.136 | 13.6
[6 empty] | -0.146 | 14.957 | 20.034 | 0.135 | 13.5
[6 empty] | -0.146 | 14.947 | 20.034 | 0.133 | 13.3
[6 empty] | -0.146 | 14.937 | 20.034 | 0.132 | 13.2
[6 empty] | -0.146 | 14.927 | 20.034 | 0.13 | 13
[6 empty] | -0.146 | 14.917 | 20.034 | 0.129 | 12.9
[6 empty] | -0.146 | 14.907 | 20.034 | 0.127 | 12.7
[6 empty] | -0.146 | 14.897 | 20.034 | 0.126 | 12.6
[6 empty] | -0.146 | 14.887 | 20.034 | 0.125 | 12.5
[6 empty] | -0.146 | 14.877 | 20.034 | 0.123 | 12.3
[6 empty] | -0.146 | 14.867 | 20.034 | 0.122 | 12.2
[6 empty] | -0.146 | 14.857 | 20.034 | 0.12 | 12
[6 empty] | -0.146 | 14.847 | 20.034 | 0.119 | 11.9
[6 empty] | -0.146 | 14.837 | 20.034 | 0.117 | 11.7
[6 empty] | -0.146 | 14.827 | 20.034 | 0.116 | 11.6
[6 empty] | -0.146 | 14.817 | 20.034 | 0.114 | 11.4
[6 empty] | -0.146 | 14.807 | 20.034 | 0.113 | 11.3
[6 empty] | -0.146 | 14.797 | 20.034 | 0.117 | 11.7
[6 empty] | -0.146 | 14.787 | 20.034 | 0.121 | 12.1
[6 empty] | -0.146 | 14.777 | 20.034 | 0.126 | 12.6
[6 empty] | -0.146 | 14.767 | 20.034 | 0.13 | 13
[6 empty] | -0.146 | 14.757 | 20.034 | 0.134 | 13.4
[6 empty] | -0.146 | 14.747 | 20.034 | 0.138 | 13.8
[6 empty] | -0.146 | 14.737 | 20.034 | 0.142 | 14.2
[6 empty] | -0.146 | 14.727 | 20.034 | 0.146 | 14.6
[6 empty] | -0.146 | 14.717 | 20.034 | 0.151 | 15.1
[6 empty] | -0.146 | 14.707 | 20.034 | 0.155 | 15.5
[6 empty] | -0.146 | 14.697 | 20.034 | 0.159 | 15.9
[6 empty] | -0.146 | 14.687 | 20.034 | 0.163 | 16.3
[6 empty] | -0.146 | 14.677 | 20.034 | 0.167 | 16.7
[6 empty] | -0.146 | 14.667 | 20.034 | 0.172 | 17.2
[6 empty] | -0.146 | 14.657 | 20.034 | 0.176 | 17.6
[6 empty] | -0.146 | 14.647 | 20.034 | 0.18 | 18
[6 empty] | -0.146 | 14.637 | 20.034 | 0.184 | 18.4
[6 empty] | -0.146 | 14.627 | 20.034 | 0.188 | 18.8
[6 empty] | -0.146 | 14.617 | 20.034 | 0.193 | 19.3
[6 empty] | -0.146 | 14.607 | 20.034 | 0.197 | 19.7
[6 empty] | -0.146 | 14.597 | 20.034 | 0.205 | 20.5
[6 empty] | -0.146 | 14.587 | 20.034 | 0.214 | 21.4
[6 empty] | -0.146 | 14.577 | 20.034 | 0.222 | 22.2
[6 empty] | -0.146 | 14.567 | 20.034 | 0.23 | 23
[6 empty] | -0.146 | 14.557 | 20.034 | 0.239 | 23.9
[6 empty] | -0.146 | 14.547 | 20.034 | 0.247 | 24.7
[6 empty] | -0.146 | 14.537 | 20.034 | 0.256 | 25.6
[6 empty] | -0.146 | 14.527 | 20.034 | 0.264 | 26.4
[6 empty] | -0.146 | 14.517 | 20.034 | 0.272 | 27.2
[6 empty] | -0.146 | 14.507 | 20.034 | 0.281 | 28.1
[6 empty] | -0.146 | 14.497 | 20.034 | 0.289 | 28.9
[6 empty] | -0.146 | 14.487 | 20.034 | 0.298 | 29.8
[6 empty] | -0.146 | 14.477 | 20.034 | 0.306 | 30.6
[6 empty] | -0.146 | 14.467 | 20.033 | 0.314 | 31.4
[6 empty] | -0.146 | 14.457 | 20.033 | 0.323 | 32.3
[6 empty] | -0.146 | 14.447 | 20.033 | 0.331 | 33.1
[6 empty] | -0.146 | 14.437 | 20.033 | 0.34 | 34
[6 empty] | -0.146 | 14.427 | 20.033 | 0.348 | 34.8
[6 empty] | -0.146 | 14.417 | 20.033 | 0.356 | 35.6
[6 empty] | -0.146 | 14.407 | 20.033 | 0.365 | 36.5
[6 empty] | -0.146 | 14.397 | 20.033 | 0.365 | 36.5
[6 empty] | -0.146 | 14.387 | 20.033 | 0.364 | 36.4
[6 empty] | -0.146 | 14.377 | 20.033 | 0.364 | 36.4
[6 empty] | -0.146 | 14.367 | 20.033 | 0.364 | 36.4
[6 empty] | -0.146 | 14.357 | 20.033 | 0.363 | 36.3
[6 empty] | -0.146 | 14.347 | 20.033 | 0.363 | 36.3
[6 empty] | -0.146 | 14.337 | 20.033 | 0.362 | 36.2
[6 empty] | -0.146 | 14.327 | 20.033 | 0.362 | 36.2
[6 empty] | -0.146 | 14.317 | 20.033 | 0.362 | 36.2
[6 empty] | -0.146 | 14.307 | 20.033 | 0.361 | 36.1
[6 empty] | -0.146 | 14.297 | 20.033 | 0.361 | 36.1
[6 empty] | -0.146 | 14.287 | 20.033 | 0.361 | 36.1
[6 empty] | -0.146 | 14.277 | 20.033 | 0.36 | 36
[6 empty] | -0.146 | 14.267 | 20.033 | 0.36 | 36
[6 empty] | -0.146 | 14.257 | 20.033 | 0.359 | 35.9
[6 empty] | -0.146 | 14.247 | 20.033 | 0.359 | 35.9
[6 empty] | -0.146 | 14.237 | 20.033 | 0.359 | 35.9
[6 empty] | -0.146 | 14.227 | 20.033 | 0.358 | 35.8
[6 empty] | -0.146 | 14.217 | 20.033 | 0.358 | 35.8
[6 empty] | -0.146 | 14.207 | 20.033 | 0.358 | 35.8
[6 empty] | -0.146 | 14.197 | 20.033 | 0.35 | 35
[6 empty] | -0.146 | 14.187 | 20.033 | 0.341 | 34.1
[6 empty] | -0.146 | 14.177 | 20.033 | 0.333 | 33.3
[6 empty] | -0.146 | 14.167 | 20.033 | 0.325 | 32.5
[6 empty] | -0.146 | 14.157 | 20.033 | 0.317 | 31.7
[6 empty] | -0.146 | 14.147 | 20.033 | 0.308 | 30.8
[6 empty] | -0.146 | 14.137 | 20.033 | 0.3 | 30
[6 empty] | -0.146 | 14.127 | 20.033 | 0.292 | 29.2
[6 empty] | -0.146 | 14.117 | 20.033 | 0.283 | 28.3
[6 empty] | -0.146 | 14.107 | 20.033 | 0.275 | 27.5
[6 empty] | -0.146 | 14.097 | 20.033 | 0.267 | 26.7
[6 empty] | -0.146 | 14.087 | 20.033 | 0.259 | 25.9
[6 empty] | -0.146 | 14.077 | 20.033 | 0.25 | 25
[6 empty] | -0.146 | 14.067 | 20.033 | 0.242 | 24.2
[6 empty] | -0.146 | 14.057 | 20.033 | 0.234 | 23.4
[6 empty] | -0.146 | 14.047 | 20.033 | 0.226 | 22.6
[6 empty] | -0.146 | 14.037 | 20.033 | 0.217 | 21.7
[6 empty] | -0.146 | 14.027 | 20.033 | 0.209 | 20.9
[6 empty] | -0.146 | 14.017 | 20.033 | 0.201 | 20.1
[6 empty] | -0.146 | 14.007 | 20.033 | 0.193 | 19.3
[6 empty] | -0.146 | 13.997 | 20.033 | 0.187 | 18.7
[6 empty] | -0.146 | 13.987 | 20.033 | 0.181 | 18.1
[6 empty] | -0.146 | 13.977 | 20.033 | 0.175 | 17.5
[6 empty] | -0.146 | 13.967 | 20.033 | 0.169 | 16.9
[6 empty] | -0.146 | 13.957 | 20.033 | 0.163 | 16.3
[6 empty] | -0.146 | 13.947 | 20.033 | 0.157 | 15.7
[6 empty] | -0.146 | 13.937 | 20.033 | 0.152 | 15.2
[6 empty] | -0.146 | 13.927 | 20.033 | 0.146 | 14.6
[6 empty] | -0.146 | 13.917 | 20.033 | 0.14 | 14
[6 empty] | -0.146 | 13.907 | 20.033 | 0.134 | 13.4
[6 empty] | -0.146 | 13.897 | 20.033 | 0.128 | 12.8
[6 empty] | -0.146 | 13.887 | 20.033 | 0.122 | 12.2
[6 empty] | -0.146 | 13.877 | 20.033 | 0.116 | 11.6
[6 empty] | -0.146 | 13.867 | 20.033 | 0.111 | 11.1
[6 empty] | -0.146 | 13.857 | 20.033 | 0.105 | 10.5
[6 empty] | -0.146 | 13.847 | 20.033 | 0.099 | 9.9
[6 empty] | -0.146 | 13.837 | 20.033 | 0.093 | 9.3
[6 empty] | -0.146 | 13.827 | 20.033 | 0.087 | 8.7
[6 empty] | -0.146 | 13.817 | 20.033 | 0.081 | 8.1
[6 empty] | -0.146 | 13.807 | 20.033 | 0.076 | 7.6
[6 empty] | -0.146 | 13.797 | 20.033 | 0.074 | 7.4
[6 empty] | -0.146 | 13.787 | 20.033 | 0.073 | 7.3
[6 empty] | -0.146 | 13.777 | 20.032 | 0.072 | 7.2
[6 empty] | -0.146 | 13.767 | 20.032 | 0.071 | 7.1
[6 empty] | -0.146 | 13.757 | 20.032 | 0.07 | 7
[6 empty] | -0.146 | 13.747 | 20.032 | 0.069 | 6.9
[6 empty] | -0.146 | 13.737 | 20.032 | 0.067 | 6.7
[6 empty] | -0.146 | 13.727 | 20.032 | 0.066 | 6.6
[6 empty] | -0.146 | 13.717 | 20.032 | 0.065 | 6.5
[6 empty] | -0.146 | 13.707 | 20.032 | 0.064 | 6.4
[6 empty] | -0.146 | 13.697 | 20.032 | 0.063 | 6.3
[6 empty] | -0.146 | 13.687 | 20.032 | 0.062 | 6.2
[6 empty] | -0.146 | 13.677 | 20.032 | 0.061 | 6.1
[6 empty] | -0.146 | 13.667 | 20.032 | 0.06 | 6
[6 empty] | -0.146 | 13.657 | 20.032 | 0.059 | 5.9
[6 empty] | -0.146 | 13.647 | 20.032 | 0.057 | 5.7
[6 empty] | -0.146 | 13.637 | 20.032 | 0.056 | 5.6
[6 empty] | -0.146 | 13.627 | 20.032 | 0.055 | 5.5
[6 empty] | -0.146 | 13.617 | 20.032 | 0.054 | 5.4
[6 empty] | -0.146 | 13.607 | 20.032 | 0.053 | 5.3
[6 empty] | -0.146 | 13.597 | 20.032 | 0.052 | 5.2
[6 empty] | -0.146 | 13.587 | 20.032 | 0.052 | 5.2
[6 empty] | -0.146 | 13.577 | 20.032 | 0.052 | 5.2
[6 empty] | -0.146 | 13.567 | 20.032 | 0.051 | 5.1
[6 empty] | -0.146 | 13.557 | 20.032 | 0.051 | 5.1
[6 empty] | -0.146 | 13.547 | 20.032 | 0.051 | 5.1
[6 empty] | -0.146 | 13.537 | 20.032 | 0.05 | 5
[6 empty] | -0.146 | 13.527 | 20.032 | 0.05 | 5
[6 empty] | -0.146 | 13.517 | 20.032 | 0.049 | 4.9
[6 empty] | -0.146 | 13.507 | 20.032 | 0.049 | 4.9
[6 empty] | -0.146 | 13.497 | 20.032 | 0.049 | 4.9
[6 empty] | -0.146 | 13.487 | 20.032 | 0.048 | 4.8
[6 empty] | -0.146 | 13.477 | 20.032 | 0.048 | 4.8
[6 empty] | -0.146 | 13.467 | 20.032 | 0.048 | 4.8
[6 empty] | -0.146 | 13.457 | 20.032 | 0.047 | 4.7
[6 empty] | -0.146 | 13.447 | 20.032 | 0.047 | 4.7
[6 empty] | -0.146 | 13.437 | 20.032 | 0.046 | 4.6
[6 empty] | -0.146 | 13.427 | 20.032 | 0.046 | 4.6
[6 empty] | -0.146 | 13.417 | 20.032 | 0.046 | 4.6
[6 empty] | -0.146 | 13.407 | 20.032 | 0.045 | 4.5
[6 empty] | -0.146 | 13.397 | 20.032 | 0.045 | 4.5
[6 empty] | -0.146 | 13.387 | 20.032 | 0.045 | 4.5
[6 empty] | -0.146 | 13.377 | 20.032 | 0.045 | 4.5
[6 empty] | -0.146 | 13.367 | 20.032 | 0.045 | 4.5
[6 empty] | -0.146 | 13.357 | 20.032 | 0.044 | 4.4
[6 empty] | -0.146 | 13.347 | 20.032 | 0.044 | 4.4
[6 empty] | -0.146 | 13.337 | 20.032 | 0.044 | 4.4
[6 empty] | -0.146 | 13.327 | 20.032 | 0.044 | 4.4
[6 empty] | -0.146 | 13.317 | 20.032 | 0.044 | 4.4
[6 empty] | -0.146 | 13.307 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.297 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.287 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.277 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.267 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.257 | 20.032 | 0.043 | 4.3
[6 empty] | -0.146 | 13.247 | 20.032 | 0.042 | 4.2
[6 empty] | -0.146 | 13.237 | 20.032 | 0.042 | 4.2
[6 empty] | -0.146 | 13.227 | 20.032 | 0.042 | 4.2
[6 empty] | -0.146 | 13.217 | 20.032 | 0.042 | 4.2
[6 empty] | -0.146 | 13.207 | 20.032 | 0.042 | 4.2
[6 empty] | -0.146 | 13.197 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.187 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.177 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.167 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.157 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.147 | 20.032 | 0.041 | 4.1
[6 empty] | -0.146 | 13.137 | 20.032 | 0.04 | 4
[6 empty] | -0.146 | 13.127 | 20.032 | 0.04 | 4
[6 empty] | -0.146 | 13.117 | 20.032 | 0.04 | 4
[6 empty] | -0.146 | 13.107 | 20.032 | 0.04 | 4
[6 empty] | -0.146 | 13.097 | 20.032 | 0.04 | 4
[6 empty] | -0.146 | 13.087 | 20.031 | 0.04 | 4
[6 empty] | -0.146 | 13.077 | 20.031 | 0.039 | 3.9
[6 empty] | -0.146 | 13.067 | 20.031 | 0.039 | 3.9
[6 empty] | -0.146 | 13.057 | 20.031 | 0.039 | 3.9
[6 empty] | -0.146 | 13.047 | 20.031 | 0.039 | 3.9
[6 empty] | -0.146 | 13.037 | 20.031 | 0.039 | 3.9
[6 empty] | -0.146 | 13.027 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 13.017 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 13.007 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.997 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.987 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.978 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.968 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.958 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.948 | 20.031 | 0.038 | 3.8
[6 empty] | -0.146 | 12.938 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.928 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.918 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.908 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.898 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.888 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.878 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.868 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.858 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.848 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.838 | 20.031 | 0.037 | 3.7
[6 empty] | -0.146 | 12.828 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.818 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.808 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.798 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.788 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.778 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.768 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.758 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.748 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.738 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.728 | 20.031 | 0.036 | 3.6
[6 empty] | -0.146 | 12.718 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.708 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.698 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.688 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.678 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.668 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.658 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.648 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.638 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.628 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.618 | 20.031 | 0.035 | 3.5
[6 empty] | -0.146 | 12.608 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.598 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.588 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.578 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.568 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.558 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.548 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.538 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.528 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.518 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.508 | 20.031 | 0.034 | 3.4
[6 empty] | -0.146 | 12.498 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.488 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.478 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.468 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.458 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.448 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.438 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.428 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.418 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.408 | 20.031 | 0.033 | 3.3
[6 empty] | -0.146 | 12.398 | 20.03 | 0.033 | 3.3
[6 empty] | -0.146 | 12.388 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.378 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.368 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.358 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.348 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.338 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.328 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.318 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.308 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.298 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.288 | 20.03 | 0.032 | 3.2
[6 empty] | -0.146 | 12.278 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.268 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.258 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.248 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.238 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.228 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.218 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.208 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.198 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.188 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.178 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.168 | 20.03 | 0.031 | 3.1
[6 empty] | -0.146 | 12.158 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.148 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.138 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.128 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.118 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.108 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.098 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.088 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.078 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.068 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.058 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.048 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.038 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.028 | 20.03 | 0.03 | 3
[6 empty] | -0.146 | 12.018 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 12.008 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.998 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.988 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.978 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.968 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.958 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.948 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.938 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.928 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.918 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.908 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.898 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.888 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.878 | 20.03 | 0.029 | 2.9
[6 empty] | -0.146 | 11.868 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.858 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.848 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.838 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.828 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.818 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.808 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.798 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.788 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.778 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.768 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.758 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.748 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.738 | 20.03 | 0.028 | 2.8
[6 empty] | -0.146 | 11.728 | 20.03 | 0.027 | 2.7
[6 empty] | -0.146 | 11.718 | 20.03 | 0.027 | 2.7
[6 empty] | -0.146 | 11.708 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.698 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.688 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.678 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.668 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.658 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.648 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.638 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.628 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.618 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.608 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.598 | 20.029 | 0.027 | 2.7
[6 empty] | -0.146 | 11.588 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.578 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.568 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.558 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.548 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.538 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.528 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.518 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.508 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.498 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.488 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.478 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.468 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.458 | 20.029 | 0.026 | 2.6
[6 empty] | -0.146 | 11.448 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.438 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.428 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.418 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.408 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.398 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.388 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.378 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.368 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.358 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.348 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.338 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.328 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.318 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.308 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.298 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.288 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.278 | 20.029 | 0.025 | 2.5
[6 empty] | -0.146 | 11.268 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.258 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.248 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.238 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.228 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.218 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.208 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.198 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.188 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.178 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.168 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.158 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.148 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.138 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.128 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.118 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.108 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.098 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.088 | 20.029 | 0.024 | 2.4
[6 empty] | -0.146 | 11.078 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.068 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.058 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.048 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.038 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.028 | 20.029 | 0.023 | 2.3
[6 empty] | -0.146 | 11.018 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 11.008 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.998 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.988 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.978 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.968 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.958 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.948 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.938 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.928 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.918 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.908 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.898 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.888 | 20.028 | 0.023 | 2.3
[6 empty] | -0.146 | 10.878 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.868 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.858 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.848 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.838 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.828 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.818 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.808 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.798 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.788 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.778 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.768 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.758 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.748 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.738 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.728 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.718 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.708 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.698 | 20.028 | 0.022 | 2.2
[6 empty] | -0.146 | 10.688 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.678 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.668 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.658 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.648 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.638 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.628 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.618 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.608 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.598 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.588 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.578 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.568 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.558 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.548 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.538 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.528 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.518 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.508 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.498 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.488 | 20.028 | 0.021 | 2.1
[6 empty] | -0.146 | 10.478 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.468 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.458 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.448 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.438 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.428 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.418 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.408 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.398 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.388 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.378 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.368 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.358 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.348 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.338 | 20.028 | 0.02 | 2
[6 empty] | -0.146 | 10.328 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.318 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.308 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.298 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.288 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.278 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.268 | 20.027 | 0.02 | 2
[6 empty] | -0.146 | 10.258 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.248 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.238 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.228 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.218 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.208 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.198 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.188 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.178 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.168 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.158 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.148 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.138 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.128 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.118 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.108 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.098 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.088 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.078 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.068 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.058 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.048 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.038 | 20.027 | 0.019 | 1.9
[6 empty] | -0.146 | 10.028 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 10.018 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 10.008 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.998 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.988 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.978 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.969 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.959 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.949 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.939 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.929 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.919 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.909 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.899 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.889 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.879 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.869 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.859 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.849 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.839 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.829 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.819 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.809 | 20.027 | 0.018 | 1.8
[6 empty] | -0.146 | 9.799 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.789 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.779 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.769 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.759 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.749 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.739 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.729 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.719 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.709 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.699 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.689 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.679 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.669 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.659 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.649 | 20.027 | 0.017 | 1.7
[6 empty] | -0.146 | 9.639 | 20.026 | 0.017 | 1.7
[6 empty] | -0.146 | 9.629 | 20.026 | 0.017 | 1.7
[6 empty] | -0.146 | 9.619 | 20.026 | 0.017 | 1.7
[6 empty] | -0.146 | 9.609 | 20.026 | 0.017 | 1.7
[6 empty] | -0.146 | 9.599 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.589 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.579 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.569 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.559 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.549 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.539 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.529 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.519 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.509 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.499 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.489 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.479 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.469 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.459 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.449 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.439 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.429 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.419 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.409 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.399 | 20.026 | 0.016 | 1.6
[6 empty] | -0.146 | 9.389 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.379 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.369 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.359 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.349 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.339 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.329 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.319 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.309 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.299 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.289 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.279 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.269 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.259 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.249 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.239 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.229 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.219 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.209 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.199 | 20.026 | 0.015 | 1.5
[6 empty] | -0.146 | 9.189 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.179 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.169 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.159 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.149 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.139 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.129 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.119 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.109 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.099 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.089 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.079 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.069 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.059 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.049 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.039 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.029 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.019 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 9.009 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 8.999 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 8.989 | 20.026 | 0.014 | 1.4
[6 empty] | -0.146 | 8.979 | 20.026 | 0.013 | 1.3
[6 empty] | -0.146 | 8.969 | 20.026 | 0.013 | 1.3
[6 empty] | -0.146 | 8.959 | 20.026 | 0.013 | 1.3
[6 empty] | -0.146 | 8.949 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.939 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.929 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.919 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.909 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.899 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.889 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.879 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.869 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.859 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.849 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.839 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.829 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.819 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.809 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.799 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.789 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.779 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.769 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.759 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.749 | 20.025 | 0.013 | 1.3
[6 empty] | -0.146 | 8.739 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.729 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.719 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.709 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.699 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.689 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.679 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.669 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.659 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.649 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.639 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.629 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.619 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.609 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.599 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.589 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.579 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.569 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.559 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.549 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.539 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.529 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.519 | 20.025 | 0.012 | 1.2
[6 empty] | -0.146 | 8.509 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.499 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.489 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.479 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.469 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.459 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.449 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.439 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.429 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.419 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.409 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.399 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.389 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.379 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.369 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.359 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.349 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.339 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.329 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.319 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.309 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.299 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.289 | 20.025 | 0.011 | 1.1
[6 empty] | -0.146 | 8.279 | 20.025 | 0.01 | 1
[6 empty] | -0.146 | 8.269 | 20.025 | 0.01 | 1
[6 empty] | -0.146 | 8.259 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.249 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.239 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.229 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.219 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.209 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.199 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.189 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.179 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.169 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.159 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.149 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.139 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.129 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.119 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.109 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.099 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.089 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.079 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.069 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.059 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.049 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.039 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.029 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.019 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 8.009 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.999 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.989 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.979 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.969 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.959 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.949 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.939 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.929 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.919 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.909 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.899 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.889 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.879 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.869 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.859 | 20.024 | 0.01 | 1
[6 empty] | -0.146 | 7.849 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.839 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.829 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.819 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.809 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.799 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.789 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.779 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.769 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.759 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.749 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.739 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.729 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.719 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.709 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.699 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.689 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.679 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.669 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.659 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.649 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.639 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.629 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.619 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.609 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.599 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.589 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.579 | 20.024 | 0.009 | 0.9
[6 empty] | -0.146 | 7.569 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.559 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.549 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.539 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.529 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.519 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.509 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.499 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.489 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.479 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.469 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.459 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.449 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.439 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.429 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.419 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.409 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.399 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.389 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.379 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.369 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.359 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.349 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.339 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.329 | 20.023 | 0.009 | 0.9
[6 empty] | -0.146 | 7.319 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.309 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.299 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.289 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.279 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.269 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.259 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.249 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.239 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.229 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.219 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.209 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.199 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.189 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.179 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.169 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.159 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.149 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.139 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.129 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.119 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.109 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.099 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.089 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.079 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.069 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.059 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.049 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.039 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.029 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.019 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 7.009 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.999 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.989 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.979 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.97 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.96 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.95 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.94 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.93 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.92 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.91 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.9 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.89 | 20.023 | 0.008 | 0.8
[6 empty] | -0.146 | 6.88 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.87 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.86 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.85 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.84 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.83 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.82 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.81 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.8 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.79 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.78 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.77 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.76 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.75 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.74 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.73 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.72 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.71 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.7 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.69 | 20.022 | 0.008 | 0.8
[6 empty] | -0.146 | 6.68 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.67 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.66 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.65 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.64 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.63 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.62 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.61 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.6 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.59 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.58 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.57 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.56 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.55 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.54 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.53 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.52 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.51 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.5 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.49 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.48 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.47 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.46 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.45 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.44 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.43 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.42 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.41 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.4 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.39 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.38 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.37 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.36 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.35 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.34 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.33 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.32 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.31 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.3 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.29 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.28 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.27 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.26 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.25 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.24 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.23 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.22 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.21 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.2 | 20.022 | 0.007 | 0.7
[6 empty] | -0.146 | 6.19 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.18 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.17 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.16 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.15 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.14 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.13 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.12 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.11 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.1 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.09 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.08 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.07 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.06 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.05 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.04 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.03 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.02 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6.01 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 6 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.99 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.98 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.97 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.96 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.95 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.94 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.93 | 20.021 | 0.007 | 0.7
[6 empty] | -0.146 | 5.92 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.91 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.9 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.89 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.88 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.87 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.86 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.85 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.84 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.83 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.82 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.81 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.8 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.79 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.78 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.77 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.76 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.75 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.74 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.73 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.72 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.71 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.7 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.69 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.68 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.67 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.66 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.65 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.64 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.63 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.62 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.61 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.6 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.59 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.58 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.57 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.56 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.55 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.54 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.53 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.52 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.51 | 20.021 | 0.006 | 0.6
[6 empty] | -0.146 | 5.5 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.49 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.48 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.47 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.46 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.45 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.44 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.43 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.42 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.41 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.4 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.39 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.38 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.37 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.36 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.35 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.34 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.33 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.32 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.31 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.3 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.29 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.28 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.27 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.26 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.25 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.24 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.23 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.22 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.21 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.2 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.19 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.18 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.17 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.16 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.15 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.14 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.13 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.12 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.11 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.1 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.09 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.08 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.07 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.06 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.05 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.04 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.03 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.02 | 20.02 | 0.006 | 0.6
[6 empty] | -0.146 | 5.01 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 5 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.99 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.98 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.97 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.96 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.95 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.94 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.93 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.92 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.91 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.9 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.89 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.88 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.87 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.86 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.85 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.84 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.83 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.82 | 20.02 | 0.005 | 0.5
[6 empty] | -0.146 | 4.81 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.8 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.79 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.78 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.77 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.76 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.75 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.74 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.73 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.72 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.71 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.7 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.69 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.68 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.67 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.66 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.65 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.64 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.63 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.62 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.61 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.6 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.59 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.58 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.57 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.56 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.55 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.54 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.53 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.52 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.51 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.5 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.49 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.48 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.47 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.46 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.45 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.44 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.43 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.42 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.41 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.4 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.39 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.38 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.37 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.36 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.35 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.34 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.33 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.32 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.31 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.3 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.29 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.28 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.27 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.26 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.25 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.24 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.23 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.22 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.21 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.2 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.19 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.18 | 20.019 | 0.005 | 0.5
[6 empty] | -0.146 | 4.17 | 20.019 | 0.004 | 0.4
[6 empty] | -0.146 | 4.16 | 20.019 | 0.004 | 0.4
[6 empty] | -0.146 | 4.15 | 20.019 | 0.004 | 0.4
[6 empty] | -0.146 | 4.14 | 20.019 | 0.004 | 0.4
[6 empty] | -0.146 | 4.13 | 20.019 | 0.004 | 0.4
[6 empty] | -0.146 | 4.12 | 20.018 | 0.004 | 0.4
[6 empty] | -0.146 | 4.11 | 20.018 | 0.004 | 0.4
[6 empty] | -0.146 | 4.1 | 20.018 | 0.004 | 0.4
[6 empty] | -0.146 | 4.09 | 20.018 | 0.004 | 0.4
[6 empty] | -0.146 | 4.08 | 20.018 | 0.004 | 0.4
[6 empty] | -0.146 | 4.07 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.06 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.05 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.04 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.03 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.02 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4.01 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 4 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 3.99 | 20.018 | 0.003 | 0.3
[6 empty] | -0.146 | 3.98 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.971 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.961 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.951 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.941 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.931 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.921 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.911 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.901 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.891 | 20.018 | 0.002 | 0.2
[6 empty] | -0.146 | 3.881 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.871 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.861 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.851 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.841 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.831 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.821 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.811 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.801 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.791 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.781 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.771 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.761 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.751 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.741 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.731 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.721 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.711 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.701 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.691 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.681 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.671 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.661 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.651 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.641 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.631 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.621 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.611 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.601 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.591 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.581 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.571 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.561 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.551 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.541 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.531 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.521 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.511 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.501 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.491 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.481 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.471 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.461 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.451 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.441 | 20.018 | 0.001 | 0.1
[6 empty] | -0.146 | 3.431 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.421 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.411 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.401 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.391 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.381 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.371 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.361 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.351 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.341 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.331 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.321 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.311 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.301 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.291 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.281 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.271 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.261 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.251 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.241 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.231 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.221 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.211 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.201 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.191 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.181 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.171 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.161 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.151 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.141 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.131 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.121 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.111 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.101 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.091 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.081 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.071 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.061 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.051 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.041 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.031 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.021 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.011 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 3.001 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.991 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.981 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.971 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.961 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.951 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.941 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.931 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.921 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.911 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.901 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.891 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.881 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.871 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.861 | 20.017 | 0.001 | 0.1
[6 empty] | -0.146 | 2.851 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.841 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.831 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.821 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.811 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.801 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.791 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.781 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.771 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.761 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.751 | 20.017 | 0 | 0
[6 empty] | -0.146 | 2.741 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.731 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.721 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.711 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.701 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.691 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.681 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.671 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.661 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.651 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.641 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.631 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.621 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.611 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.601 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.591 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.581 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.571 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.561 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.551 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.541 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.531 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.521 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.511 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.501 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.491 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.481 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.471 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.461 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.451 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.441 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.431 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.421 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.411 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.401 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.391 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.381 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.371 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.361 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.351 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.341 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.331 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.321 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.311 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.301 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.291 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.281 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.271 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.261 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.251 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.241 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.231 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.221 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.211 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.201 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.191 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.181 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.171 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.161 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.151 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.141 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.131 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.121 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.111 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.101 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.091 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.081 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.071 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.061 | 20.016 | 0 | 0
[6 empty] | -0.146 | 2.051 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.041 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.031 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.021 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.011 | 20.015 | 0 | 0
[6 empty] | -0.146 | 2.001 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.991 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.981 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.971 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.961 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.951 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.941 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.931 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.921 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.911 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.901 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.891 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.881 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.871 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.861 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.851 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.841 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.831 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.821 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.811 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.801 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.791 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.781 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.771 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.761 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.751 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.741 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.731 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.721 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.711 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.701 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.691 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.681 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.671 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.661 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.651 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.641 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.631 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.621 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.611 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.601 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.591 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.581 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.571 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.561 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.551 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.541 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.531 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.521 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.511 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.501 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.491 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.481 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.471 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.461 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.451 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.441 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.431 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.421 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.411 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.401 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.391 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.381 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.371 | 20.015 | 0 | 0
[6 empty] | -0.146 | 1.361 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.351 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.341 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.331 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.321 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.311 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.301 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.291 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.281 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.271 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.261 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.251 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.241 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.231 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.221 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.211 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.201 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.191 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.181 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.171 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.161 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.151 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.141 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.131 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.121 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.111 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.101 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.091 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.081 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.071 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.061 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.051 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.041 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.031 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.021 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.011 | 20.014 | 0 | 0
[6 empty] | -0.146 | 1.001 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.991 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.981 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.972 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.962 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.952 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.942 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.932 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.922 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.912 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.902 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.892 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.882 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.872 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.862 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.852 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.842 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.832 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.822 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.812 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.802 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.792 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.782 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.772 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.762 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.752 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.742 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.732 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.722 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.712 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.702 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.692 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.682 | 20.014 | 0 | 0
[6 empty] | -0.146 | 0.672 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.662 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.652 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.642 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.632 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.622 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.612 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.602 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.592 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.582 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.572 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.562 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.552 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.542 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.532 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.522 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.512 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.502 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.492 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.482 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.472 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.462 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.452 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.442 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.432 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.422 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.412 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.402 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.392 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.382 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.372 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.362 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.352 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.342 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.332 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.322 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.312 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.302 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.292 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.282 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.272 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.262 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.252 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.242 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.232 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.222 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.212 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.202 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.192 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.182 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.172 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.162 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.152 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.142 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.132 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.122 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.112 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.102 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.092 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.082 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.072 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.062 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.052 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.042 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.032 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.022 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.012 | 20.013 | 0 | 0
[6 empty] | -0.146 | 0.002 | 20.013 | 0 | 0
[6 empty] | -0.146 | -0.008 | 20.013 | 0 | 0
[6 empty] | -0.146 | -0.018 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.028 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.038 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.048 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.058 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.068 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.078 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.088 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.098 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.108 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.118 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.128 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.138 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.148 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.158 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.168 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.178 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.188 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.198 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.208 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.218 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.228 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.238 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.248 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.258 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.268 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.278 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.288 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.298 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.308 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.318 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.328 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.338 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.348 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.358 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.368 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.378 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.388 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.398 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.408 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.418 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.428 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.438 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.448 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.458 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.468 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.478 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.488 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.498 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.508 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.518 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.528 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.538 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.548 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.558 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.568 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.578 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.588 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.598 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.608 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.618 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.628 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.638 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.648 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.658 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.668 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.678 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.688 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.698 | 20.012 | 0 | 0
[6 empty] | -0.146 | -0.708 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.718 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.728 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.738 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.748 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.758 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.768 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.778 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.788 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.798 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.808 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.818 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.828 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.838 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.848 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.858 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.868 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.878 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.888 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.898 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.908 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.918 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.928 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.938 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.948 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.958 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.968 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.978 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.988 | 20.011 | 0 | 0
[6 empty] | -0.146 | -0.998 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.008 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.018 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.028 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.038 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.048 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.058 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.068 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.078 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.088 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.098 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.108 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.118 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.128 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.138 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.148 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.158 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.168 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.178 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.188 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.198 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.208 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.218 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.228 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.238 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.248 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.258 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.268 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.278 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.288 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.298 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.308 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.318 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.328 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.338 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.348 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.358 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.368 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.378 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.388 | 20.011 | 0 | 0
[6 empty] | -0.146 | -1.398 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.408 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.418 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.428 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.438 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.448 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.458 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.468 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.478 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.488 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.498 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.508 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.518 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.528 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.538 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.548 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.558 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.568 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.578 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.588 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.598 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.608 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.618 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.628 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.638 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.648 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.658 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.668 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.678 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.688 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.698 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.708 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.718 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.728 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.738 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.748 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.758 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.768 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.778 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.788 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.798 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.808 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.818 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.828 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.838 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.848 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.858 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.868 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.878 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.888 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.898 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.908 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.918 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.928 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.938 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.948 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.958 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.968 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.978 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.988 | 20.01 | 0 | 0
[6 empty] | -0.146 | -1.998 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.008 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.018 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.027 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.037 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.047 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.057 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.067 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.077 | 20.01 | 0 | 0
[6 empty] | -0.146 | -2.087 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.097 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.107 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.117 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.127 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.137 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.147 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.157 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.167 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.177 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.187 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.197 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.207 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.217 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.227 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.237 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.247 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.257 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.267 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.277 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.287 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.297 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.307 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.317 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.327 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.337 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.347 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.357 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.367 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.377 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.387 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.397 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.407 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.417 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.427 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.437 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.447 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.457 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.467 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.477 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.487 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.497 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.507 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.517 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.527 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.537 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.547 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.557 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.567 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.577 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.587 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.597 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.607 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.617 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.627 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.637 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.647 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.657 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.667 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.677 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.687 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.697 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.707 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.717 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.727 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.737 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.747 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.757 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.767 | 20.009 | 0 | 0
[6 empty] | -0.146 | -2.777 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.787 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.797 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.807 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.817 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.827 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.837 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.847 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.857 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.867 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.877 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.887 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.897 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.907 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.917 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.927 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.937 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.947 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.957 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.967 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.977 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.987 | 20.008 | 0 | 0
[6 empty] | -0.146 | -2.997 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.007 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.017 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.027 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.037 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.047 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.057 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.067 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.077 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.087 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.097 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.107 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.117 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.127 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.137 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.147 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.157 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.167 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.177 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.187 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.197 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.207 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.217 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.227 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.237 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.247 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.257 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.267 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.277 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.287 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.297 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.307 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.317 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.327 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.337 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.347 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.357 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.367 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.377 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.387 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.397 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.407 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.417 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.427 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.437 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.447 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.457 | 20.008 | 0 | 0
[6 empty] | -0.146 | -3.467 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.477 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.487 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.497 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.507 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.517 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.527 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.537 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.547 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.557 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.567 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.577 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.587 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.597 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.607 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.617 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.627 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.637 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.647 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.657 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.667 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.677 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.687 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.697 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.707 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.717 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.727 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.737 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.747 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.757 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.767 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.777 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.787 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.797 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.807 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.817 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.827 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.837 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.847 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.857 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.867 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.877 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.887 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.897 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.907 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.917 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.927 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.937 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.947 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.957 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.967 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.977 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.987 | 20.007 | 0 | 0
[6 empty] | -0.146 | -3.997 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.007 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.017 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.027 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.037 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.047 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.057 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.067 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.077 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.087 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.097 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.107 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.117 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.127 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.137 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.147 | 20.007 | 0 | 0
[6 empty] | -0.146 | -4.157 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.167 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.177 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.187 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.197 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.207 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.217 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.227 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.237 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.247 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.257 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.267 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.277 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.287 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.297 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.307 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.317 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.327 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.337 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.347 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.357 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.367 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.377 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.387 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.397 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.407 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.417 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.427 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.437 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.447 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.457 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.467 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.477 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.487 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.497 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.507 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.517 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.527 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.537 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.547 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.557 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.567 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.577 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.587 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.597 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.607 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.617 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.627 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.637 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.647 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.657 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.667 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.677 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.687 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.697 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.707 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.717 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.727 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.737 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.747 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.757 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.767 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.777 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.787 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.797 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.807 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.817 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.827 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.837 | 20.006 | 0 | 0
[6 empty] | -0.146 | -4.847 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.857 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.867 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.877 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.887 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.897 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.907 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.917 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.927 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.937 | 20.005 | 0 | 0
[6 empty] | -0.146 | -4.947 | 20.005 | 0 | 0
[6 empty] | ####################################################### |  |  |  | 0