## Slide 1: RayStation 12A FFF Beam Modelling
Liam Stubbington, RT Physicist

## Slide 2: Why FFF?
FFF beams:
  Higher dose rate
    Shorter fraction/treatment times (for hypofractionated treatment courses)
    Less opportunity for intrafraction motion
  Decrease in treatment head leakage per MU  lower dose outside the field
  Little to no off-axis beam softening  easier to model in TPS
  Decrease in beam quality with Varian machines
  Higher dose per pulse  higher ion recombination
  Reduced output factor variation with field size
Also eventually replace Elements with RayStation:

## Slide 3: Varian vs Elekta
Xiao, Ying, Stephen F. Kry, Richard Popple, Ellen Yorke, Niko Papanikolaou, Sotirios Stathakis, Ping Xia, et al. “Flattening Filter-Free Accelerators: A Report from the AAPM Therapy Emerging Technology Assessment Work Group.” Journal of Applied Clinical Medical Physics 16, no. 3 (2015): 12–29. https://doi.org/10.1120/jacmp.v16i3.5219.
Notes: Dose-rates
  10FFF 2400 MU/min
  6FFF 1400 MU/min

## Slide 4: But which energy?
Why not have beam models for both?
  Main reason is for practice
6FFF for SRS work
  Narrow penumbra
10FFF for SBRT & SABR
  High dose rate

## Slide 5: RayStation 12 A Beam Commissioning module
Copy the template TrueBeam model
  Why would you do anything else?
  Verify against archive commissioning reports & manufacturer data
Copy existing 6MV, 10MV model, confirm merging possible
  1 Varian machine model for RS users
Tick box that says export Beam Dose (330A,0084) as nominal contribution
  Confirm with AH

## Slide 6: RayStation 12 A Beam Commissioning module
Scope of Beam Model
  No wedges
  No cones
  No blocks
  Will routinely be used for IMRT
    Small field profiles and PDDs
  What about largest field size?
    Primary collimator leads to rounded corners
      40x40cm
      We will need to model corners for new machines

## Slide 7: Recommended Beam Data
Square fields
  2cm, 3cm, 5cm, 10cm, 15cm, 20cm, 30cm, 40cm
Rectangular fields
  5×20cm, 20×5cm
Curve types
  PDDs
  Profiles @ 5, 10 & 20cm depth
    Shift curves by half an MLC, interleaf leakage
  One diagonal for the largest field size
Output factors
  Used to normalize the curves, not dose computation so less important
Arbitrary test fields
  I suggest Picket Fence & DLG (synchronous and asynchronous)
These should be centered about CAX and corrected for detector volume/noise
  Use BEAMSCAN

## Slide 8: RayStation 12A Beam Model parameters
Available through auto-modelling scripts
Primary source
  Width
Flatting Filter Source
  Non-zero for FFF
  Effective distance, weight, width
Electron contamination
  Weight, width, weight of flattening filter electron source, E0 and C

## Slide 9: RayStation 12A Beam Model parameters
Available through auto-modelling scripts
Off Axis
  Beam Profile Corrections
  Off axis softening
Photon Energy Spectrum fluence
Output factor corrections
  Not the same as what is measured
  RayStation calculates a Field Measure used to scale MUs with field size
  Not derived from measure OFs

## Slide 10: RayStation 12A Beam Model parameters
Has to be done manually…
MLC
  Effective distance
  Transmission
  Tongue and groove
  Leaf tip width
  Collimator calibration (MLCs, X Jaws, Y Jaws)
Picket Fence fields & dosimetric leaf gap
  SMLC & DMLC

## Slide 11: Measuring the MLC modelling parameters
Essentially 4 parameters characterize the MLCs in RayStation
  Leaf tip width
    Rounded for consistent penumbra
  Tongue-and-groove width
    Reduce inter-leaf leakage
  X-position offset
  Transmission
TPSs require these
  Some rely on matching beam profiles for static MLC defined fields
  Others use 2D Gamma for a set of test beams
Saez, J. et al. (2020) ‘A novel procedure for determining the optimal MLC configuration parameters in treatment planning systems based on measurements with a Farmer chamber’, Physics in Medicine & Biology, 65(15), p. 155006. Available at: https://doi.org/10.1088/1361-6560/ab8cd5.

## Slide 12: Measuring the MLC modelling parameters
Sweeping gap
  DLG or offset
  Sliding slit field with all leaf pairs aligned travels a distance d at constant speed
Asynchornous sweeping gap
  Tongue-and-groove
Saez, J. et al. (2020) ‘A novel procedure for determining the optimal MLC configuration parameters in treatment planning systems based on measurements with a Farmer chamber’, Physics in Medicine & Biology, 65(15), p. 155006. Available at: https://doi.org/10.1088/1361-6560/ab8cd5.
Notes: We can measure these parameters experimentally.

## Slide 13: Measuring the MLC modelling parameters
Saez, J. et al. (2020) ‘A novel procedure for determining the optimal MLC configuration parameters in treatment planning systems based on measurements with a Farmer chamber’, Physics in Medicine & Biology, 65(15), p. 155006. Available at: https://doi.org/10.1088/1361-6560/ab8cd5.

## Slide 14: Measuring the MLC modelling parameters
Obtain average transmission - ratio of open field to MLC closed Use EPID and Farmer
Obtain a series of measurements with sweeping gaps with different gap sizes, plot this against chamber reading and derive the nominal gap and intercept
Obtain a series of readings with asynchronous gaps and use this to derive the and
Compute
Saez, J. et al. (2020) ‘A novel procedure for determining the optimal MLC configuration parameters in treatment planning systems based on measurements with a Farmer chamber’, Physics in Medicine & Biology, 65(15), p. 155006. Available at: https://doi.org/10.1088/1361-6560/ab8cd5.

## Slide 15: Measuring the MLC modelling parameters
Saez, J. et al. (2020) ‘A novel procedure for determining the optimal MLC configuration parameters in treatment planning systems based on measurements with a Farmer chamber’, Physics in Medicine & Biology, 65(15), p. 155006. Available at: https://doi.org/10.1088/1361-6560/ab8cd5.
Chamber reading at the CAX varies linearly with the nominal gap width for SG measurements.
Drop in chamber reading (due to T&G) varies linearly with the leaf offset for aSG, then saturates.
Use this to derive expressions for , and . measured directly.
Curvature and gain handled carefully.

## Slide 16: Sweeping gap fields
Method:
  Modified Varian DLG plans
  Load in Machine QA mode
Measurements:
  A1SL Cheese Phantom
  EPID
    Dosimetry mode

## Slide 17: Sweeping gap fields

## Slide 18: RESULTS: Chamber
To do

## Slide 19: RESULTS: EPID
To do

## Slide 20: RayStation Profile & PDD extraction
Objective:
  Automatic extraction of PDDs and Profiles from RayStation phantom
  Convert to .mcc format
Advantages:
  Manual method is time consuming & limited by manual dexterity
  Make use of Scanned Data Analysis & BEAMSCAN tools

## Slide 21: RayStation Profile & PDD extraction
Run phantom_curves.py script using Script creation in RayStation 12A
Project directory structure
root/
phantom_curves.py
mcc/
template/
PROFILE.mcc
PDD.mcc
raystation_data_out/
ptw_mcc.py
rs_macros/
raystation_curves.py

## Slide 22: RayStation Profile & PDD extraction
Patient Name: RayTraining 40x40x40 Phantom Case: LSt WaterTank Phantom

## Slide 23: RayStation Profile & PDD extraction
Patient Name: RayTraining 40x40x40 Phantom Case: LSt WaterTank Phantom
  3D-CRT Beams
  Add POIs
    GUN, TARGET, RIGHT, LEFT, AXIS
    Depths for profile extraction
  All Beams in current BeamSet written to .mcc
    File name includes:
      Beam Name & Description
      Curve Type (PDD, Profile)
      Depth
Use descriptive beam names!

## Slide 24: RayStation Profile & PDD extraction
Beams.mcc files can be opened in Scanned Data Analysis for review

## Slide 25: RayStation Profile & PDD extraction
Location of POIs is key:
  Crossplane profiles go from RIGHT POI to LEFT POI.
  Inplane profiles go from GUN to TARGET POI.
  Profiles are taken at the depths of all POIs that have "cm" in their label.
  PDDs are taken from the AXIS POI in the ANT POST direction.
Known limitations:
  Bare bones template.mcc files are manipulated Some PTW functionality may not be supported
  Energy – 6X FFF comes in as 6.3 etc.
  Initial Jaw Positions used as Field Size label
  3D-CRT only
  Resulting .mcc will have:
    Collimator and gantry at zero
    Measured in water
    Field type: IRREGULAR
    Reference field: 10x10, 5cm deep, ISOCENTER
    SCD = 1000mm

## Slide 26: RayStation Profile & PDD extraction
Validation
  Two Beams 15x10cm and MLC Striped pattern
  Verified manually in Excel
  See spreadsheet for details.

## Slide 27: RayStation Profile & PDD extraction
Chart: Depth [cm]

## Slide 28: RayStation Profile & PDD extraction
Chart: X [cm]
CROSSPLANE
Chart: Y [cm]

## Slide 29: RayStation Profile & PDD extraction
Chart: Y [cm]

## Slide 30: Thanks for listening!
What have I missed?

## Slide 31: Some useful tools
BeamCommissioningFieldsList
  Only accessibly via scripting
  Allows you to add arbitrary fields to the dose curve computation process
  Very useful for MLC characterization
    Impossible to do with square fields
BEAMSCAN has a button to
  Auto-centre
  Denoise diode measurements, deconvolve chamber volume

## Slide 32: Some useful tools
Varian Golden Beam Data
  Use our own measured data
HomeBrew
  Automatic extraction of profiles, PDDs from water phantom & export as .mcc
    Avoids the need for pixel accuracy with the cursor
    Use the 1D Gamma analysis button in PTW software
  Calculation of 1D Gamma between calculated and measured curves
    PyMedPhys

## Slide 33: Some useful tools
Measurements in other phantoms
  Plans delivered to Cheese with
  Arbitrary fields delivered to:
    Octavius
    EPID
  2D Gamma Analysis comparing different beam models
    Add as Custom Metric in ProKnow