## Notes
References: | Hernandez, V. et al. (2018) ‘A new method for modelling the tongue-and-groove in treatment planning systems’, Physics in Medicine & Biology, 63(24), p. 245005. Available at: https://doi.org/10.1088/1361-6560/aaf098.
LINAC: | LA8
DATE: | 2023-01-02
DICOM Files: | U:\Liam Stubbington\beam-model\sweeping_gap\varian_dcm\
SETUP: | 10×10cm square field. Fields delivered to the EPID in Dosimetry mode.
NOTES:

## ImageJ
MACRO: | C:\Program Files\ImageJ\macros\SweepingGapTools.ijm
DLG
6X DLG
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 6x [MV], 400 [MU/min] | 22500 | 0.01195 | 0.000475 | 0.01186 | 0.01051 | 0.02051 | 595 | 595 | 268.9 | 0.0119 | 100 | 268.9 | 0.2 | 0.026597
4mm : 6x [MV], 400 [MU/min] | 22500 | 0.01915 | 0.0004783 | 0.01909 | 0.01769 | 0.02754 | 595 | 595 | 430.9 | 0.0191 | 100 | 430.9 | 0.4 | 0.042621
6mm : 6x [MV], 400 [MU/min] | 22500 | 0.02653 | 0.0004689 | 0.02642 | 0.02508 | 0.03472 | 595 | 595 | 596.9 | 0.02648 | 100 | 596.9 | 0.6 | 0.059041
10mm: 6x [MV], 400 [MU/min] | 22500 | 0.04139 | 0.0004618 | 0.0413 | 0.03991 | 0.0494 | 595 | 595 | 931.3 | 0.04134 | 100 | 931.3 | 1 | 0.092117
14mm: 6x [MV], 400 [MU/min] | 22500 | 0.05623 | 0.0004588 | 0.05612 | 0.05468 | 0.06446 | 595 | 595 | 1265 | 0.05619 | 100 | 1265 | 1.4 | 0.12512
16mm: 6x [MV], 400 [MU/min] | 22500 | 0.06367 | 0.0004578 | 0.06352 | 0.0621 | 0.07232 | 595 | 595 | 1432 | 0.06363 | 100 | 1432 | 1.6 | 0.14164
20mm: 6x [MV], 400 [MU/min] | 22500 | 0.07848 | 0.0004607 | 0.07845 | 0.07688 | 0.08754 | 595 | 595 | 1766 | 0.07845 | 100 | 1766 | 2 | 0.17468
Open: 6x [MV], 400 [MU/min] | 22500 | 0.4493 | 0.0009185 | 0.4496 | 0.443 | 0.4562 | 595 | 595 | 10110 | 0.4494 | 100 | 10110 | OPEN | 1
TransmA : 6x [MV], 400 [MU/min] | 22500 | 0.004006 | 0.0005898 | 0.003838 | 0.002365 | 0.01269 | 595 | 595 | 90.13 | 0.003945 | 100 | 90.13 | TRANS_A | 0.0089149
TransmB : 6x [MV], 400 [MU/min] | 22500 | 0.003854 | 0.0005357 | 0.003723 | 0.002407 | 0.01217 | 595 | 595 | 86.72 | 0.00379 | 100 | 86.72 | TRANS_B | 0.0085776
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X | 0.8746% | 0.0146 | 0.99125 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X | 0.082391 | 0.0098007 | 0.11895 | 0.013071
6X FFF
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.009998 | 0.0003332 | 0.009869 | 0.009027 | 0.01609 | 595 | 595 | 225 | 0.00997 | 100 | 225 | 0.2 | 0.021469
4mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.01749 | 0.0003549 | 0.01749 | 0.01634 | 0.02361 | 595 | 595 | 393.5 | 0.01747 | 100 | 393.5 | 0.4 | 0.037548
6mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.02515 | 0.0003903 | 0.02505 | 0.02385 | 0.03108 | 595 | 595 | 565.9 | 0.02514 | 100 | 565.9 | 0.6 | 0.053998
10mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.04063 | 0.0004845 | 0.04074 | 0.03917 | 0.04605 | 595 | 595 | 914.1 | 0.04064 | 100 | 914.1 | 1 | 0.087223
14mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.05606 | 0.0005931 | 0.05627 | 0.05422 | 0.06135 | 595 | 595 | 1261 | 0.05611 | 100 | 1261 | 1.4 | 0.12032
16mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.0638 | 0.0006572 | 0.06402 | 0.06183 | 0.06926 | 595 | 595 | 1435 | 0.06386 | 100 | 1435 | 1.6 | 0.13693
20mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.07922 | 0.0007821 | 0.07973 | 0.07689 | 0.08442 | 595 | 595 | 1782 | 0.07931 | 100 | 1782 | 2 | 0.17004
Open: 6xFFF [MV], 400 [MU/min] | 22500 | 0.4658 | 0.004229 | 0.4693 | 0.455 | 0.4763 | 595 | 595 | 10480 | 0.4663 | 100 | 10480 | OPEN | 1
TransmA : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002687 | 0.0003836 | 0.002499 | 0.001622 | 0.009228 | 595 | 595 | 60.45 | 0.002636 | 100 | 60.45 | TRANS_A | 0.0057681
TransmB : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002549 | 0.0003444 | 0.002433 | 0.001589 | 0.008677 | 595 | 595 | 57.35 | 0.002498 | 100 | 57.35 | TRANS_B | 0.0054723
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X FFF | 0.562% |  | 0.99438 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X FFF | 0.082673 | 0.0046117 | 0.055783 | -0.012041
10X FFF
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.01559 | 0.000581 | 0.01549 | 0.01378 | 0.02319 | 595 | 595 | 350.8 | 0.01555 | 100 | 350.8 | 0.2 | 0.029729
4mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.02399 | 0.0006609 | 0.02385 | 0.02198 | 0.03149 | 595 | 595 | 539.8 | 0.02397 | 100 | 539.8 | 0.4 | 0.045746
6mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.03257 | 0.0007504 | 0.03271 | 0.03045 | 0.03971 | 595 | 595 | 732.9 | 0.03257 | 100 | 732.9 | 0.6 | 0.06211
10mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.04989 | 0.0009795 | 0.05017 | 0.04717 | 0.05658 | 595 | 595 | 1123 | 0.04994 | 100 | 1123 | 1 | 0.095169
14mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.06713 | 0.001234 | 0.06785 | 0.06383 | 0.07367 | 595 | 595 | 1510 | 0.06723 | 100 | 1510 | 1.4 | 0.12797
16mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.07576 | 0.001354 | 0.07631 | 0.07212 | 0.08157 | 595 | 595 | 1705 | 0.07589 | 100 | 1705 | 1.6 | 0.14449
20mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.09304 | 0.001618 | 0.09365 | 0.08883 | 0.09843 | 595 | 595 | 2093 | 0.09322 | 100 | 2093 | 2 | 0.17737
Open: 10xFFF [MV], 400 [MU/min] | 22500 | 0.5246 | 0.008596 | 0.5291 | 0.5043 | 0.5474 | 595 | 595 | 11800 | 0.5253 | 100 | 11800 | OPEN | 1
TransmB : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006185 | 0.0005687 | 0.005936 | 0.004581 | 0.01383 | 595 | 595 | 139.2 | 0.006119 | 100 | 139.2 | TRANS_B | 0.011797
TransmA : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006352 | 0.0006326 | 0.006026 | 0.004666 | 0.01447 | 595 | 595 | 142.9 | 0.006269 | 100 | 142.9 | TRANS_A | 0.01211
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 10X FFF | 1.195% |  | 0.98805 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 10X FFF | 0.082144 | 0.013021 | 0.15851 | 0.013336
10X
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 10x [MV], 400 [MU/min] | 22500 | 0.01426 | 0.0005068 | 0.01404 | 0.0126 | 0.0233 | 595 | 595 | 320.8 | 0.01421 | 100 | 320.8 | 0.2 | 0.031952
4mm : 10x [MV], 400 [MU/min] | 22500 | 0.02139 | 0.0005039 | 0.02113 | 0.01969 | 0.03003 | 595 | 595 | 481.3 | 0.02134 | 100 | 481.3 | 0.4 | 0.047938
6mm : 10x [MV], 400 [MU/min] | 22500 | 0.02869 | 0.0005005 | 0.02854 | 0.02699 | 0.0373 | 595 | 595 | 645.6 | 0.02864 | 100 | 645.6 | 0.6 | 0.064303
10mm: 10x [MV], 400 [MU/min] | 22500 | 0.0434 | 0.0004975 | 0.04329 | 0.04166 | 0.05221 | 595 | 595 | 976.4 | 0.04335 | 100 | 976.4 | 1 | 0.097251
14mm: 10x [MV], 400 [MU/min] | 22500 | 0.05807 | 0.0004924 | 0.05808 | 0.05636 | 0.06672 | 595 | 595 | 1307 | 0.05803 | 100 | 1307 | 1.4 | 0.13018
16mm: 10x [MV], 400 [MU/min] | 22500 | 0.06543 | 0.0004862 | 0.06536 | 0.0638 | 0.0742 | 595 | 595 | 1472 | 0.06539 | 100 | 1472 | 1.6 | 0.14661
20mm: 10x [MV], 400 [MU/min] | 22500 | 0.08007 | 0.0004844 | 0.07997 | 0.07839 | 0.0886 | 595 | 595 | 1802 | 0.08003 | 100 | 1802 | 2 | 0.17948
Open: 10x [MV], 400 [MU/min] | 22500 | 0.4464 | 0.000894 | 0.4468 | 0.4406 | 0.4542 | 595 | 595 | 10040 | 0.4465 | 100 | 10040 | OPEN | 1
TransmB : 10x [MV], 400 [MU/min] | 22500 | 0.005598 | 0.0005701 | 0.005509 | 0.003779 | 0.01445 | 595 | 595 | 126 | 0.005538 | 100 | 126 | TRANS_B | 0.01255
TransmA : 10x [MV], 400 [MU/min] | 22500 | 0.005762 | 0.0006279 | 0.005699 | 0.003893 | 0.01531 | 595 | 595 | 129.7 | 0.005698 | 100 | 129.7 | TRANS_A | 0.012918
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 10X | 1.273% |  | 98.73% | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 10X | 0.082093 | 0.015235 | 0.18558 | 0.030804
aSG
6X aSG
Label | Area | Mean | StdDev | Mode | IntDen | Median | %Area | RawIntDen | s [cm] | EPID Signal | Delta
TransmB : 6x [MV], 400 [MU/min] | 22500 | 0.003834 | 0.0005344 | 0.003679 | 86.27 | 0.00377 | 100 | 86.27 | TRANS_B | 0.0085416 | 0.16614
TransmA : 6x [MV], 400 [MU/min] | 22500 | 0.003984 | 0.0005847 | 0.003784 | 89.64 | 0.003921 | 100 | 89.64 | TRANS_A | 0.0088752 | 0.1658
aSG_10mm: 6x [MV], 400 [MU/min] | 22500 | 0.05861 | 0.002857 | 0.05918 | 1319 | 0.05902 | 100 | 1319 | 1 | 0.13059 | 0.044084
aSG_7mm : 6x [MV], 400 [MU/min] | 22500 | 0.06496 | 0.001952 | 0.06539 | 1462 | 0.06527 | 100 | 1462 | 0.7 | 0.14475 | 0.029926
aSG_5mm : 6x [MV], 400 [MU/min] | 22500 | 0.06929 | 0.001355 | 0.06988 | 1559 | 0.06952 | 100 | 1559 | 0.5 | 0.15436 | 0.020322
aSG_3mm : 6x [MV], 400 [MU/min] | 22500 | 0.07349 | 0.0008473 | 0.07373 | 1654 | 0.0736 | 100 | 1654 | 0.3 | 0.16376 | 0.010916
aSG_2mm : 6x [MV], 400 [MU/min] | 22500 | 0.07551 | 0.0006347 | 0.0756 | 1699 | 0.07554 | 100 | 1699 | 0.2 | 0.16822 | 0.0064607
aSG_1mm : 6x [MV], 400 [MU/min] | 22500 | 0.07724 | 0.0004991 | 0.07716 | 1738 | 0.07721 | 100 | 1738 | 0.1 | 0.17208 | 0.0025993
Open: 6x [MV], 400 [MU/min] | 22500 | 0.4488 | 0.0009268 | 0.4491 | 10100 | 0.4489 | 100 | 10100 | OPEN | 1 | -0.82532
EPID Signal
20mm | 0.17468
TRANSMISSION 6X | 0.8708% |  | DLG 6X | 0.013071
| GRADIENT | INTERCEPT | INTERCEPT/GRADIENT | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
0 < s ≤ 1.0 cm | 0.046529 | -0.002662 | -0.057212 | 0.057212 | 0.82207 | 0.0566 | 0.084534 | 0.0016993
6FFF aSG
Label | Area | Mean | StdDev | Mode | IntDen | Median | %Area | RawIntDen | s [cm] | EPID Signal | Delta
TransmB : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002595 | 0.000345 | 0.002553 | 58.39 | 0.002545 | 100 | 58.39 | TRANS_B | 0.0055716 | 0.16447
TransmA : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002726 | 0.0003838 | 0.002555 | 61.34 | 0.002675 | 100 | 61.34 | TRANS_A | 0.0058531 | 0.16419
aSG_10mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.05799 | 0.003272 | 0.05884 | 1305 | 0.05815 | 100 | 1305 | 1 | 0.12452 | 0.045515
aSG_7mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.06457 | 0.002291 | 0.06461 | 1453 | 0.0647 | 100 | 1453 | 0.7 | 0.13865 | 0.031393
aSG_5mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.06909 | 0.001651 | 0.06929 | 1555 | 0.06919 | 100 | 1555 | 0.5 | 0.14838 | 0.02166
aSG_3mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.07358 | 0.001106 | 0.07368 | 1656 | 0.07364 | 100 | 1656 | 0.3 | 0.15802 | 0.012023
aSG_2mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.07571 | 0.0009095 | 0.07589 | 1703 | 0.07576 | 100 | 1703 | 0.2 | 0.1625 | 0.0075382
aSG_1mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.07769 | 0.0008038 | 0.0781 | 1748 | 0.07776 | 100 | 1748 | 0.1 | 0.16679 | 0.0032443
Open: 6xFFF [MV], 400 [MU/min] | 22500 | 0.4658 | 0.004225 | 0.4691 | 10480 | 0.4662 | 100 | 10480 | OPEN | 1 | -0.82996
EPID Signal
20mm | 0.17004
TRANSMISSION 6FFF | 0.5712% |  | DLG 6FFF | 0.012041 | Absolute value taken…
| GRADIENT | INTERCEPT | INTERCEPT/GRADIENT | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
0 < s ≤ 1.0 cm | 0.047299 | -0.001844 | -0.038987 | 0.038987 | 0.85455 | 0.05535 | 0.069835 | 0.0032981
10X aSG
Label | Area | Mean | StdDev | Mode | IntDen | Median | %Area | RawIntDen | s [cm] | EPID Signal | Delta
TransmB : 10x [MV], 400 [MU/min] | 22500 | 0.005564 | 0.0005631 | 0.005344 | 125.2 | 0.0055 | 100 | 125.2 | TRANS_B | 0.01247 | 0.16701
TransmA : 10x [MV], 400 [MU/min] | 22500 | 0.00567 | 0.0006198 | 0.005576 | 127.6 | 0.005599 | 100 | 127.6 | TRANS_A | 0.012709 | 0.16677
aSG_10mm: 10x [MV], 400 [MU/min] | 22500 | 0.06118 | 0.002289 | 0.06237 | 1377 | 0.0614 | 100 | 1377 | 1 | 0.13715 | 0.042331
aSG_7mm : 10x [MV], 400 [MU/min] | 22500 | 0.06734 | 0.001576 | 0.06813 | 1515 | 0.06753 | 100 | 1515 | 0.7 | 0.1509 | 0.028586
aSG_5mm : 10x [MV], 400 [MU/min] | 22500 | 0.07146 | 0.001114 | 0.07172 | 1608 | 0.07159 | 100 | 1608 | 0.5 | 0.16016 | 0.019323
aSG_3mm : 10x [MV], 400 [MU/min] | 22500 | 0.0755 | 0.0007321 | 0.07575 | 1699 | 0.07556 | 100 | 1699 | 0.3 | 0.16922 | 0.010259
aSG_2mm : 10x [MV], 400 [MU/min] | 22500 | 0.07739 | 0.0005886 | 0.07739 | 1741 | 0.07739 | 100 | 1741 | 0.2 | 0.17341 | 0.0060757
aSG_1mm : 10x [MV], 400 [MU/min] | 22500 | 0.07902 | 0.0005026 | 0.07897 | 1778 | 0.07899 | 100 | 1778 | 0.1 | 0.17709 | 0.0023904
Open: 10x [MV], 400 [MU/min] | 22500 | 0.4462 | 0.0009093 | 0.4465 | 10040 | 0.4463 | 100 | 10040 | OPEN | 1 | -0.82052
EPID Signal
20mm | 0.17948
TRANSMISSION 10X | 1.259% |  | DLG 10X | 0.030804
| GRADIENT | INTERCEPT | INTERCEPT/GRADIENT | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
0 < s ≤ 1.0 cm | 0.044798 | -0.0027448 | -0.061272 | 0.061272 | 0.78818 | 0.056837 | 0.099453 | 0.0093083
10FFF aSG
Label | Area | Mean | StdDev | Mode | IntDen | Median | %Area | RawIntDen | s [cm] | EPID Signal | Delta
TransmB : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006232 | 0.0005702 | 0.006069 | 140.2 | 0.006162 | 100 | 140.2 | TRANS_B | 0.011891 | 0.16548
TransmA : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006387 | 0.0006344 | 0.006176 | 143.7 | 0.006303 | 100 | 143.7 | TRANS_A | 0.012188 | 0.16518
aSG_10mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.07086 | 0.00366 | 0.07124 | 1594 | 0.07109 | 100 | 1594 | 1 | 0.1352 | 0.042174
aSG_7mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.07794 | 0.002674 | 0.07844 | 1754 | 0.07812 | 100 | 1754 | 0.7 | 0.14877 | 0.028603
aSG_5mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08274 | 0.002109 | 0.08314 | 1862 | 0.08283 | 100 | 1862 | 0.5 | 0.15793 | 0.019442
aSG_3mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08741 | 0.001708 | 0.0885 | 1967 | 0.08752 | 100 | 1967 | 0.3 | 0.16684 | 0.010537
aSG_2mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08963 | 0.001613 | 0.09073 | 2017 | 0.08979 | 100 | 2017 | 0.2 | 0.17108 | 0.0062957
aSG_1mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.0916 | 0.001589 | 0.09255 | 2061 | 0.09179 | 100 | 2061 | 0.1 | 0.17481 | 0.0025637
Open: 10xFFF [MV], 400 [MU/min] | 22500 | 0.5242 | 0.008617 | 0.5295 | 11790 | 0.5249 | 100 | 11790 | OPEN | 1 | -0.82263
EPID Signal
20mm | 0.17737
TRANSMISSION 10FFF | 1.204% |  | DLG 10FFF | 0.013336
| GRADIENT | INTERCEPT | INTERCEPT/GRADIENT | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
0 < s ≤ 1.0 cm | 0.044371 | -0.0024376 | -0.054935 | 0.054935 | 0.79259 | 0.055983 | 0.09754 | 0.0013097

## ROISize
Use different ROI areas…
aSG 10FFF
SMALL ROI
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | EPID
TransmB : 10xFFF [MV], 400 [MU/min] | 2500 | 0.006011 | 0.0005809 | 0.005929 | 0.004641 | 0.008132 | 595 | 595 | 15.03 | 0.005916 | 100 | 15.03 | 0.011797
TransmA : 10xFFF [MV], 400 [MU/min] | 2500 | 0.006159 | 0.0005845 | 0.005861 | 0.004867 | 0.008029 | 595 | 595 | 15.4 | 0.006072 | 100 | 15.4 | 0.012088
aSG_10mm: 10xFFF [MV], 400 [MU/min] | 2500 | 0.06865 | 0.003582 | 0.07057 | 0.06165 | 0.07597 | 595 | 595 | 171.6 | 0.06881 | 100 | 171.6 | 0.13469
aSG_7mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.0756 | 0.002416 | 0.077 | 0.07069 | 0.08056 | 595 | 595 | 189 | 0.07581 | 100 | 189 | 0.14835
aSG_5mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.08029 | 0.001653 | 0.08044 | 0.07651 | 0.08404 | 595 | 595 | 200.7 | 0.08044 | 100 | 200.7 | 0.15754
aSG_3mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.08487 | 0.0009385 | 0.08514 | 0.0824 | 0.08706 | 595 | 595 | 212.2 | 0.08497 | 100 | 212.2 | 0.16656
aSG_2mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.08705 | 0.0006678 | 0.08711 | 0.08487 | 0.08865 | 595 | 595 | 217.6 | 0.08709 | 100 | 217.6 | 0.1708
aSG_1mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.08899 | 0.0005418 | 0.08894 | 0.08712 | 0.09074 | 595 | 595 | 222.5 | 0.08898 | 100 | 222.5 | 0.17465
Open: 10xFFF [MV], 400 [MU/min] | 2500 | 0.5098 | 0.001945 | 0.5108 | 0.5036 | 0.5165 | 595 | 595 | 1274 | 0.5097 | 100 | 1274 | 1
LARGE ROI
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | EPID | ERROR [%]
TransmB : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006232 | 0.0005702 | 0.006069 | 0.004641 | 0.0139 | 595 | 595 | 140.2 | 0.006162 | 100 | 140.2 | 0.011891 | -0.79%
TransmA : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006387 | 0.0006344 | 0.006176 | 0.00485 | 0.01441 | 595 | 595 | 143.7 | 0.006303 | 100 | 143.7 | 0.012188 | -0.8236%
aSG_10mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.07086 | 0.00366 | 0.07124 | 0.06165 | 0.07925 | 595 | 595 | 1594 | 0.07109 | 100 | 1594 | 0.1352 | -0.3739%
aSG_7mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.07794 | 0.002674 | 0.07844 | 0.07069 | 0.08441 | 595 | 595 | 1754 | 0.07812 | 100 | 1754 | 0.14877 | -0.2813%
aSG_5mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08274 | 0.002109 | 0.08314 | 0.07651 | 0.0881 | 595 | 595 | 1862 | 0.08283 | 100 | 1862 | 0.15793 | -0.2502%
aSG_3mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08741 | 0.001708 | 0.0885 | 0.0824 | 0.09285 | 595 | 595 | 1967 | 0.08752 | 100 | 1967 | 0.16684 | -0.1644%
aSG_2mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.08963 | 0.001613 | 0.09073 | 0.08487 | 0.09533 | 595 | 595 | 2017 | 0.08979 | 100 | 2017 | 0.17108 | -0.1617%
aSG_1mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.0916 | 0.001589 | 0.09255 | 0.08712 | 0.09723 | 595 | 595 | 2061 | 0.09179 | 100 | 2061 | 0.17481 | -0.09289%
Open: 10xFFF [MV], 400 [MU/min] | 22500 | 0.5242 | 0.008617 | 0.5295 | 0.5036 | 0.5471 | 595 | 595 | 11790 | 0.5249 | 100 | 11790 | 1 | 0%
SG 10FFF
SMALL ROI
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | EPID
TransmB : 10xFFF [MV], 400 [MU/min] | 2500 | 0.005952 | 0.0005764 | 0.005681 | 0.004581 | 0.008123 | 595 | 595 | 14.88 | 0.00585 | 100 | 14.88 | 0.011671
TransmA : 10xFFF [MV], 400 [MU/min] | 2500 | 0.006118 | 0.0005847 | 0.006044 | 0.004666 | 0.008042 | 595 | 595 | 15.3 | 0.006034 | 100 | 15.3 | 0.012
2mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.01513 | 0.0004774 | 0.0149 | 0.01378 | 0.01678 | 595 | 595 | 37.83 | 0.01508 | 100 | 37.83 | 0.029671
4mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.02331 | 0.0004867 | 0.02327 | 0.02198 | 0.02502 | 595 | 595 | 58.27 | 0.02327 | 100 | 58.27 | 0.045702
6mm : 10xFFF [MV], 400 [MU/min] | 2500 | 0.03165 | 0.0004768 | 0.03161 | 0.03045 | 0.03326 | 595 | 595 | 79.12 | 0.03161 | 100 | 79.12 | 0.062055
10mm: 10xFFF [MV], 400 [MU/min] | 2500 | 0.04848 | 0.0004953 | 0.04851 | 0.04717 | 0.05024 | 595 | 595 | 121.2 | 0.04845 | 100 | 121.2 | 0.095059
14mm: 10xFFF [MV], 400 [MU/min] | 2500 | 0.06523 | 0.0005225 | 0.06499 | 0.06383 | 0.06702 | 595 | 595 | 163.1 | 0.06521 | 100 | 163.1 | 0.12792
16mm: 10xFFF [MV], 400 [MU/min] | 2500 | 0.07364 | 0.000529 | 0.07367 | 0.07212 | 0.07542 | 595 | 595 | 184.1 | 0.07363 | 100 | 184.1 | 0.14439
20mm: 10xFFF [MV], 400 [MU/min] | 2500 | 0.09044 | 0.0005615 | 0.0904 | 0.08883 | 0.09209 | 595 | 595 | 226.1 | 0.09043 | 100 | 226.1 | 0.17733
Open: 10xFFF [MV], 400 [MU/min] | 2500 | 0.5101 | 0.001951 | 0.5083 | 0.5043 | 0.5167 | 595 | 595 | 1275 | 0.51 | 100 | 1275 | 1
LARGE ROI
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | EPID | ERROR [%]
TransmB : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006185 | 0.0005687 | 0.005936 | 0.004581 | 0.01383 | 595 | 595 | 139.2 | 0.006119 | 100 | 139.2 | 0.011797 | -1.068%
TransmA : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006352 | 0.0006326 | 0.006026 | 0.004666 | 0.01447 | 595 | 595 | 142.9 | 0.006269 | 100 | 142.9 | 0.01211 | -0.9097%
2mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.01559 | 0.000581 | 0.01549 | 0.01378 | 0.02319 | 595 | 595 | 350.8 | 0.01555 | 100 | 350.8 | 0.029729 | -0.1959%
4mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.02399 | 0.0006609 | 0.02385 | 0.02198 | 0.03149 | 595 | 595 | 539.8 | 0.02397 | 100 | 539.8 | 0.045746 | -0.09575%
6mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.03257 | 0.0007504 | 0.03271 | 0.03045 | 0.03971 | 595 | 595 | 732.9 | 0.03257 | 100 | 732.9 | 0.06211 | -0.08898%
10mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.04989 | 0.0009795 | 0.05017 | 0.04717 | 0.05658 | 595 | 595 | 1123 | 0.04994 | 100 | 1123 | 0.095169 | -0.1163%
14mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.06713 | 0.001234 | 0.06785 | 0.06383 | 0.07367 | 595 | 595 | 1510 | 0.06723 | 100 | 1510 | 0.12797 | -0.0348%
16mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.07576 | 0.001354 | 0.07631 | 0.07212 | 0.08157 | 595 | 595 | 1705 | 0.07589 | 100 | 1705 | 0.14449 | -0.06877%
20mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.09304 | 0.001618 | 0.09365 | 0.08883 | 0.09843 | 595 | 595 | 2093 | 0.09322 | 100 | 2093 | 0.17737 | -0.0223%
Open: 10xFFF [MV], 400 [MU/min] | 22500 | 0.5246 | 0.008596 | 0.5291 | 0.5043 | 0.5474 | 595 | 595 | 11800 | 0.5253 | 100 | 11800 | 1 | 0%

## Transmission
LINAC: | LA7
DATE: | 2023-01-16
DETECTOR:
ELECTROMETER: | E13
DICOM Files: | U:\Liam Stubbington\beam-model\sweeping_gap\varian_dcm
SETUP: | 10×10cm square field. Chamber @ +5.0cm depth in blocks. Chamber orientated perpendicular to the direction of MLC travel. For closed MLC fields, plan MU overridden in service to 500MU. Factor 0.2 applied in calculated transmission.
6X
| R1 | R2 | R3 | AVERAGE [nC]
OPEN | 1.66 | 1.66 |  | 1.66
MLC_A | 0.097 | 0.095 |  | 0.096
MLC_B | 0.093 | 0.093 |  | 0.093
TRANSMISSION [%] |  |  |  | 1.139%
6X FFF
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]
10X
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]
10X FFF
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]