## Notes
References: | Hernandez, V. et al. (2018) ‘A new method for modelling the tongue-and-groove in treatment planning systems’, Physics in Medicine & Biology, 63(24), p. 245005. Available at: https://doi.org/10.1088/1361-6560/aaf098.
LINAC: | LA8
DATE: | 2023-01-02
DETECTOR: | A9
ELECTROMETER: | E17
DICOM Files: | U:\Liam Stubbington\beam-model\sweeping_gap\varian_dcm\MODIFIED
SETUP: | 10×10cm square field. Chamber @ +5.0cm depth point in Cheese. Chamber orientated perpendicular to the direction of MLC travel.
NOTES: | For 6. beams, transmission was determined by delivering 500MU of the same field and multiplying the chamber reading by 0.2 For 10. beams, the transmission field was delivered with the electrometer on LOW RANGE. The 6. beams were delivered in service, the 10. beams were delivered in Machine QA mode. With the electrometer on MEDIUM RANGE, the error is taken to be half the smallest increment as chamber readings showed high reproducibilty.

## Analysis
1.) Determine the mean MLC transmission (inter and intra-leaf transmission)
| 6FFF | ± [%] | 10FFF | ± [%] | 6cFF | ± [%] | 10cFF | ± [%]
OPEN FIELD | 1.5093 | 0.027048 | 1.7137 | 0.023823 | 1.5537 | 0.052553 | 1.6483 | 0.065528
MLC_A | 0.0144 | 0.98209 | 0.01955 | 0.27307 | 0.0164 | 3.0488 | 0.02102 | 0.03364
MLC_B | 0.015533 | 0.52564 | 0.020833 | 0.23103 | 0.017733 | 2.3022 | 0.022177 | 0.13275
TRANSMISSION_A | 0.0095406 | 0.98247 | 0.011408 | 0.27411 | 0.010556 | 3.0492 | 0.012752 | 0.073659
TRANSMISSION_B | 0.010292 | 1.1139 | 0.012157 | 0.35769 | 0.011414 | 3.8203 | 0.013454 | 0.13694
TRANSMISSION | 0.9916% | 0.74667 | 1.178% | 0.22729 | 1.098% | 2.4669 | 1.31% | 0.078915
Relative error of transmission ratios are presented…
2.) Determine the gradient and intercept of the chamber reading versus gap width plots.
| GRADIENT [nC/cm] | INTERCEPT [nC] | TRANSMISSION | 1-T | d [cm] × T | INTERCEPT/GRADIENT [cm] | DLG [cm] | k nC/cm2
6FFF | 0.11587 | 0.026388 | 0.0099161 | 0.99008 | 0.11899 | 0.22773 | 0.10755 | 0.46813
| 0.0029415 | 0.0045427
10FFF | 0.13157 | 0.035126 | 0.011783 | 0.98822 | 0.14139 | 0.26698 | 0.1239 | 0.53255
| 0.0034344 | 0.005304
6cFF | 0.11859 | 0.028663 | 0.010985 | 0.98902 | 0.13182 | 0.2417 | 0.10842 | 0.47962
| 0.003887 | 0.0065505
10cFF | 0.12503 | 0.036453 | 0.013103 | 0.9869 | 0.15724 | 0.29156 | 0.13224 | 0.50675
| 0.0041203 | 0.0069924
Goodness of fit needs evaluating….
Distance travelled, d, by the sliding gap, is taken to be 12cm.
The leaf-tip-width value, for the central leaves of a TrueBeamSTx, is taken to be 0.25cm.
3.) Derive the leaf tip width, tongue and groove width and offset from the aSG measurements.
s [cm] | 6FFF [nC] | DELTA / k [cm2] | 10FFF [nC] | DELTA / k [cm2] | 6cFF [nC] | DELTA / k [cm2] | 10cFF [nC] | DELTA / k [cm2]
0 | 0.268 | #N/A | 0.309 | #N/A | 0.278 | #N/A | 0.299 | #N/A
0.1 | 0.266 | 0.0042723 | 0.307 | 0.0037555 | #N/A | #N/A | #N/A | #N/A
0.2 | 0.259 | 0.019225 | 0.3 | 0.0169 | 0.27133 | 0.0139 | 0.29 | 0.009
0.3 | 0.253 | 0.032042 | 0.293 | 0.030044 | #N/A | #N/A | #N/A | #N/A
0.5 | 0.24 | 0.059812 | 0.279 | 0.056333 | 0.251 | 0.056294 | 0.269 | 0.03
0.7 | 0.227 | 0.087582 | 0.264 | 0.0845 | #N/A | #N/A | #N/A | #N/A
1 | 0.208 | 0.12817 | 0.243 | 0.12393 | 0.217 | 0.12718 | 0.233 | 0.066
2 | 0.2005 | 0.14419 | 0.235 | 0.13896 | 0.209 | 0.14386 | 0.225 | 0.074
3 | 0.2 | 0.14526 | #N/A | #N/A | 0.209 | 0.14386 | #N/A | #N/A
|  |  |  | s [cm] | 6cFF [nC] | DELTA / k [cm2] | 10cFF [nC] | DELTA / k [cm2]
|  |  |  | 0.2 | 0.27133 | 0.0139 | 0.29 | 0.009
|  |  |  | 0.5 | 0.251 | 0.056294 | 0.269 | 0.03
|  |  |  | 1 | 0.217 | 0.12718 | 0.233 | 0.066
Take the gradient and intercept for 0 < s < 1cm….
| GRADIENT [cm] | INTERCEPT [cm2] | INTERCEPT/GRADIENT [cm] | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
6FFF | 0.13736 | -0.0089172 | -0.064919 | 0.064919 | 0.81076 | 0.084711 | 0.090561 | 0.10167
| 0.00066302 | 0.00037113
10FFF | 0.13395 | -0.0099347 | -0.074165 | 0.074165 | 0.79469 | 0.084282 | 0.097919 | 0.11664
| 0.00068831 | 0.00038529
6cFF | 0.14162 | -0.01446 | -0.1021 | 0.1021 | 0.80137 | 0.088363 | 0.094866 | 0.098736
| 0.00012283 | 8.0547e-05
10cFF | 0.071327 | -0.0054184 | -0.075966 | 0.075966 | 0.78417 | 0.045479 | 0.10271 | 0.12443
| 0.00053022 | 0.00034769

## Analysis_2
1.) Determine the mean MLC transmission (inter and intra-leaf transmission)
| 6FFF | 10FFF | 6cFF | 10cFF
OPEN FIELD | 1.5093 | 1.7137 | 1.5537 | 1.6483
TRANSMISSION | 0.9916% | 1.178% | 1.098% | 1.31%
Relative error of transmission ratios are presented…
2.) Determine the gradient and intercept of the chamber reading versus gap width plots.
| GRADIENT [nC/cm] | INTERCEPT [nC] | TRANSMISSION | 1-TRANSMISSION | d [CM] | d [cm] × T | INTERCEPT/GRADIENT [cm] | DLG [cm]
6FFF | 0.082146 | 0.012996 | 0.9916% | 0.99008 | 12 | 0.11899 | 0.15821 | 0.038023
10FFF | 0.082295 | 0.01589 | 1.178% | 0.98822 | 12 | 0.14139 | 0.19309 | 0.050012
6cFF | 0.082141 | 0.014574 | 1.098% | 0.98902 | 12 | 0.13182 | 0.17743 | 0.044144
10cFF | 0.081784 | 0.017774 | 1.31% | 0.9869 | 12 | 0.15724 | 0.21733 | 0.058008
Distance travelled, d, by the sliding gap, is taken to be 12cm.
3.) Derive the leaf tip width, tongue and groove width and offset from the aSG measurements.
Take the gradient and intercept for 0 < s ≤ 1cm….
| GRADIENT [cm] | INTERCEPT [cm2] | INTERCEPT/GRADIENT [cm] | ltip [cm] | 1 - 2SQRT(T) + T | wTG [cm] | SQRT(T) - T / (1 - T) | C [cm]
6FFF | 0.042603 | -0.0027657 | -0.064919 | 0.064919 | 0.81076 | 0.052547 | 0.090561 | 0.013132
10FFF | 0.041628 | -0.0030874 | -0.074165 | 0.074165 | 0.79469 | 0.052383 | 0.097919 | 0.017744
6cFF |  |  | #DIV/0! | #DIV/0! | 0.80137 | 0 | 0.094866
10cFF |  |  | #DIV/0! | #DIV/0! | 0.78417 | 0 | 0.10271
No factor 2 included in wTG

## ImageJ
MACRO: | C:\Program Files\ImageJ\macros\SweepingGapTools.ijm
DLG
6X DLG
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 6x [MV]; 400 [MU/min] | 22500 | 0.01315 | 0.000451 | 0.01298 | 0.01197 | 0.01503 | 595 | 595 | 295.9 | 0.01311 | 100 | 295.9 | 0.2 | 0.0287
4mm : 6x [MV]; 400 [MU/min] | 22500 | 0.02045 | 0.0004475 | 0.02028 | 0.01926 | 0.02239 | 595 | 595 | 460.1 | 0.02041 | 100 | 460.1 | 0.4 | 0.044627
6mm : 6x [MV]; 400 [MU/min] | 22500 | 0.02798 | 0.0004433 | 0.02782 | 0.02667 | 0.02992 | 595 | 595 | 629.6 | 0.02794 | 100 | 629.6 | 0.6 | 0.061067
10mm: 6x [MV]; 400 [MU/min] | 22500 | 0.04314 | 0.0004361 | 0.04297 | 0.04178 | 0.045 | 595 | 595 | 970.6 | 0.0431 | 100 | 970.6 | 1 | 0.094142
14mm: 6x [MV]; 400 [MU/min] | 22500 | 0.05825 | 0.0004392 | 0.05818 | 0.05671 | 0.06045 | 595 | 595 | 1311 | 0.05822 | 100 | 1311 | 1.4 | 0.12716
16mm: 6x [MV]; 400 [MU/min] | 22500 | 0.06584 | 0.0004359 | 0.06573 | 0.06435 | 0.06782 | 595 | 595 | 1481 | 0.0658 | 100 | 1481 | 1.6 | 0.14365
20mm: 6x [MV]; 400 [MU/min] | 22500 | 0.08098 | 0.0004381 | 0.08095 | 0.07942 | 0.08302 | 595 | 595 | 1822 | 0.08095 | 100 | 1822 | 2 | 0.17672
TransmA : 6x [MV]; 400 [MU/min] | 22500 | 0.004411 | 0.000503 | 0.0041 | 0.00337 | 0.006314 | 595 | 595 | 99.25 | 0.0043 | 100 | 99.25 | TRANS_A | 0.0096266
TransmB : 6x [MV]; 400 [MU/min] | 22500 | 0.004313 | 0.0004408 | 0.004044 | 0.00334 | 0.006474 | 595 | 595 | 97.05 | 0.004239 | 100 | 97.05 | TRANS_B | 0.0094132
Open: 6x [MV]; 400 [MU/min] | 22500 | 0.4582 | 0.0009256 | 0.4585 | 0.4532 | 0.4628 | 595 | 595 | 10310 | 0.4583 | 100 | 10310 | OPEN | 1
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X | 0.952% | 0.0146 | 0.99048 | 10
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X | 0.082375 | 0.011852 | 0.14388 | 0.047763
6X FFF DLG
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal | Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen
TransmB : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.002249 | 0.0003162 | 0.002053 | 0.001442 | 0.004162 | 595 | 595 | 50.61 | 0.002207 | 100 | 50.61 | TRANS_A | 0.0047255 | TransmB : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.002249 | 0.0003162 | 0.002053 | 0.001442 | 0.004162 | 595 | 595 | 50.61 | 0.002207 | 100 | 50.61
TransmA : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.002324 | 0.0003547 | 0.002188 | 0.001524 | 0.003747 | 595 | 595 | 52.3 | 0.002256 | 100 | 52.3 | TRANS_B | 0.0048833 | TransmA : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.002324 | 0.0003547 | 0.002188 | 0.001524 | 0.003747 | 595 | 595 | 52.3 | 0.002256 | 100 | 52.3
Open: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.4762 | 0.004312 | 0.4791 | 0.4632 | 0.4873 | 595 | 595 | 10710 | 0.4767 | 100 | 10710 | OPEN | 1 | Open: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.4762 | 0.004312 | 0.4791 | 0.4632 | 0.4873 | 595 | 595 | 10710 | 0.4767 | 100 | 10710
6mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.02588 | 0.00042 | 0.02584 | 0.02467 | 0.0276 | 595 | 595 | 582.4 | 0.02586 | 100 | 582.4 | 0.6 | 0.054379 | 6mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.02588 | 0.00042 | 0.02584 | 0.02467 | 0.0276 | 595 | 595 | 582.4 | 0.02586 | 100 | 582.4
4mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.01802 | 0.0003873 | 0.01794 | 0.01691 | 0.01978 | 595 | 595 | 405.5 | 0.018 | 100 | 405.5 | 0.4 | 0.037862 | 4mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.01802 | 0.0003873 | 0.01794 | 0.01691 | 0.01978 | 595 | 595 | 405.5 | 0.018 | 100 | 405.5
2mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.0104 | 0.0003594 | 0.01039 | 0.009421 | 0.01222 | 595 | 595 | 234 | 0.01038 | 100 | 234 | 0.2 | 0.021849 | 2mm : 6xFFF [MV]; 400 [MU/min] | 22500 | 0.0104 | 0.0003594 | 0.01039 | 0.009421 | 0.01222 | 595 | 595 | 234 | 0.01038 | 100 | 234
20mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.08119 | 0.0008218 | 0.08164 | 0.07883 | 0.08395 | 595 | 595 | 1827 | 0.08127 | 100 | 1827 | 2 | 0.17059 | 20mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.08119 | 0.0008218 | 0.08164 | 0.07883 | 0.08395 | 595 | 595 | 1827 | 0.08127 | 100 | 1827
16mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.06537 | 0.0006952 | 0.0657 | 0.06333 | 0.06764 | 595 | 595 | 1471 | 0.06542 | 100 | 1471 | 1.6 | 0.13735 | 16mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.06537 | 0.0006952 | 0.0657 | 0.06333 | 0.06764 | 595 | 595 | 1471 | 0.06542 | 100 | 1471
14mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.05747 | 0.0006324 | 0.05753 | 0.05559 | 0.05974 | 595 | 595 | 1293 | 0.0575 | 100 | 1293 | 1.4 | 0.12073 | 14mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.05747 | 0.0006324 | 0.05753 | 0.05559 | 0.05974 | 595 | 595 | 1293 | 0.0575 | 100 | 1293
10mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.04166 | 0.0005173 | 0.04158 | 0.04013 | 0.04355 | 595 | 595 | 937.4 | 0.04166 | 100 | 937.4 | 1 | 0.087526 | 10mm: 6xFFF [MV]; 400 [MU/min] | 22500 | 0.04166 | 0.0005173 | 0.04158 | 0.04013 | 0.04355 | 595 | 595 | 937.4 | 0.04166 | 100 | 937.4
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X FFF | 0.4804% |  | 0.9952 | 10
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X FFF | 0.082763 | 0.004912 | 0.05935 | 0.011074
10X FFF DLG
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
10mm: 10xFFF [MV]; 400 [MU/min] | 22500 | 0.04591 | 0.0008534 | 0.04622 | 0.04364 | 0.04871 | 595 | 595 | 1033 | 0.04593 | 100 | 1033 | 1 | 0.093231
14mm: 10xFFF [MV]; 400 [MU/min] | 22500 | 0.06217 | 0.001083 | 0.06266 | 0.05934 | 0.06543 | 595 | 595 | 1399 | 0.06224 | 100 | 1399 | 1.4 | 0.12626
16mm: 10xFFF [MV]; 400 [MU/min] | 22500 | 0.07033 | 0.001208 | 0.07122 | 0.06718 | 0.07396 | 595 | 595 | 1583 | 0.07041 | 100 | 1583 | 1.6 | 0.14287
20mm: 10xFFF [MV]; 400 [MU/min] | 22500 | 0.08659 | 0.001459 | 0.08716 | 0.08286 | 0.09075 | 595 | 595 | 1948 | 0.08671 | 100 | 1948 | 2 | 0.17581
2mm : 10xFFF [MV]; 400 [MU/min] | 22500 | 0.01374 | 0.0004667 | 0.01367 | 0.01244 | 0.01589 | 595 | 595 | 309.1 | 0.0137 | 100 | 309.1 | 0.2 | 0.027897
4mm : 10xFFF [MV]; 400 [MU/min] | 22500 | 0.02156 | 0.0005413 | 0.02163 | 0.02004 | 0.02372 | 595 | 595 | 485.2 | 0.02154 | 100 | 485.2 | 0.4 | 0.043791
6mm : 10xFFF [MV]; 400 [MU/min] | 22500 | 0.02964 | 0.0006302 | 0.02985 | 0.02788 | 0.03203 | 595 | 595 | 667 | 0.02964 | 100 | 667 | 0.6 | 0.060199
Open: 10xFFF [MV]; 400 [MU/min] | 22500 | 0.4923 | 0.007907 | 0.498 | 0.472 | 0.5131 | 595 | 595 | 11080 | 0.493 | 100 | 11080 | OPEN | 1
TransmA : 10xFFF [MV]; 400 [MU/min] | 22500 | 0.004273 | 0.0004822 | 0.004016 | 0.00318 | 0.006173 | 595 | 595 | 96.15 | 0.004181 | 100 | 96.15 | TRANS_A | 0.0086778
TransmB : 10xFFF [MV]; 400 [MU/min] | 22500 | 0.004252 | 0.0003949 | 0.004023 | 0.003265 | 0.007267 | 595 | 595 | 95.66 | 0.004197 | 100 | 95.66 | TRANS_B | 0.0086336
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 10X FFF | 0.8656% |  | 0.99134 | 10
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 10X FFF | 0.082349 | 0.011022 | 0.13384 | 0.046532

## 6FFF_measurements
0.) Open field chamber readings
No MLCs…
# | R [nC]
1 | 1.51
2 | 1.509
3 | 1.509
AVERAGE [nC] | 1.5093
STD ERR. [nC] | 0.00040825
1.) MLC Leaf bank mean transmission measurements
Leafs closed behind jaws…
500MUs
| Leaf Bank A | Leaf Bank B
# | R [nC] | R [nC]
1 | 0.072 | 0.078
2 | 0.073 | 0.077
3 | 0.071 | 0.078
AVERAGE [nC] | 0.0144 | 0.015533
STD ERR. [nC] | 0.00014142 | 8.165e-05
TRANSMISSION | 0.9541% | 1.029% | 0.9916%
2.) Synchronous sweeping gap measurements
Constant gap width, MLCs move at constant speed.
d [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO | RESIDUAL | RESIDUAL 2
0.2 | 0.045 | 0.045 |  | 0.045 | 0.0005 | 0.029814 | -0.0030226 | 0.00038932 | 0.029425
0.4 | 0.069 | 0.069 |  | 0.069 | 0.0005 | 0.045716 | -0.0024757 | -0.00013874 | 0.045854
0.6 | 0.094 | 0.094 |  | 0.094 | 0.0005 | 0.062279 | -0.0012662 | -4.2471e-06 | 0.062283
1 | 0.143 | 0.143 |  | 0.143 | 0.0005 | 0.094744 | 0.00049023 | -0.00039781 | 0.095142
1.4 | 0.193 | 0.193 |  | 0.193 | 0.0005 | 0.12787 | 0.0029092 | -0.00012883 | 0.128
1.6 | 0.218 | 0.218 |  | 0.218 | 0.0005 | 0.14443 | 0.0041187 | 5.6628e-06 | 0.14443
2 | 0.268 | 0.268 |  | 0.268 | 0.0005 | 0.17756 | 0.0065376 | 0.00027464 | 0.17729
3 | 0.363 | 0.363 |  | 0.363 | 0.0005 | 0.2405 | -0.0072913 | -0.018929 | 0.25943
|  |  |  | GRADIENT | INTERCEPT
|  |  | 0 < d ≤ 3.0 cm | 0.076771 | 0.017483
|  |  | 0 < d ≤ 2.0 cm | 0.082146 | 0.012996
On this basis we cannot use the reading for d = 3.0cm
3.) Asynchrnous sweeping gap measurements
Adjacent leaves have different positions, constant gap width 20mm, MLCs move at constant speed.
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | RATIO | DELTA | CALCULATED | RESIDUAL
0 | 0.268 | 0.268 |  | 0.268 | 0.17756 | 0 | -0.0027657 | 0.0027657
0.1 | 0.266 | 0.266 |  | 0.266 | 0.17624 | 0.0013251 | 0.0014946 | -0.00016949
0.2 | 0.259 | 0.259 |  | 0.259 | 0.1716 | 0.0059629 | 0.0057549 | 0.00020801
0.3 | 0.253 | 0.253 |  | 0.253 | 0.16762 | 0.0099382 | 0.010015 | -7.704e-05
0.5 | 0.24 | 0.24 |  | 0.24 | 0.15901 | 0.018551 | 0.018536 | 1.5408e-05
0.7 | 0.227 | 0.227 |  | 0.227 | 0.1504 | 0.027164 | 0.027056 | 0.00010786
1 | 0.208 | 0.208 |  | 0.208 | 0.13781 | 0.039753 | 0.039837 | -8.4744e-05
2 | 0.2 | 0.201 |  | 0.2005 | 0.13284 | 0.044722 | 0.082441 | -0.037719
3 | 0.2 |  |  | 0.2 | 0.13251 | 0.045053 | 0.12504 | -0.079991
|  |  |  | GRADIENT | INTERCEPT
|  |  | 0 < s ≤ 1.0 cm | 0.042603 | -0.0027657
|  |  |  | 0.042368 | -0.0025804
4.) Asynchrnous sweeping gap measurements
Different gap widths, with different s values.
This is to show that the reduction in fluence at the chamber is independent of the leaf gap used.
d = 0.5cm
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC]
0.2
0.5
1
2
3
d = 3.0cm
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC]
0.2
0.5
1
2
3

## 10FFF_measurements
0.) Open field chamber readings
No MLCs…
# | R [nC]
1 | 1.714
2 | 1.714
3 | 1.713
AVERAGE [nC] | 1.7137
STD ERR. [nC] | 0.00040825
1.) MLC Leaf bank mean transmission measurements
Leafs closed behind jaws…
Low RANGE note pC not nC
| Leaf Bank A | Leaf Bank B
# | R [pC] | R [pC]
1 | 19.48 | 20.81
2 | 19.54 | 20.91
3 | 19.63 | 20.78
AVERAGE [nC] | 0.01955 | 0.020833
STD ERR. [nC] | 5.3385e-05 | 4.8132e-05
TRANSMISSION | 1.141% | 1.216% | 1.178%
2.) Synchronous sweeping gap measurements
Constant gap width, MLCs move at constant speed.
d [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO | RESIDUAL | RESIDUAL 2
0.2 | 0.056 | 0.056 |  | 0.056 |  | 0.032678 | -0.0031742 | 0.00032918 | 0.032349
0.4 | 0.083 | 0.083 |  | 0.083 |  | 0.048434 | -0.0027736 | -0.00037407 | 0.048808
0.6 | 0.112 | 0.112 |  | 0.112 |  | 0.065357 | -0.001206 | 8.9776e-05 | 0.065267
1 | 0.168 | 0.168 |  | 0.168 |  | 0.098035 | 0.00076221 | -0.00014963 | 0.098185
1.4 | 0.224 | 0.224 |  | 0.224 |  | 0.13071 | 0.0027304 | -0.00038903 | 0.1311
1.6 | 0.254 | 0.254 |  | 0.254 |  | 0.14822 | 0.0048816 | 0.00065836 | 0.14756
2 | 0.309 | 0.309 |  | 0.309 |  | 0.18032 | 0.0062663 | -0.00016459 | 0.18048
3 | 0.417 | 0.417 |  | 0.417 |  | 0.24334 | -0.0074867 | -0.019437 | 0.26277
|  |  |  | GRADIENT | INTERCEPT
|  |  | 0 < d ≤ 3.0 cm | 0.076776 | 0.020498
|  |  | 0 < d ≤ 2.0 cm | 0.082295 | 0.01589
3.) Asynchrnous sweeping gap measurements
Adjacent leaves have different positions, constant gap width 20mm, MLCs move at constant speed.
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | RATIO | DELTA | CALCULATED | RESIDUAL
0 | 0.309 | 0.309 |  | 0.309 | 0.18032 | 0 | -0.0030874 | 0.0030874
0.1 | 0.307 | 0.307 |  | 0.307 | 0.17915 | 0.0011671 | 0.0010755 | 9.1603e-05
0.2 | 0.3 | 0.3 |  | 0.3 | 0.17506 | 0.0052519 | 0.0052383 | 1.3571e-05
0.3 | 0.293 | 0.293 |  | 0.293 | 0.17098 | 0.0093367 | 0.0094012 | -6.4461e-05
0.5 | 0.279 | 0.279 |  | 0.279 | 0.16281 | 0.017506 | 0.017727 | -0.00022053
0.7 | 0.264 | 0.264 |  | 0.264 | 0.15406 | 0.026259 | 0.026053 | 0.00020695
1 | 0.243 | 0.243 |  | 0.243 | 0.1418 | 0.038514 | 0.038541 | -2.7142e-05
2 | 0.235 | 0.235 |  | 0.235 | 0.13713 | 0.043182 | 0.080169 | -0.036987
|  |  |  | GRADIENT | INTERCEPT
|  |  | 0 < s ≤ 1.0 cm | 0.041628 | -0.0030874

## 6cFF_measurements
0.) Open field chamber readings
No MLCs…
# | R [nC]
1 | 1.553
2 | 1.555
3 | 1.553
AVERAGE [nC] | 1.5537
STD ERR. [nC] | 0.0008165
1.) MLC Leaf bank mean transmission measurements
Leafs closed behind jaws…
500MU - factor 5 included in the average
| Leaf Bank A | Leaf Bank B
# | R [nC] | R [nC]
1 | 0.082 | 0.089
2 | 0.082 | 0.089
3 | 0.082 | 0.088
AVERAGE [nC] | 0.0164 | 0.017733
STD ERR. [nC] | 0.0005 | 0.00040825
TRANSMISSION | 1.056% | 1.141% | 1.098%
2.) Synchronous sweeping gap measurements
Constant gap width, MLCs move at constant speed.
d [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO | RESIDUAL | RESIDUAL 2
0.2 | 0.049 | 0.048 | 0.048 | 0.048333 | 0.00013608 | 0.031109 | -0.0026052 | 0.00010727
0.4 | 0.073 | 0.074 | 0.074 | 0.073667 | 0.00013608 | 0.047415 | -0.0015653 | -1.5325e-05
1 | 0.15 | 0.15 | 0.15 | 0.15 | 0 | 0.096546 | 0.0017689 | -0.00016857
2 | 0.278 | 0.278 | 0.278 | 0.278 | 0 | 0.17893 | 0.0078266 | 7.6624e-05
3 | 0.376 | 0.376 | 0.376 | 0.376 | 3.9252e-17 | 0.24201 | -0.005425 | -0.018987
|  |  |  | GRADIENT | INTERCEPT
|  |  | 0 < d ≤ 3.0 cm | 0.076328 | 0.018449
|  |  | 0 < d ≤ 2.0 cm | 0.082141 | 0.014574
3.) Asynchrnous sweeping gap measurements
Adjacent leaves have different positions, constant gap width 20mm, MLCs move at constant speed.
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO
0.2 | 0.271 | 0.272 | 0.271 | 0.27133 | 0.00036004 | 0.17464
0.5 | 0.251 | 0.251 | 0.251 | 0.251 | 0 | 0.16155
1 | 0.217 | 0.217 | 0.217 | 0.217 | 0 | 0.13967
2 | 0.209 | 0.209 | 0.209 | 0.209 | 0 | 0.13452
3 | 0.209 | 0.209 | 0.209 | 0.209 | 0 | 0.13452

## 10cFF_measurements
0.) Open field chamber readings
No MLCs…
# | R [nC]
1 | 1.65
2 | 1.648
3 | 1.647
AVERAGE [nC] | 1.6483
STD ERR. [nC] | 0.0010801
1.) MLC Leaf bank mean transmission measurements
Leafs closed behind jaws…
pC
| Leaf Bank A | Leaf Bank B
# | R [pC] | R [pC]
1 | 21.02 | 22.13
2 | 21.03 | 22.19
3 | 21.01 | 22.21
AVERAGE [nC] | 0.02102 | 0.022177
STD ERR. [nC] | 7.0711e-06 | 2.9439e-05
TRANSMISSION | 1.275% | 1.345% | 1.31%
2.) Synchronous sweeping gap measurements
Constant gap width, MLCs move at constant speed.
d [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO
0.2 | 0.05646 |  |  | 0.05646 | #DIV/0! | 0.034253
0.6 | 0.11 |  |  | 0.11 | #DIV/0! | 0.066734
1 | 0.164 |  |  | 0.164 | #DIV/0! | 0.099494
2 | 0.299 |  |  | 0.299 | #DIV/0! | 0.1814
3 | 0.403 |  |  | 0.403 | #DIV/0! | 0.24449
3.) Asynchrnous sweeping gap measurements
Adjacent leaves have different positions, constant gap width 20mm, MLCs move at constant speed.
s [cm] | R1 | R2 | R3 | AVERAGE [nC] | STD ERR. [nC] | RATIO
0.2 | 0.29 |  |  | 0.29 | #DIV/0! | 0.17594
0.5 | 0.269 |  |  | 0.269 | #DIV/0! | 0.1632
1 | 0.233 |  |  | 0.233 | #DIV/0! | 0.14135
2 | 0.225 |  |  | 0.225 | #DIV/0! | 0.1365