## Information
This excel workbook is an implementation of the procedure described in the paper that can be found at:
https://iopscience.iop.org/article/10.1088/1361-6560/ab8cd5
to obtain the optimal parameters for VARIAN MLCs in the RayStation TPS extended to Elekta's Agility.
MEASUREMENT SETUP
Use a Farmer-type ionization chamber. Longitudinal axis perpendicular to leaf motion.
Water or plastic phantom
SSD = 90 cm
Depth = 10 cm
Measure at two different longitudinal (Gun-Target) positions (Round 1 and Round 2) at (-x) cm and (+x) cm + 0.5*leafwidth.
Agility: | Round1 @ | -1.00 cm
| Round2 @ | +1.25cm

## 6 MV
SG-aSG Workflow for obtaining 'optimal' MLC configuration parameters in RayStation
1. Import the DICOM set provided into RayStation.
2. Calculate all the beams. Use the same number of MUs on every field.
3. Measure the data from the plan SG and aSG20 (Only these beams are require to obtain all the parameters)
4. Perform the measurements as described :
Measuring instructions
1. Farmer chamber. Long axis perpendicular to leaf motion.
2. Isocentric set-up. SSD 90 cm, depth 10 cm.
3. Water or real water.
4. (Recommended) Use two different measuring positions:
Round1 @ 1.00 cm towards Gun
Round2 @ 1.125 cm towards Target
(This is to further average the response taking into account more leafs)
5. Fill the data on the corresponding cells in columns C and D
6. Cells on J29:J34 contain the 'optimal parameters' calculated from your measured data.
(Option1): Verify that the optimal parameters minimized dose differences in SG/aSG beams | [6 empty] | Image not to scale.
7. Use the 'optimal parameters' in RayStation. Recalculate all the beams.
8. Fill the values on column 'E' with the average doses calculated by RayStation. Use the script provided to extract per beam doses.
9. Cells on M29:M34 should calculate the parameters you just introduce.
(Option 2): Assess the differences from your original set of parameters.
7. Use you 'original parameters' in RayStation. Recalculate all the beams.
8. Fill the values on column 'E' with the average doses calculated by RayStation. Use the script provided to extract per beam doses.
9. Cells on M29:M34 should calculate the parameters you just introduce.
Clinic:
Linac name :
Linac model:
MLC model: | HD120 |  |  | Parameters used in your clinical model: |  |  | Parameters recommended (measurement): |  | Parameters obtained (from TPS):
Energy (MV): | 6 |  |  | T = |  |  | T = | 0.01025 | T = | 0.01025
WFF/FFF: |  |  |  | LT(cm) = |  |  | LT(cm) = | 0.10284 | LT(cm) = | 0.10284
Leaf width (mm) | 2.5 |  |  | TG (cm) = |  |  | TG (cm) = | 0.042287 | TG (cm) = | 0.042287
|  |  |  | Offset (cm) = |  |  | Offset (cm) = | 0.019255 | Offset (cm) = | 0.019255
Contact person: |  |  |  | Gain (cm) = |  |  | Gain (cm) = | 0 | Gain (cm) = | 0
email: |  |  |  | Curvature (cm) = |  |  | Curvature (cm) = | 0 | Curvature (cm) = | 0
[11 empty] | Data from Sweeping Gaps.. |  |  | Data from Asynchronous Sweeping Gaps..
SG (DLG beams, s=0) |  |  |  |  | SG (DLG beams, s=0) | [6 empty] | Calc | Measured |  | Calc | Measured | The slope of the linear part of Delta_phi defines the tongue and groove width. The intercept with the x-axis of this linear part determines the leaf tip width.
Beam | ROUND1 | ROUND2 | TPS dose |  | AVERAGE | Measured dose | Diff (%) |  |  |  | lambda = | 0.61431 | 0.61431 | Slope = | 0.32224 | 0.32224
10x10 | 74.4917 |  | 74.4917 |  | 74.4917 | 74.4917 | 0% |  |  |  | gamma = | 1.0116 | 1.0116 | Intercept = | -0.33139 | -0.33139
5 | 4.0839 |  | 4.0839 |  | 4.0839 | 4.0839 | 0% |  |  |  | k = | 0.24827 | 0.24827
10 | 7.1524 |  | 7.1524 |  | 7.1524 | 7.1524 | 0% |  |  |  | delta [mm] = | 0.20201 | 0.20201 | LeafTipWidth (cm) | 0.10284 | 0.10284
20 | 13.3004 |  | 13.3004 |  | 13.3004 | 13.3004 | 0% | [6 empty] | TongueAndGroove (cm) | 0.042287 | 0.042287
30 | 19.4397 |  | 19.4397 |  | 19.4397 | 19.4397 | 0%
Transm_A | 0.76354 |  | 0.76354 |  | 0.76354 | 0.76354 | 0% | [6 empty] | xoff (cm) | 0.019255 | 0.019255 | The x-offset parameter can be determined from the leaf tip width and delta calculated previously with the SG data.
Transm_B | 0.76354 |  | 0.76354 |  | 0.76354 | 0.76354 | 0%
aSG20 |  |  |  |  | aSG20 | [5 empty] | Delta_phi (s)
s (mm) | ROUND1 | ROUND2 | TPS dose |  | AVERAGE | Measured dose | Diff (%) |  |  |  | s | Calc | Measured
0 | 13.3161 |  | 13.3161 |  | 13.3161 | 13.3161 | 0% |  |  |  | 0 | 0 | 0
2 | 13.0968 |  | 13.0968 |  | 13.0968 | 13.0968 | 0% |  |  |  | 2 | 0.44171 | 0.44171
5 | 12.679 |  | 12.679 |  | 12.679 | 12.679 | 0% |  |  |  | 5 | 1.2831 | 1.2831
10 | 11.8853 |  | 11.8853 |  | 11.8853 | 11.8853 | 0% |  |  |  | 10 | 2.8816 | 2.8816
15 | 11.076 |  | 11.076 |  | 11.076 | 11.076 | 0% |  |  |  | 15 | 4.5114 | 4.5114
20 | 10.282 |  | 10.282 |  | 10.282 | 10.282 | 0% |  |  |  | 20 | 6.1105 | 6.1105
Only for plotting purposes:
Optimal model
[18 empty] | s (mm) | DeltaPhis
[18 empty] | 0 | 0
[18 empty] | 1.0284 | 0
[18 empty] | 20 | 6.1135
Calculation model
[18 empty] | s (mm) | DeltaPhis
[18 empty] | 0 | 0
[18 empty] | 1.0284 | 0
[18 empty] | 20 | 6.1135
Note on B36: @ isocenter
Note on I44: (calc-meas)/meas
Note on I54: (calc-meas)/meas

## Other energies
SG-aSG tests (duplicate sheet for more energies)
Please return the excel workbook to your contact person at RaySearch.
Clinic: |  |  |  | Did you decide to use the model clinically?
Linac name : |  |  |  | Used for techniques:
Linac model:
MLC model: | Agility |  |  | Parameters used in your clinical model:
Energy (MV): | 6 |  |  | T =
WFF/FFF: |  |  |  | LT =
Leaf width (mm) | 5 |  |  | TG =
Offset =
Contact person: |  |  |  | Gain =
email: |  |  |  | Curvature =
SG (DLG beams, s=0) | [8 empty] | Calc | Measured |  | Delta_phi
Beam | ROUND1 | ROUND2 | TPS dose |  | AVERAGE | Measured dose | Diff (%) | lambda = | #DIV/0! | #DIV/0! | s | TPS | Measured
10x10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | gamma = | #DIV/0! | #DIV/0! | 0 | #DIV/0!
5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | k = | #DIV/0! | #DIV/0! | 2 | #DIV/0! | #NUM!
10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | delta [mm] = | #DIV/0! | #DIV/0! | 5 | #DIV/0! | #NUM!
20 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 8 | #DIV/0! | #NUM!
30 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 10 | #DIV/0! | #NUM!
Transm_A |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 15 | #DIV/0!
Transm_B |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 20 | #DIV/0! | #NUM!
[9 empty] | Delta_phi (s) |  | 30 | #DIV/0!
aSG20 | [7 empty] | s | Calc | Measured | 40 | #DIV/0!
s (mm) | ROUND1 | ROUND2 | TPS dose |  |  |  |  | 0 | #DIV/0!
0 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 2 | #DIV/0!
2 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 5 | #DIV/0!
5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 8 | #DIV/0!
8 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 10 | #DIV/0!
10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 15 | #DIV/0!
15 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 20 | #DIV/0!
20 |  |  |  |  | #DIV/0! |  | #VALUE!
25 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
30 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
aSG40 | [7 empty] | s | Calc | Delta_phi (s)
s (mm) | ROUND1 | ROUND2 | TPS dose |  |  |  |  | 0 | #DIV/0!
0 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 10 | #DIV/0!
10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 20 | #DIV/0!
20 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 30 | #DIV/0!
30 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 40 | #DIV/0!
40 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
If a compromise is needed, prioritize the region between 0 and 20 mm
aSG10 (optional) | [7 empty] | s | Calc | Delta_phi (s)
s (mm) | ROUND1 | ROUND2 | TPS dose |  |  |  |  | 0 | #DIV/0!
0 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 2 | #DIV/0!
2 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 5 | #DIV/0!
5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 8 | #DIV/0!
8 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 10 | #DIV/0!
10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
15 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
20 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
aSG5 (optional) | [7 empty] | s | Calc | Delta_phi (s)
s (mm) | ROUND1 | ROUND2 | TPS dose |  |  |  |  | 0 | #DIV/0!
0 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 2 | #DIV/0!
2 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 5 | #DIV/0!
5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
8 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
10 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
15 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
20 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
Note on B12: @ isocenter
Note on I20: (calc-meas)/meas