## Notes
References: | Hernandez, V. et al. (2018) ‘A new method for modelling the tongue-and-groove in treatment planning systems’, Physics in Medicine & Biology, 63(24), p. 245005. Available at: https://doi.org/10.1088/1361-6560/aaf098.
LINAC: | LA8
DATE: | 2023-01-02
DICOM Files: | U:\Liam Stubbington\beam-model\sweeping_gap\varian_dcm\
SETUP: | 10×10cm square field. Fields delivered to the EPID in Dosimetry mode.
NOTES:

## ImageJ
MACRO: | C:\Program Files\ImageJ\macros\SweepingGapTools.ijm
DLG
6X DLG
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 6x [MV], 400 [MU/min] | 22500 | 0.01195 | 0.000475 | 0.01186 | 0.01051 | 0.02051 | 595 | 595 | 268.9 | 0.0119 | 100 | 268.9 | 0.2 | 0.026597
4mm : 6x [MV], 400 [MU/min] | 22500 | 0.01915 | 0.0004783 | 0.01909 | 0.01769 | 0.02754 | 595 | 595 | 430.9 | 0.0191 | 100 | 430.9 | 0.4 | 0.042621
6mm : 6x [MV], 400 [MU/min] | 22500 | 0.02653 | 0.0004689 | 0.02642 | 0.02508 | 0.03472 | 595 | 595 | 596.9 | 0.02648 | 100 | 596.9 | 0.6 | 0.059041
10mm: 6x [MV], 400 [MU/min] | 22500 | 0.04139 | 0.0004618 | 0.0413 | 0.03991 | 0.0494 | 595 | 595 | 931.3 | 0.04134 | 100 | 931.3 | 1 | 0.092117
14mm: 6x [MV], 400 [MU/min] | 22500 | 0.05623 | 0.0004588 | 0.05612 | 0.05468 | 0.06446 | 595 | 595 | 1265 | 0.05619 | 100 | 1265 | 1.4 | 0.12512
16mm: 6x [MV], 400 [MU/min] | 22500 | 0.06367 | 0.0004578 | 0.06352 | 0.0621 | 0.07232 | 595 | 595 | 1432 | 0.06363 | 100 | 1432 | 1.6 | 0.14164
20mm: 6x [MV], 400 [MU/min] | 22500 | 0.07848 | 0.0004607 | 0.07845 | 0.07688 | 0.08754 | 595 | 595 | 1766 | 0.07845 | 100 | 1766 | 2 | 0.17468
Open: 6x [MV], 400 [MU/min] | 22500 | 0.4493 | 0.0009185 | 0.4496 | 0.443 | 0.4562 | 595 | 595 | 10110 | 0.4494 | 100 | 10110 | OPEN | 1
TransmA : 6x [MV], 400 [MU/min] | 22500 | 0.004006 | 0.0005898 | 0.003838 | 0.002365 | 0.01269 | 595 | 595 | 90.13 | 0.003945 | 100 | 90.13 | TRANS_A | 0.0089149
TransmB : 6x [MV], 400 [MU/min] | 22500 | 0.003854 | 0.0005357 | 0.003723 | 0.002407 | 0.01217 | 595 | 595 | 86.72 | 0.00379 | 100 | 86.72 | TRANS_B | 0.0085776
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X | 0.8746% | 0.0146 | 0.99125 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X | 0.082391 | 0.0098007 | 0.11895 | 0.013071
6X FFF
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.009998 | 0.0003332 | 0.009869 | 0.009027 | 0.01609 | 595 | 595 | 225 | 0.00997 | 100 | 225 | 0.2 | 0.021469
4mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.01749 | 0.0003549 | 0.01749 | 0.01634 | 0.02361 | 595 | 595 | 393.5 | 0.01747 | 100 | 393.5 | 0.4 | 0.037548
6mm : 6xFFF [MV], 400 [MU/min] | 22500 | 0.02515 | 0.0003903 | 0.02505 | 0.02385 | 0.03108 | 595 | 595 | 565.9 | 0.02514 | 100 | 565.9 | 0.6 | 0.053998
10mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.04063 | 0.0004845 | 0.04074 | 0.03917 | 0.04605 | 595 | 595 | 914.1 | 0.04064 | 100 | 914.1 | 1 | 0.087223
14mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.05606 | 0.0005931 | 0.05627 | 0.05422 | 0.06135 | 595 | 595 | 1261 | 0.05611 | 100 | 1261 | 1.4 | 0.12032
16mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.0638 | 0.0006572 | 0.06402 | 0.06183 | 0.06926 | 595 | 595 | 1435 | 0.06386 | 100 | 1435 | 1.6 | 0.13693
20mm: 6xFFF [MV], 400 [MU/min] | 22500 | 0.07922 | 0.0007821 | 0.07973 | 0.07689 | 0.08442 | 595 | 595 | 1782 | 0.07931 | 100 | 1782 | 2 | 0.17004
Open: 6xFFF [MV], 400 [MU/min] | 22500 | 0.4658 | 0.004229 | 0.4693 | 0.455 | 0.4763 | 595 | 595 | 10480 | 0.4663 | 100 | 10480 | OPEN | 1
TransmA : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002687 | 0.0003836 | 0.002499 | 0.001622 | 0.009228 | 595 | 595 | 60.45 | 0.002636 | 100 | 60.45 | TRANS_A | 0.0057681
TransmB : 6xFFF [MV], 400 [MU/min] | 22500 | 0.002549 | 0.0003444 | 0.002433 | 0.001589 | 0.008677 | 595 | 595 | 57.35 | 0.002498 | 100 | 57.35 | TRANS_B | 0.0054723
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 6X FFF | 0.562% |  | 0.99438 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 6X FFF | 0.082673 | 0.0046117 | 0.055783 | -0.012041
10X FFF
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.01559 | 0.000581 | 0.01549 | 0.01378 | 0.02319 | 595 | 595 | 350.8 | 0.01555 | 100 | 350.8 | 0.2 | 0.029729
4mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.02399 | 0.0006609 | 0.02385 | 0.02198 | 0.03149 | 595 | 595 | 539.8 | 0.02397 | 100 | 539.8 | 0.4 | 0.045746
6mm : 10xFFF [MV], 400 [MU/min] | 22500 | 0.03257 | 0.0007504 | 0.03271 | 0.03045 | 0.03971 | 595 | 595 | 732.9 | 0.03257 | 100 | 732.9 | 0.6 | 0.06211
10mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.04989 | 0.0009795 | 0.05017 | 0.04717 | 0.05658 | 595 | 595 | 1123 | 0.04994 | 100 | 1123 | 1 | 0.095169
14mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.06713 | 0.001234 | 0.06785 | 0.06383 | 0.07367 | 595 | 595 | 1510 | 0.06723 | 100 | 1510 | 1.4 | 0.12797
16mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.07576 | 0.001354 | 0.07631 | 0.07212 | 0.08157 | 595 | 595 | 1705 | 0.07589 | 100 | 1705 | 1.6 | 0.14449
20mm: 10xFFF [MV], 400 [MU/min] | 22500 | 0.09304 | 0.001618 | 0.09365 | 0.08883 | 0.09843 | 595 | 595 | 2093 | 0.09322 | 100 | 2093 | 2 | 0.17737
Open: 10xFFF [MV], 400 [MU/min] | 22500 | 0.5246 | 0.008596 | 0.5291 | 0.5043 | 0.5474 | 595 | 595 | 11800 | 0.5253 | 100 | 11800 | OPEN | 1
TransmB : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006185 | 0.0005687 | 0.005936 | 0.004581 | 0.01383 | 595 | 595 | 139.2 | 0.006119 | 100 | 139.2 | TRANS_B | 0.011797
TransmA : 10xFFF [MV], 400 [MU/min] | 22500 | 0.006352 | 0.0006326 | 0.006026 | 0.004666 | 0.01447 | 595 | 595 | 142.9 | 0.006269 | 100 | 142.9 | TRANS_A | 0.01211
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 10X FFF | 1.195% |  | 0.98805 | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 10X FFF | 0.082144 | 0.013021 | 0.15851 | 0.013336
10X
Label | Area | Mean | StdDev | Mode | Min | Max | X | Y | IntDen | Median | %Area | RawIntDen | Gap [cm] | EPID Signal
2mm : 10x [MV], 400 [MU/min] | 22500 | 0.01426 | 0.0005068 | 0.01404 | 0.0126 | 0.0233 | 595 | 595 | 320.8 | 0.01421 | 100 | 320.8 | 0.2 | 0.031952
4mm : 10x [MV], 400 [MU/min] | 22500 | 0.02139 | 0.0005039 | 0.02113 | 0.01969 | 0.03003 | 595 | 595 | 481.3 | 0.02134 | 100 | 481.3 | 0.4 | 0.047938
6mm : 10x [MV], 400 [MU/min] | 22500 | 0.02869 | 0.0005005 | 0.02854 | 0.02699 | 0.0373 | 595 | 595 | 645.6 | 0.02864 | 100 | 645.6 | 0.6 | 0.064303
10mm: 10x [MV], 400 [MU/min] | 22500 | 0.0434 | 0.0004975 | 0.04329 | 0.04166 | 0.05221 | 595 | 595 | 976.4 | 0.04335 | 100 | 976.4 | 1 | 0.097251
14mm: 10x [MV], 400 [MU/min] | 22500 | 0.05807 | 0.0004924 | 0.05808 | 0.05636 | 0.06672 | 595 | 595 | 1307 | 0.05803 | 100 | 1307 | 1.4 | 0.13018
16mm: 10x [MV], 400 [MU/min] | 22500 | 0.06543 | 0.0004862 | 0.06536 | 0.0638 | 0.0742 | 595 | 595 | 1472 | 0.06539 | 100 | 1472 | 1.6 | 0.14661
20mm: 10x [MV], 400 [MU/min] | 22500 | 0.08007 | 0.0004844 | 0.07997 | 0.07839 | 0.0886 | 595 | 595 | 1802 | 0.08003 | 100 | 1802 | 2 | 0.17948
Open: 10x [MV], 400 [MU/min] | 22500 | 0.4464 | 0.000894 | 0.4468 | 0.4406 | 0.4542 | 595 | 595 | 10040 | 0.4465 | 100 | 10040 | OPEN | 1
TransmB : 10x [MV], 400 [MU/min] | 22500 | 0.005598 | 0.0005701 | 0.005509 | 0.003779 | 0.01445 | 595 | 595 | 126 | 0.005538 | 100 | 126 | TRANS_B | 0.01255
TransmA : 10x [MV], 400 [MU/min] | 22500 | 0.005762 | 0.0006279 | 0.005699 | 0.003893 | 0.01531 | 595 | 595 | 129.7 | 0.005698 | 100 | 129.7 | TRANS_A | 0.012918
| EPID | RayStation | (1-T) | d [cm]
TRANSMISSION 10X | 1.273% |  | 98.73% | 12
| GRADIENT [cm-1] | INTERCEPT | INTER/GRAD | DLG [cm]
DLG 10X | 0.082093 | 0.015235 | 0.18558 | 0.030804

## Transmission
LINAC: | LA7
DATE: | 2023-01-16
DETECTOR:
ELECTROMETER: | E13
DICOM Files: | U:\Liam Stubbington\beam-model\sweeping_gap\varian_dcm
SETUP: | 10×10cm square field. Chamber @ +5.0cm depth in blocks. Chamber orientated perpendicular to the direction of MLC travel. For closed MLC fields, plan MU overridden in service to 500MU. Factor 0.2 applied in calculated transmission.
6X
| R1 | R2 | R3 | AVERAGE [nC]
OPEN | 1.66 | 1.66 |  | 1.66
MLC_A | 0.097 | 0.095 |  | 0.096
MLC_B | 0.093 | 0.093 |  | 0.093
TRANSMISSION [%] |  |  |  | 1.139%
6X FFF
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]
10X
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]
10X FFF
| R1 | R2 | R3 | AVERAGE [nC]
OPEN
MLC_A
MLC_B
TRANSMISSION [%]